17975 | 1220648-2024-19738 | 20396908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396908&pc=OZD | 11/23/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 11/23/2022 | 10/08/2024 |  |  | 08/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254602 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped in left ventricle position; with no restart possible, the impella was pulled back. they muted the alarms, position and retrograde flow, but aic was still alarming. pump was unable to be explanted due to deregulated coagulation and explanation via manta was not possible after 3 weeks support due to risk of infection. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. no clinical details provided were sufficient to determine a root cause. impella cp used beyond indication for use. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
17976 | 1220648-2024-19737 | 20396914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20396914&pc=OZD | 09/11/2024 | 11/19/2024 |  | 10/08/2024 |  |  | 09/11/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533763 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed for support of a patient admitted into a tertiary center from a community hospital. the patient was in acute myocardial infarction/cardiogenic shock. the pump support was for 29 hours and then explanted with pump thrombosis. the patient survived the event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation has been completed. the pump was returned for investigation. there was heavy biomaterial observed in the outflow cage. the failure could not be reproduced since the pump was cut from the catheter. data logs were returned and investigated. at the end of the case, a motor current spike was seen followed by low flow suggesting biomaterial ingestion. therefore, the root cause of the low pump flow issue was biomaterial. b.2 and h.1 revised as the previous selection was incorrect on initial manufacturer device report number. h.6 revised helath effect - impact code as the previous selection was incorrect on initial manufacturer device report number.
17977 | 1220648-2024-19739 | 20397034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397034&pc=OZD | 12/12/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 12/12/2022 | 10/08/2024 |  |  | 06/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238033 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the broken purge sidearm was broken. it was noted that the staff broke the purge sidearm in half during insertion. consequently, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation into the purge leak - pump has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the purge leak - pump was most likely a broken sidearm reservoir-to-filter joint due to improper handling of the impella. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
17978 | 1220648-2024-19740 | 20397095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397095&pc=OZD | 01/29/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 01/29/2023 | 10/08/2024 |  |  | 08/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257954 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. post implant, the device did not start. therefore, device was replaced with another impella cp pump. there was no reported patient harm. | [5 empty] | The investigation into the impella cp pump did not start has been completed. the impella pump was returned for investigation. during investigation, it was noted that the device was able to be started but stopped after a few minutes when the system was not purging. it was then run again with the system purging, and it ran at p-9 for 10 minutes without any problems. the motor was torn down and there was no evidence of biomaterial found, and all other parameters were within specifications. the cause of the reported issue was not able to be determined. this report is being filed as part of a retrospective review of historical records.
17979 | 1220648-2024-19742 | 20397140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397140&pc=OZD | 01/27/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 01/27/2023 | 10/08/2024 |  |  | 05/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229225 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. physician was unable to deliver pump due to severe peripheral vascular disease; the physician was unable to pass through vasculature. impella procedure was aborted. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17980 | 1220648-2024-19741 | 20397168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397168&pc=OZD | 01/27/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 01/27/2023 | 10/08/2024 |  |  | 04/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225821 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was performing with lower-than-expected flows. troubleshooting was unable to resolve the issue. a kink was noted in the device. the pump was removed and replaced with intra-aortic balloon pump support with no reported harm to the patient. | [5 empty] | The investigation for the reported product damage (cannula kink) has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were sufficient to determine a root cause. based on the clinical account, the root cause of the low flows were due to a cannula kink which was a result of the patient's left ventricle and ascending aorta anatomy. this report is being filed as part of a retrospective review of historical records.
17981 | 1220648-2024-19743 | 20397213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397213&pc=OZD | 09/10/2024 | 08/21/2025 |  | 10/08/2024 |  |  | 09/09/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526672 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Sepsis | Ischemia Stroke | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during impella 5.5 support for 38-year-old male patient, there was bleeding at the groin site. the repo sheath angle was lowered, dressing was changed, and bleeding subsided. the patient experienced fever and there were concerns of septic shock source. patient was on argatroban due to being heparin induced thrombocytopenia (hit) positive. suctions alarms were present. patient was given 1 unit of blood to resolve. additionally, the patient required thrombectomy in the setting of acute ischemic stroke. pump was ultimately explanted and a new impella 5.5 was placed. patient is noted to be stable. | The complainant reported the patient was on impella cp support and during support, the patient had hemolysis. the p level was reduced due to the hemolysis. the pump was explanted and replaced with the impella 5.5. it is unknown if the pump was explanted due to the hemolysis. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. b1 revised to only represent the adverse event as malfunction erroneously entered on the initial manufacturer device report number. b3 date of event revised as erroneously entered on the initial manufacturer device report number. b5 describe event or problem revised as previous information erroneously entered on the initial manufacturer device report number. d1 brand name was erroneously entered on the initial manufacturer device report number. d4 model number was erroneously entered on the initial manufacturer device report number. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h6 health effect clinical revised as erroneously entered on the initial manufacturer device report number. h6 health effect impact code revised as erroneously entered on the initial manufacturer device report number. h6 investigation conclusions revised from the initial submission in accordance with updated procedures. h10 additional manufacturer narrative revised from the initial submission in accordance with updated procedures.
17982 | 1220648-2024-19745 | 20397370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397370&pc=OZD | 01/31/2023 | 10/08/2025 |  | 10/08/2024 |  |  | 01/31/2023 | 10/08/2024 |  |  | 03/03/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200406 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A japan complainant reported that that a patient (pt) was placed on impella cp smart assist support for cardiogenic shock (cgs)/acute myocardial infarction (ami). it was reported that the impella due to an aortic dissection. after removal, it was noted that the pigtail was kinked. | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while trying to insert the impella, the physician was unable to complete due to an aortic dissection. after removal, it was noted that the pigtail was kinked. no further information is available. |  |  |  |  | Tthis report is being filed as part of a retrospective review of historical records.he investigation for product damage (cannula kink) has been completed. the device was not returned for evaluation. the cause of the pigtail kink observed after removal was most likely use related. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
17983 | 1220648-2024-19744 | 20397380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397380&pc=OZD | 09/17/2024 | 11/05/2024 |  | 10/08/2024 |  |  | 09/17/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528839 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp had high purge pressure system blocked alarm. tissue plasminogen activator protocol started but a decision was made to remove impella in the cath lab. the patient outcome is unknown, | [5 empty] | The investigation for the reported high purge pressure has been completed. neither the impella device nor the data logs were returned for evaluation. additionally, clinical details provided were limited to determine the root cause. therefore, the root cause of the high purge pressure could not be determined. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi states the following: section: using the automated impella controller with the impella rp with smartassist. system catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ added- h6 impact code 4644. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
17984 | 1220648-2024-19748 | 20397407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397407&pc=OZD | 02/13/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 02/13/2023 | 10/08/2024 |  |  | 06/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236755 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hematoma and bleeding. it was reported the patient was brought to the catheterization laboratory for an angiogram of the left groin prior to the impella removal and in case there was need for balloon tamponade or covered stent. there was no extravasating noted on the angiogram. the impella was removed, and manual compression held for seventy-five minutes, with femostop placed for minor track oozing. the patient tolerated the procedure well. the patient received a total of four units of packed red blood cells throughout the case. there was no further bleeding or hematoma, and the patient was returned to the intensive care unit (icu) with femostop in place to belt pressure. the bleeding was felt to be related to the patient being on integrillin drip. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation into the reported access site bleeding has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the reported issue was not determined. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
17985 | 1220648-2024-19752 | 20397441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397441&pc=OZD | 09/11/2024 | 11/04/2024 |  | 10/08/2024 |  |  | 09/11/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485862 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an impella cp implant and the flows were 3.2 in auto. it was found that the pump was kinked in the cannula in the ascending aorta. shortly after, there was a suction event with decreased flows. the pump was attempted to be repositioned without success. the impella was removed and attempted to be placed again with no avail. the pump was explanted and replaced successfully. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of low pump flow has been completed. the data logs confirmed the failure mode and clinical narrative. logs showed after the pump started and run for four minutes, low flow occurred that triggered auto suction response and suction alarms. logs also showed the pump was in improper position and a pump stop alarm triggered but resolved quickly. this event remained for the rest of the case.visual inspection found a kink on the cannula about 15 millimeters from cage and cannula bonding site. no other issues were found on the rest of the pump. the pump ran without issue under bench test. the root cause of low flow was a kinked cannula. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
17986 | 1220648-2024-19746 | 20397442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397442&pc=OZD | 02/13/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 02/13/2023 | 10/08/2024 |  |  | 12/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024296383 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Failure to Pump | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported by the customer that the pump did not start, but upon further investigation, it was clarified that the pump was working but experienced low pump flows. as a result, device was replaced with another impella cp pump. no patient harm was reported. | [5 empty] | The investigation has been completed. the lt log shows that the pump was run but was experiencing low pump flows. the returned pump ran within specification in the lab. no visual abnormalities were observed. the cause of the low pump flow was most likely patient condition related.
17987 | 1220648-2024-19747 | 20397446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397446&pc=OZD | 09/16/2024 | 11/01/2024 |  | 10/08/2024 |  |  | 09/16/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482979 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump for a patient admitted in cgs and cardiomyopathy. it was reported that upon explant of impella 5.5 there was concern for thrombus in right axillary due to dampened right radial arterial line. after left ventricular assist device placed, computed tomography surgery team performed a right axillary thrombectomy in operating room to retrieve thrombus with success and improved pulsitility on right radial arterial line. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
17988 | 1220648-2024-19753 | 20397467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397467&pc=OZD | 09/16/2024 | 10/14/2024 |  | 10/08/2024 |  |  | 09/16/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 230410 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, the automated impella controller (aic) displayed purge disc not detected alarm. two cassettes were tried displaying the same alarm. the console was changed and the issue was resolved. no reported patient harm. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. after reviewing the logs, the issues with the aic not being able to detect the purge disc was confirmed. upon examining aic for any visible defects, it was evident that the purge disc flag was broken (missing at blue retainer). the root cause of the issue was a broken purge disc flag. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-19753. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
17989 | 1220648-2024-19754 | 20397475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397475&pc=OZD | 02/14/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 02/14/2023 | 10/08/2024 |  |  | 10/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290385 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported during placement of the impella the physician experienced difficulty pulling the guide wire out during pump delivery. the decision was made to replace the pump, which was successful. no patient harm was reported. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella pump was returned for investigation. the visual investigation of the returned guidewire noted the guidewire was frayed. the cause of the delivery issue was use related damaged frayed guidewire. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
17990 | 1220648-2024-19749 | 20397487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397487&pc=OZD | 09/16/2024 | 11/07/2024 |  | 10/08/2024 |  |  | 09/16/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526152 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and they had hemolysis. the patient had dark red-brown urine and it was suspected that the patient had hemolysis due to previously deep impella placement. the pump was repositioned appropriately. continuous venovenous hemofiltration (cvvh) was started due to no urine output and the patient's lactate dehydrogenase (ldh) was 2400. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis issues has been completed. data logs were reviewed: finding no motor current (mc) spike observed. position alarms occurred in aorta. placement signal (ps) low alarmed as suction alarms. clinical details confirmed that impella repositioned with echo tech due to deep placement (from about 7cm av to inlet to 5.7cm av to inlet). impella was deeper than ideal position. clinical details didn't confirm if the issue solved but no impella related issue confirmed after several attempts to repositioning. the cause of the hemolysis issue most likely was improper position related due to improper technique.
17991 | 1220648-2024-19751 | 20397495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397495&pc=PYX | 09/16/2024 | 11/06/2024 |  | 10/08/2024 |  |  | 09/16/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025508325 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella rp support ongoing for 6 days when there were signs of hemolysis. labs hemolyzing only rectified with p level adjustment. the pump was instead explanted to resolve the hemolysis. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the hemolysis has been completed. data logs were investigated and there were several abnormalities. there is a motor current spike and speed deviation that does not align with any p-level changes. additionally, just leading up to these spikes, there is noise/spikes in the cvp. these two factors indicate that there was likely something in the cannula that was interacting with the pressure sensors and also impeded the rotation of the impeller. this event is not followed by any change in the purge pressures/flows or ps, however. on the same days as patient's urine color change, several high purge pressure (hpp) spike events are seen. there are 3 distinct spikes in the purge pressure that all reach close to or above 1000 mmhg. none of these events correlate with a purge bag/cassette change, and they all resolve. the only abnormality leading up to these events is a drop in the motor current, rather than a spike up, while the pump is running constant at p-6. additionally, during the time of pp spikes, the cvp spikes up, indicating something could be interacting with the sensor faces. lastly, from on (b)(6) through to the end of the case, the pump is not able to achieve flows within specification for the p-levels it is running at. all these factors indicate that starting on (b)(6) through the time that hemolysis symptoms are noted, there could be biomaterial ingestion/build up in the cannula. returned product was investigated and biomaterial was seen to be in the purge gap and wrapped around the impeller, even after being soaked in water/sporacidin. based on the data logs and returned product, it was determined that the root cause of the hemolysis was biomaterial.
17992 | 1220648-2024-19755 | 20397526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397526&pc=OZD | 09/19/2024 | 11/19/2024 |  | 10/08/2024 |  |  | 09/19/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025542554 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 62-year-old male patient with cardiomyopathy and other systemic comorbidities was implanted with an impella cp device for mechanical circulatory support. during impella support, there was significant damage to the catheter when the patient was turned. the damage resulted in a pump stop. the pump was removed. there was no patient harm reported. | [5 empty] | The investigation of pump stop has been completed. the pump was returned for investigation. the data log shows several 115-impella stopped alarms with a motor current spike at the end of the case run. the catheter was found completely detached with the kink protector from the red handle, which caused all three electrical lines to be open, resulting in the pump stopping. according to the clinical notes, the catheter was damaged while turning the patient. the cause of the pump stop issue was an open electrical line found at rend handle. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
17993 | 1220648-2024-19750 | 20397565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397565&pc=OZD | 09/16/2024 | 11/22/2024 |  | 10/08/2024 |  |  | 09/16/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025491252 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. it was reported that serosanguinous fluid was found inside the sterile sleeve of the 5.5. no reported patient harm and the patient remains on support. | The decision was made to withdraw care and the patient expired. there was no sepsis or infection against the device. the patients condition was poor peri-operatively. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into product damage (sterile sleeve damage) has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-19750. b5 patient outcome and mso statement added d6b missing e4 should have been left blank.
17994 | 1220648-2024-19756 | 20397654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397654&pc=OZD | 09/16/2024 | 11/12/2024 |  | 10/08/2024 |  |  | 09/16/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521385 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp placed for a patient admitted in with multiple cardiac risk factors for a high risk percutaneous coronary intervention. the cp had low flows and was replaced with a new pump after just over an hour of support. no patient harm reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the low pump flow has been completed. pump was returned for analysis. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue during bench test. the root cause of the low flow was most likely a patient condition as no issue was found on the pump.
17995 | 1220648-2024-19760 | 20397664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397664&pc=OZD | 09/19/2024 | 11/18/2024 |  | 10/08/2024 |  |  | 09/19/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025460362 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant had an impella cp pump placed along with ecmo to support a 64-year-old male patient admitted suffering from an acute myocardial infarction / cardiogenic shock (ami/cgs). the cp was placed as escalation from the intra-aortic balloon pump (iabp). the team found high purge pressures/purge blocked after 12 days of support. the team used tissue plasminogen activator (tpa) in the purge to troubleshoot. |  |  |  |  | The investigation for the reported high purge pressure has been completed. neither the impella device nor the data logs were returned for evaluation. additionally, clinical details provided were limited to determine the root cause. therefore, the root cause of the high purge pressure could not be determined. added-b1-selected adverse event, b2-selected death and added date of death, b5-mso review, h6 impact code 4644, h6 impact code 1802 corrections to the initial mfr report: h1-type of reportable event changed to death. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
17996 | 1220648-2024-19764 | 20397681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397681&pc=OZD | 09/17/2024 | 09/26/2025 |  | 10/08/2024 |  |  | 09/17/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1438016 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during prep of the automated impella controller there was a red alarm error. the controller was replaced and no patient harm was reported. | [5 empty] | The investigation for the console (aic) controller failure-red alarm has been completed. there was no evidence in the logs to confirm the reported failure mode. data logs p (psnr) ¿placement signal not reliable (opm invalid signal, optical pump connected]) ¿ alarm,¿ event #1036 during the support. it resolved automatically in 10 minutes. event #1036 occurred after 4 days 19 hours since the pump started, but there were 1036 error logs ¿processsampleforopminvalidsignalerror: calculator placement signal 1036¿ throughout the case, which is unrelated to the reported failure mode. a placement signal issue was observed. there was no placement signal issue with the cases performed post-27 sep 2024. this console returned for evaluation and obtained an abnormal envelope - without connecting either the optical test pump or test cavity with aic and abnormal fringe pattern - with the optical test cavity connected to the aic and indicating defective optical bench. the cause of the controller failure was corrupted epm firmware. it is undetermined what caused the initial corruption of the firmware. the root cause for psnr was most likely due to the defective optical bench. added-d9 device return date corrections to the initial mfr report: d2 updated common device name g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
17997 | 1220648-2024-19757 | 20397716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397716&pc=OZD | 09/16/2024 | 10/14/2024 |  | 10/08/2024 |  |  | 09/16/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1805850 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller when turned on and plugged in, gives a battery alarm. the controller was returned for investigation. there was no patient involvement. | [5 empty] | The investigation for the console (aic) battery failure red alarm has been completed. the reported ¿ac on battery¿ issue was confirmed in the logs. the reported issue was not an alarm but a prompt that appears when the aic is disconnected from ac power. there were numerous ¿ac power disconnected - controller is running on battery power¿ (event set #109) prompts observed on the reported complaint date. the console was returned for evaluation. the reported issue was able to be reproduced in the lab. the aic would not charge or recognize ac power when connected to the wall outlet. this issue was determined to be caused by a defective ac power cord. console was tested with a different power cord and the aic was able recognize ac power without any issues. corrections to the initial mfr report: d2-common device name. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
17998 | 1220648-2024-19761 | 20397721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397721&pc=OZD | 09/17/2024 | 09/22/2025 |  | 10/08/2024 |  |  | 09/17/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481902 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was on impella 5.5 support and there were suction alarms above p-4. the patient has no other issues her purge flow and pressure is all within normal limits and no other issues, flow is appropriate for p level currently at. the patient remained on support and stable until successful weaning of device. | The automated impella controller (aic) was also starting to have impella position unknown alarms. the doctor confirmed the position with echocardiography. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. no product was returned for analysis. the clinical states that there was concern for patient volume issues. there were no spikes in motor current seen, but a trend in the placement signal. lower flows were seen at higher p-levels. suction alarms were seen during the changing of p-levels but then they stabilized at a lower p-level, indicating the patient may not have been able to tolerate higher p-levels. based on the clinical details provided and the data log analysis, the root cause of the low pump flow is most likely due to the patients condition. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. at p-7 the placement signal trends with the lower flows and the signal-to-noise ratio (snr) varies. when the p-level changes from 7 to 4, suction, positioning and placement signal low alarms are seen. there are two placement signal unreliable alarms that are triggered in the case when the pump is turned off. the rest of the 5s parameters are stable throughout the case. based on this trend in the placement signal and flows, the root cause is most likely due to the patient¿s condition. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-19761. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. e4 erroneously entered as "no" on the initial report. g1 reporting contact email updated.
17999 | 1220648-2024-19758 | 20397723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397723&pc=OZD | 02/14/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 02/14/2023 | 10/08/2024 |  |  | 10/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024288747 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. during implant, device had delivery issues. physician was unable to cross the device across aortic valve (av). reportedly, the impella placement was aborted due to the patient's anatomy. an intra-aortic balloon pump was placed due to patient being on bypass for eight hours and being unstable. the impella placement was aborted. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not received from the customer, therefore, investigation of the device was not possible. the clinical investigation determined that the root cause of the delivery issues was related to the patient condition. this report is being filed as part of a retrospective review of historical records.
18000 | 1220648-2024-19763 | 20397727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397727&pc=OZD | 02/14/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 02/14/2023 | 10/08/2024 |  |  | 09/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275743 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported a temporary pacer was placed due to the patient¿s heart rate during the delivery and placement of the impella. there was no harm to the patient. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. per the clinical investigation, the root cause of the delivery issue was most likely patient condition related to right bundle branch block of the heart. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18001 | 1220648-2024-19768 | 20397749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397749&pc=OZD | 09/19/2024 | 11/08/2024 |  | 10/08/2024 |  |  | 09/19/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544161 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the 81-year-old female patient was having an impella cp implant and when the staff were attempting to push blue dilator through valve in peel away sheath, it appeared to not have a hole to pass through. user attempted multiple times and wasn¿t able to get dilator in the sheath. there was no patient harm reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the introducer damage has been completed. pump and introducer were not returned for investigation. data log review was not required for this case. as per the clinical information, the introducer did not have a hole to pass through the dilator, and the doctor wasn¿t able to insert the dilator into the sheath. the cause of the introducer damage issue was not determined since product was not returned and sufficient clinical information was not provided.
18002 | 1220648-2024-19769 | 20397752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397752&pc=OZD | 09/19/2024 | 11/15/2024 |  | 10/08/2024 |  |  | 09/19/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025542912 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 55-year-old male patient was on impella 5.5 support due to st elevated myocardial infarction (stemi). while on support, there were issues with flows due to a small left ventricle. the pump was running at p4 and the patient was receiving fluid which helped with flows. the impella was positioned against mitral apparatus and doctor was unable to reposition it away from wall. additionally, the patient was cardioverted out of atrial fibrillation (afib)/ rapid ventricular response (rvr). the pump was successfully explanted after four days of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The investigation for the reported low pump flow, atrial arrhythmia, and positioning issues has been completed. product was not returned for investigation. data logs were reviewed when investigating the low pump flows, and there was only one positioning alarm throughout the entire case and it occurred on the second to last day of support, so it was not related to the complication reported at case start. clinical details provided also report that the patient had a clearly small lv and that any attempts to reposition the pump did not prevent contact with the heart wall/mitral apparatus. based on clinical details provided and data logs, the root cause of the positioning issues was determined to be patient condition. there are 13 suction alarms triggered. additionally, the ps appears relatively noisy. that being said, there are no motor current spikes that don't align with p-level changes, and there is only one positioning alarm throughout the entire case and it occurs on the second to last day of support. throughout the case, pump flows are rarely below the minimum specification for the p-level that the pump is running at. however, for the majority of the case, pump flows are fluctuating and variable even if p-level is being held constant for a matter of days or hours. clinical details report that the patient had a relatively small lv and this was believed to be the reason for what they considered were low flows. finally, it was reported that the patient was being given fluid which helped improve flows. the pump was not returned for investigation. based on clinical details provided and data log review, the root cause of the low pump flows was determined to most likely be patient condition. the root cause of the atrial arrhythmia could not be determined without additional clinical details.
18003 | 1220648-2024-19765 | 20397763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397763&pc=OZD | 09/17/2024 | 11/01/2024 |  | 10/08/2024 |  |  | 09/17/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025513201 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed along with extracorporeal membrane oxygenation for a patient undergoing thromboembolectomy and left lower extremity thrombosis. it was reported that the patient experienced a cerebral vascular accident while on support. neuro protective protocol implemented. currently not a candidate for advanced therapies. a few days later the pump was explanted and the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. added- h6 impact code 4641.
18004 | 1220648-2024-19766 | 20397779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397779&pc=OZD | 09/17/2024 | 09/24/2025 |  | 10/08/2024 |  |  | 09/17/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528132 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was on impella 5.5 support and they had a thrombectomy in their right arm. the patient remained stable and on support until the patient was successfully weaned off the device. | The complainant also reported that there were placement signal issues. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-19766. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. product was not returned for investigation. data logs show the placement signal not reliable and controller error, along with an oscillating trend in contrast. this trend periodically returned to normal during case run. upon review of data logs, no problem was found with the pump and it is apparent the console was the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-19766. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
18005 | 1220648-2024-19770 | 20397813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397813&pc=OZD | 09/23/2024 | 12/09/2024 |  | 10/08/2024 |  |  | 09/23/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025508265 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 21 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | BALLOON PUMP | EXTRA CORPOREAL MEMBRANE OXYGENATION | [17 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed limb ischemia and the impella was removed. the patient remained stable and survived the event. | [5 empty] | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly on manufacturer device report 1220648-2024-19770. e4 should have been left blank. g1 reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18006 | 1220648-2024-19759 | 20397837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397837&pc=OZD | 09/16/2024 | 11/12/2024 |  | 10/08/2024 |  |  | 09/16/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471108 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during repositioning, the impella 5.5 pulled back across the valve and the impella was in the aorta. the physician attempted to cross the aorta valve (av) with impella in intensive care unit but was unable to position back in the left ventricle (lv). the patient was taken back to the operation room and the 5.5 was unable to cross av. the impella had a small kink due to the attempt of crossing the lv. the impella was remove and a new impella 5.5 was placed. the patient was noted to be stable and no reported patient harm. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issues was most likely use-related. e4 should have been left blank on manufacturer device report 1220648-2024-19759. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
18007 | 1220648-2024-19771 | 20397863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397863&pc=OZD | 02/15/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 02/15/2023 | 10/08/2024 |  |  | 10/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285491 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella flows were low. the physician attempted to use a new automated impella controller and purge cassette with no resolution. reportedly, impella flows were continuing to drop with multiple alarms despite correct placement of the impella. device was removed and replaced with a new impella pump. no further issues were noted and there was no harm reported to the patient. | [5 empty] | The investigation into the low pump flow has been completed. the impella pump was returned for investigation. the product evaluation showed that there was a kink in the cannula as well as a kink in the catheter. the root cause of the low pump flow was patient condition related. this report is being filed as part of a retrospective review of historical records.
18008 | 1220648-2024-19762 | 20397871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397871&pc=OZD | 09/11/2024 | 11/07/2024 |  | 10/08/2024 |  |  | 09/11/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482909 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when placing the impella 5.5, the cannula had difficulty passing through the axillary and subclavian arteries and would not pass the origin of innominate. the placement of the impella 5.5 was unsuccessful despite numerous wire attempts, including buddy wire. the decision was made to abort the procedure. no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the impella pump implant was aborted due to patient's anatomy. h.6 code 2682 was inadvertently omitted from initial manufacturer device report number 1220648-2024-19762 and was added. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
18009 | 1220648-2024-19767 | 20397885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397885&pc=OZD | 09/18/2024 | 11/07/2024 |  | 10/08/2024 |  |  | 09/18/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108096 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported that during prep the rear housing of the automated impella controller (aic) was found to be cracked. there was no patient involvement. | [5 empty] | The investigation for the console damage has been completed. data logs did not contain any data relevant to this complaint. the console was returned for evaluation finding a cracked rear housing was confirmed. further inspection of the aic revealed that the front housing and purge insert were also damaged. aside from physical damage of the aic and the issue downloading logs, the aic functioned as normal. the root cause of the damage was use related. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
18010 | 1220648-2024-19772 | 20397893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20397893&pc=OZD | 09/11/2024 | 05/13/2025 |  | 10/08/2024 |  |  | 09/11/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534708 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Contamination | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was opened but the impella 5.5 was contaminated before use. the case was completed using another device, extracorporeal membrane oxygenation. no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of packaging issue has been completed. clinical reports impella 5.5 was open and contaminated before use. the cause of the packaging issue sterility was not determined since product was not returned and lack of clinical details. d6a and d6b should have been left blank on the initial manufacturer device report number 1220648-2024-19772, as the product was not used. g1 reporting contact email revised on manufacturer device report 1220648-2024-19772 in accordance with updated procedures. h6 code 4642 was reported incorrectly on manufacturer device report 1220648-2024-19772. new code was added to health effect - impact code.
18011 | 1220648-2024-19774 | 20398161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20398161&pc=OZD | 09/11/2024 | 11/19/2024 |  | 10/08/2024 |  |  | 09/11/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025524551 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and there were positioning issues. the pump migrated into the aorta. it was not addressed and the patient arrested without support. the pump was explanted and the patient was put on extracorporeal membrane oxygenation. the patient eventually expired after care was withdrawn. upon review by abiomed's medical safety officer, the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue most likely was use related due to improper technique as the impella pulled out to aorta. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
18012 | 1220648-2024-19773 | 20398262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20398262&pc=OZD | 09/11/2024 | 11/06/2024 |  | 10/08/2024 |  |  | 09/11/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531283 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a hematoma was noted in intensive care unit at the impella cp insertion site. the hematoma was originally present when intra-aortic balloon pump was inserted. issues seem to have been resolved but overnight the hematoma grew. hemoglobin and hematocrit dropped. the patient became hypotensive requiring blood products and vasopressors. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the access site bleeding could not be determined. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18013 | 1220648-2024-19775 | 20398349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20398349&pc=OZD | 09/15/2024 | 11/18/2024 |  | 10/08/2024 |  |  | 09/15/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486841 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella 5.5 pump to support a patient admitted in with cardiomyopathy. the pump was supporting but there were purge alarms. the team noted purge pressure high and purge blocked. the team attempted troubleshooting with tissue plasminogen activator in the purge. however, it is unknow if this resolved the alarm. there was no reported patient harm and the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for the high purge pressure has been completed. clinical details provided limited information that could explain why the purge system blocked alarm would have triggered. additionally, the means of addressing/resolving the complication are unclear. neither data logs nor the product were returned for investigation. due to the lacking details and data logs and the product not being returned, the root cause of the high purge pressure could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
18014 | 1220648-2024-19776 | 20398560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20398560&pc=OZD | 02/16/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 02/16/2023 | 10/08/2024 |  |  | 09/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276500 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during implant, device had delivery issues due to the tortuosity of the patient's descending aorta, secondary to an aortic aneurysm. despite best efforts, the physician was unable to deliver the impella into proper position and elected to abort the procedure all together. the impella placement was aborted, there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. per clinical review, the physician was unable to deliver impella cp into proper position due to the tortuosity of the patient's descending aorta, secondary to an aortic aneurysm. the cause of the delivery issue was tortuosity and aortic aneurysm patient condition related. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography gu section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
18015 | 1220648-2024-19778 | 20398638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20398638&pc=OZD | 02/18/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 02/18/2023 | 10/08/2024 |  |  | 12/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024288290 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was no flow, despite normal motor current, normal pressure signal and normal aortic signal. it was noted that the left ventricle curve was decreased; a thrombus was suspected but not confirmed. the placement of the impella was aborted. there was no patient harm. | [5 empty] | The investigation into the low pump flow issue has been completed. the impella pump was returned for investigation. visual inspection of the returned pump revealed the presence of biomaterial on the impeller within the cannula. no other damage or abnormalities were found. it was determined that the cause of the low pump flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
18016 | 1220648-2024-19779 | 20398640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20398640&pc=OZD | 09/11/2024 | 11/19/2024 |  | 10/08/2024 |  |  | 09/11/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501505 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp pulled the impella back into the descending aorta. the impella cp was explanted and a new impella cp was placed. the patient remained stable and survived the surgery. | [5 empty] | The investigation for the reported positioning issue has been completed. clinical details mentioned that the pump was pulled back by the patient into the aorta. the pump was not returned. data logs were investigated and "position in aorta" alarms were observed confirming positioning issue. therefore, the root cause of the positioning issue was patient movement related. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
18017 | 1220648-2024-19780 | 20399255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20399255&pc=OZD | 09/14/2024 | 11/19/2024 |  | 10/08/2024 |  |  | 09/14/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544052 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp to support a patient post arrest for a cosmetic surgery. the pump was placed (as well as extracorporeal membrane oxygenation) and the patient was coding. the patient was transferred to the tertiary center after three days. the pump malfunctioned at transport when the purge sidearm broke and began to leak. the pump additionally stopped. the pump was explanted and the patient survived. | [5 empty] | The investigation for the pump stop has been completed. tpurge leak - pump is a casading event which led to the pump stop. the returned product was inspected and a crack along the yellow luer was observed. the purge leak was reproduced. the returned product had biomaterial in the purge gap which did not clear after soaking efforts. the data logs confirmed motor current high pump stop and showed a motor current spike before pump stop. the cause of the pump stop issue was due to purge leak with sidearm yellow luer damage leading to lack of purge. the cause of sidearm yellow luer damage was most likely occurring during use due to the high tensions in combination with disinfectants and certain drug ingredients (example oil, alcohol, lipids, etc.) the root cause of the pump stop was determined to be due to a damaged sidearm yellow luer. h.6 per investigation results, code 1250 fluid/blood leak should not have been selected on initial manufacturer device report number 1220648-2024-19780 as it is a casading event which led to the pump stop. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
18018 | 1220648-2024-19781 | 20400368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20400368&pc=PYX | 09/12/2024 | 11/19/2024 |  | 10/08/2024 |  |  | 09/12/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025531500 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon initial setup of the impella rp flex and insertion, purge fluid would not prime, but after straightening the tubing and disconnecting and reconnecting, prime was completed and impella rp flex was inserted without issue. later in the afternoon, there was a question about whether the purge fluid was leaking at the yellow hub connection so the purge cassette was changed and leaking seemed to have stopped and purge pressures and flow were maintained within range. the following day, the impella rp flex experienced a motor current high and pump stop. slightly leaking purge fluid at connector site was noted. attempts to restart at p6 were unsuccessful as well as attempts to restart at p2. the physician removed the impella rp flex at bedside. during this time, blood pressure dropped and pulsatility decreased and temporarily they had to restart levophed. as soon as the impella rp flex was removed, hemodynamics began to improve and return to baseline. the patient survived the explant. | [5 empty] | The investigation has been completed. the impella rp flex was returned for evaluation. upon visual inspection, no abnormalities were noted on the pump. biomaterial was present within internal wall of inflow cage. the pump was attempted to be run at p-9 and immediate pump stop occurred. additionally, high purge pressure alarms were triggered. the pump was soaked in tergazyme to remove biomaterial. a clean pump was attempted to be run again and immediate pump stop occurred again. a window was cut in the catheter just proximal to the motor and the purge line showed blood ingress. the purge sidearm was inspected further and a leak at the purge reservoir was noted. data logs from the field were reviewed and long term logs showed purge pressure step down with increases in purge flow at each step after a purge cassette change occured five hours into support. later in the case, motor current began to rise over a period of nine hours before motor current high pump stop occurred. the pump flow remained stable throughout the case. clinical details noted that a leak was seen at the yellow luer connection to the purge sidearm. the root cause of the pump stop was due to a cascading events as a result of a purge leak at the reservoir resulting in blood ingress into the purge line causing a rise in motor current to pump stop. the purge leak reported by the customer was investigated under the failure mode of "pump stop". b.7 revised a word that was misspelled on initial manufacturer device report number 1220648-2024-19781. g.4 revised pma/510(k)number as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-19781. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella rp flex with smart assist system with automated impella controller section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (eg, infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.¿
18019 | 1220648-2024-19782 | 20400609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20400609&pc=OZD | 09/07/2021 | 10/08/2024 |  | 10/08/2024 |  |  | 09/07/2021 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNK |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was presented to the medical facility for impella support. it was stated that the impella pump flows were not calculating properly when attempting to obtain cardiac output. during support, the thermodilution was reading a co of 2.5 l/min while impella flow was reading 3.5 l/min. there was no patient harm reported. | [5 empty] | D4: unknown serial number, lot number, udi, expiration date. d6a/d6b: date of implant/explant is unknown. h4: date of manufacturing is unknown. the investigation for the reported low pump flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. product and logs were not returned. the root cause of the low flows was likely due to aic overreporting. this report is being filed as part of a retrospective review of historical records.
18020 | 1220648-2024-19783 | 20400624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20400624&pc=OZD | 08/05/2021 | 10/08/2024 |  | 10/08/2024 |  |  | 08/05/2021 | 10/08/2024 |  |  | 12/31/2022 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022031319 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the impella was accidentally started outside of the body while the easy guide lumen was still in place. the pump was unplugged from the automated impella controller (aic), this pump was removed and a new impella was successfully placed. there was no patient harm reported, and this device was replaced. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the easy guide lumen was seen still in the pump. the root cause of the out of body pump stop was the failure to remove the easy guide lumen before turning on the pump. impella cp with smartassist for use during cardiogenic shock and high risk cpi instructions for use for the related event are as follows: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18021 | 1220648-2024-19784 | 20400661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20400661&pc=OZD | 08/08/2021 | 10/08/2024 |  | 10/08/2024 |  |  | 08/08/2021 | 10/08/2024 |  |  | 04/30/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079226 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for hemodynamic support. while on support, there was a pump stop. the controller was exchanged out and the pump is now running at p5 flows with appropriate waveforms. | [5 empty] | The investigation into the reported pump stop issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the pump stop was determined to be unrelated to the impella pump since replacing the aic resolved the issue. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18022 | 1220648-2024-19785 | 20400917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20400917&pc=OZD | 09/12/2024 | 11/21/2024 |  | 10/08/2024 |  |  | 09/12/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535074 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and the patient developed hematuria. plasma-free hemogloblin was administrated. the patient was stabilized. the impella was to be explanted due to the hemolysis. the patient survived the explant. | [5 empty] | The investigation for the hemolysis has been completed. product and data logs were not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | A.5 is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
18023 | 1220648-2024-19787 | 20401111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20401111&pc=OZD | 07/20/2022 | 10/13/2025 |  | 10/08/2024 |  |  | 07/20/2022 | 10/08/2024 |  |  | 04/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218266 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient required red blood cells and send out plasma-free hemoglobin (measure for hemolysis) due to the patient having pink urine. the patient also experienced access site bleeding and anticoagulation medical was withheld. the bleeding issue was reportedly resolved. it was reported that the patient survived the procedure. | The us complainant reported that at patient was placed on impella cp hemodynamic support. it was reported that the pt was going to get red blood cells and send out pfhg due to the pt having pink urine. the automated impella controller (aic) was intermittently alarming impella in aorta but the tee found that the position was normal with 3.3 cm from ao annulus. no repositioning was needed at that time. surgeon was concerned about the impella¿s sensor. additionally, it was noted that there was no systemic anticoagulation due to previous bleeding at the access site, now resolved. the returned pump was evaluated. the outflow cage and sensor were covered in a gel like material. as the gel dried it began to crystalize and when the pump was placed in water the pressure reading was artificially high. as the substance dissolved in the water the pressure reading returned to normal. the substance is most likely a build up of dextrose. the sensor operated without issue. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The product was returned and the optical sensor was found to be functional. the logs were reviewed and placement signal was confirmed to be abnormal with spiking. the root cause of the false impella in aorta alarm and inaccurate map was most likely patient condition as the patient condition was poor. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g1: corrected mfg email address.
18024 | 1220648-2024-19788 | 20401136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20401136&pc=OZD | 09/07/2021 | 10/08/2024 |  | 10/08/2024 |  |  | 09/07/2021 | 10/08/2024 |  |  | 02/28/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2022045998 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella cp device for mechanical circulatory support. during placement of the impella a positioning issue occurred when a wire crossed the aortic valve and then came back into the ascending aorta. additionally, as the team gained access into the left ventricle and tried to advance the impella it continued to get stuck and would not advance any further. while removing the impella for re-insertion the team did not notice that the sheath had folded in on itself and was severely kinked. a third attempt was performed, and the team noted that the pigtail was almost straight and did not want to attempt to re-insert the impella again. the physician chose to do a mini sternotomy and place an impella 5.5. there was no patient harm reported, and this device was replaced. | [5 empty] | The investigation for the reported delivery issues and cannula/pigtail detachment issues has been completed. the reported delivery issues were most likely due to the patient¿s condition. however, the root cause of the reported introducer sheath kink could not be determined as no product was provided for evaluation. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18025 | 1220648-2024-19786 | 20401601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20401601&pc=OZD | 08/22/2021 | 10/08/2024 |  | 10/08/2024 |  |  | 08/23/2021 | 10/08/2024 |  |  | 03/31/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022057183 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a cerebral hemorrhage during impella management, the following day the patients¿ vitals were reported to be stable. additionally, it was unclear whether the cause of the stroke was due to anticoagulated balance, as anti-coagulation is known to be a contributing factor of cerebral hemorrhage. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18026 | 1220648-2024-19789 | 20401635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20401635&pc=OZD | 09/08/2021 | 10/08/2024 |  | 10/08/2024 |  |  | 03/23/2022 | 10/08/2024 |  |  | 05/31/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083565 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced an ischemic stoke one week after the impella was removed from a successful procedure outcome. the relationship with the impella was indicated as ¿possibly related to the impella.¿ due to the facility restrictions no further information was able to be obtained. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18027 | 1220648-2024-19790 | 20401802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20401802&pc=OZD | 09/27/2021 | 10/08/2024 |  | 10/08/2024 |  |  | 09/27/2021 | 10/08/2024 |  |  | 06/30/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094275 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella was placed without any issues, then the patient coded and cardio-pulmonary resuscitation was initiated resulting in the impella moving into the aorta. the impella was removed with the intention to re-insert the impella, however, the patient stabilized. the physician made the decision not to re-insert the impella, then noted that the motor housing felt hot and he wanted to send the impella in for investigation. there was no patient harm reported, and the device was removed. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs were reviewed and show placement signal and flow drop to zero while motor current and motor speed remain. this suggests that the pump was removed while still running at a low p-level. a pump running outside the body will produce more heat. the situation was reproduced by pulling the pump out of a bucket of water while running at p-1. the motor housing heated as the pump ran in the air. the root cause of the warm motor housing was determined to be a use related issue. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18028 | 1220648-2024-19791 | 20401879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20401879&pc=OZD | 07/24/2022 | 03/10/2025 |  | 10/08/2024 |  |  | 07/24/2022 | 10/08/2024 |  |  | 02/29/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197052 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Increase in Pressure | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that when the patient sat up, the pump position was lost. the lines were pulled, prompting the team to readvance the pump to the left ventricle position. subsequently, signs of hemolysis appeared, necessitating troubleshooting and repositioning of the pump. the hemolysis resolved following these interventions. additionally, the team noted a rise in purge pressure, leading to an educational session on the use of tissue plasminogen activator in the bicarbonate purge. the administration of tissue plasminogen activator in the bicarbonate resulted in immediate positive outcomes, effectively preventing further increases in purge pressures and declines in purge flow. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hemolysis, high purge pressure and positioning issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. after reviewing the clinical information, it was determined that the cause of the positioning issues was patient movement. regarding the purge issues, data logs showed the "purge pressure high" alarm triggered after the purge pressure gradually increased. at the time of the alarms, the motor current was high, and the pump flow was saturated and non-pulsatile. it was determined that the cause of the high purge pressure was most likely ingested biomaterial since the patient likely had a clotting issue and tpa helped resolve the issue. this report is being filed as part of a retrospective review of historical records.
18029 | 1220648-2024-19792 | 20401944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20401944&pc=OZD | 07/25/2022 | 10/17/2025 |  | 10/08/2024 |  |  | 07/27/2022 | 10/08/2024 |  |  | 11/30/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023160816 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  |  | Asian | Injury | Yes | Yes | Restricted Flow rate | Difficult to Insert | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was observed that there was difficulty inserting the 14fr sheath due to the sheath having a crack in the sheath. when the pump was removed, a thrombus was found, which may have been part of the papillary muscle. the patient continued on support until the device was successfully weaned. no further information is available. | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | Instructions for use: impella cp with smart assist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves, ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." g1 reporting contact email revised on manufacturer device report 1220648-2024-19792 in accordance with updated procedures. h6 medical device problem code 1248, component code 4742, and investigation findings (b)(4) was inadvertently omitted on the initial manufacturer device report number 1220648-2024-19792. h6 investigation conclusions 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-19792. h10 instructions for use was inadvertently omitted on the initial manufacture device report number 1220648-2024-19792
18030 | 1220648-2024-19794 | 20402013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20402013&pc=OZD | 12/02/2021 | 10/08/2024 |  | 10/08/2024 |  |  | 12/02/2021 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Packaging Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the carton was crushed when the hospital received the shipment. the pump set itself were not damaged only their outer carton. action taken to resolve the issue are unknown. no patient is involved . | [5 empty] | The cause of the crushed packaging was damage during transport since the packaging is rated and tested for standard shipping requirements. device sn ,udi manufacturing, expiry dates and lot # information is unknown. this report is being filed as part of a retrospective review of historical records.
18031 | 1220648-2024-19793 | 20402280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20402280&pc=OZD | 11/19/2021 | 10/08/2024 |  | 10/08/2024 |  |  | 11/30/2021 | 10/08/2024 |  |  | 02/28/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUM |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022044870 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis, no further information regarding the treatment or outcome of the hemolysis was provided. the patient outcome was recovered, the patient remained on impella support and was successfully weaned | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the hemolysis has been completed. no clinical details provided were sufficient to determine a root cause. no clinical details, data logs, or devices were returned for this investigation and therefore the relationship between impella and the suspected hemolysis cannot be determined. the cause of the patient¿s injury was patient condition and the relation to impella was determined to be unknown/unrelated. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
18032 | 1220648-2024-19798 | 20402494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20402494&pc=OZD | 08/31/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 08/31/2022 | 10/08/2024 |  |  | 04/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216258 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a failed delivery into a calcified vasculature. insertion was aborted without any harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was most likely calcified vessels and large pannus. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18033 | 1220648-2024-19795 | 20402595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20402595&pc=OZD | 03/17/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 03/17/2022 | 10/08/2024 |  |  | 10/31/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2022150881 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Infusion or Flow Problem | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced ventricular fibrillation and ventricular tachycardia, the resolution for this issue is unknown. also, post intervention reveled some occlusion with the repositioning sheath, and a femoral bypass was performed to resolve this issue. additionally, the device exhibited flow saturation. the motor current was noted to be higher than expected for the p-level on the 9th day of impella cp support. the next day flows the motor current was higher, and flows were saturated. the decision was made to explant and proactively replace the pump as a result. it was reported that the patient survived explant. no further information is available. | [5 empty] | The investigation for the reported ventricular fibrillation (vf), ventricular tachycardia (vt), thrombus, saturation flow and has been completed. the product was returned, and the evaluation completed. per the results of the investigation "the data logs were returned and show a gradual increase in motor current in the last 3 days of the case. this increase in motor current correlates to an increase in calculated flow. the product was returned, and biomaterial was found built up under the impeller. a high motor current pump stop occurred when an attempt was made to start the pump. the root cause of the saturated flows was determined to be biomaterial". per the investigation for the vt/vf issue "any device contacting the heart wall in conjunction with a poor patient condition could result in vf/vt but since fluoroscopic imaging and/or sufficient clinical details were not provided, a relation between impella and the vf/vt is unable to be determined". as noted in the impella ifu: impella cp with smart assist system impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. this report is being filed as part of a retrospective review of historical records.
18034 | 1220648-2024-19796 | 20402687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20402687&pc=OZD | 08/30/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 08/30/2022 | 10/08/2024 |  |  | 04/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223271 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported lower than expected flows during use of the device. the pump was explanted due to persistent low flows. a second impella was placed with optimal flows achieved. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the logs confirm low pump flow and show suction alarms and impella position in ventricle alarms after the pump was started. the cause of the low pump flow was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
18035 | 1220648-2024-19797 | 20402690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20402690&pc=OZD | 12/28/2021 | 10/08/2024 |  | 10/08/2024 |  |  | 12/28/2021 | 10/08/2024 |  |  | 08/31/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022134812 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the impella was placed, the clinician noticed the motor current was flat, and the left ventricle tracing was fluctuating. the pump was repositioned with no improvement resulting in the decision to remove this impella and place a new impella. additionally, it was reported that upon removal of this impella there was a clot, thrombus, in both the inlet and outlet cages. there was no patient harm reported, and this device was successfully replaced. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. after reviewing the clinical information, it was determined that the cause of the low pump flow was most likely ingested biomaterial stemming from improper anticoagulation. this report is being filed as part of a retrospective review of historical records.
18036 | 1220648-2024-19801 | 20402831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20402831&pc=OZD | 01/10/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 01/10/2022 | 10/08/2024 |  |  | 10/31/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022155878 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, low flows were experienced, and the team confirmed proper positioning of the device and adequate volume status. the maximum flows achieved were reported to be 1.6 to 1.8 liters per minute. the flows remained low despite pressure level adjustments. the decision was made to remove and replace with a new impella. there was no patient harm reported. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18037 | 1220648-2024-19803 | 20402971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20402971&pc=OZD | 09/28/2021 | 10/08/2024 |  | 10/08/2024 |  |  | 01/12/2022 | 10/08/2024 |  |  | 05/31/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022092809 |  | 0048-0014-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, hemolysis and renal insufficiency occurred. the patient was weaned and the impella was removed. no other information was provided. | [5 empty] | The investigation for the reported hemolysis and renal failure issues has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. (there is no evidence linking the hemolysis and renal failure the impella and thus relation is unknown.) impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
18038 | 1220648-2024-19799 | 20403083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403083&pc=OZD | 01/03/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 01/03/2022 | 10/08/2024 |  |  | 10/31/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151789 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding which was managed with redressing the site with a better angle and surgiseal. it was also reported that the impella was positioned in the papillary muscle, the physician attempted to reposition the impella which resulted in the impella popping across the aortic valve. the physician made the decision to remove and replace the impella after multiple attempts to reposition the impella failed. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the initial investigation from 11-feb2022 states: the root cause of the positioning issue was use related. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18039 | 1220648-2024-19805 | 20403084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403084&pc=OZD | 10/04/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 10/04/2022 | 10/08/2024 |  |  | 03/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023205421 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an acute patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump stopped due to high motor current. the pump was exchanged with a new pump and the patient remained stable. | [5 empty] | The investigation for the reported pump stop issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause of the pump stop was bearing wear. this report is being filed as part of a retrospective review of historical records.
18040 | 1220648-2024-19800 | 20403115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403115&pc=OZD | 09/27/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 09/27/2022 | 10/08/2024 |  |  | 04/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247140 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Restricted Flow rate | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the pump was emitting purge alarms with no flow detected. the patient was also starting argatroban since patient came back heparin induced thrombocytopenia (hit) positive and bicarb in purge. the pump was removed and replaced. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has been returned for evaluation. the pump and purge cassette were returned and both operated without issue. clinical details were sufficient to determine the cause of the issue were user and patient condition related. this report is being filed as part of a retrospective review of historical records.
18041 | 1220648-2024-19802 | 20403130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403130&pc=OZD | 09/23/2024 | 11/08/2024 |  | 10/08/2024 |  |  | 09/23/2024 | 10/08/2024 |  |  |  |  | 10/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276202 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | Eu complainant reported the automated impella controller (aic) was being prepped for use and the purge disk holder of the aic broke and the aic couldn´t be used. the aic was replaced. no patient harm has been reported. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
18042 | 1220648-2024-19808 | 20403190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403190&pc=OZD | 09/18/2021 | 10/08/2024 |  | 10/08/2024 |  |  | 01/02/2022 | 10/08/2024 |  |  | 07/31/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105871 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis and an infection. it was noted that full cardiac function was recovered following the percutaneous coronary intervention. due to facility restrictions, no further information was able to be obtained. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation for the reported hemolysis and infection/sepsis issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions, including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18043 | 1220648-2024-19804 | 20403247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403247&pc=OZD | 01/21/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 01/21/2022 | 10/08/2024 |  |  | 10/31/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022155355 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, race and ethnicity unknown, noted as high-risk percutaneous coronary intervention was attempted to implant with an impella cp device for mechanical circulatory support. it was reported the impella cannula kinked, and maximum flows were only 1.5 regardless of pressure. the placement of the impella pump was aborted. no patient harm was reported. | [5 empty] | The investigation for the reported low flow issue has been completed. evidence of a cannula kink was found. the pump ran to specification in the lab. the root cause of the low flows was a cannula kink. this report is being filed as part of a retrospective review of historical records.
18044 | 1220648-2024-19806 | 20403310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403310&pc=OZD | 02/25/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 02/25/2022 | 10/08/2024 |  |  | 08/31/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022129748 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Failure to Sense | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. cerebral hemorrhage was confirmed during impella management. anti-coagulation is known to be a contributing factor to cerebral hemorrhage. there was also a placement signal and left ventricle waveforms suddenly stopped appearing. after 10 minutes the waveforms suddenly returned and the pump was operating normally. action taken to resolve the reported issues are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18045 | 1220648-2024-19814 | 20403325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403325&pc=OZD | 08/10/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 08/10/2022 | 10/08/2024 |  |  | 05/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218763 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a red alarm pump stop. the pump was explanted. the patient continued on support until the device was successfully weaned. no patient harm was reported due to the issue. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the device was not received for evaluation. the root cause of the pump stop was unable to be determined as no product or logs were returned for analysis. this report is being filed as part of a retrospective review of historical records.
18046 | 1220648-2024-19817 | 20403407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403407&pc=OZD | 11/14/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 11/14/2022 | 10/08/2024 |  |  | 03/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023204935 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after connecting the power cable, the step-by-step explanation was aborted and the start screen became active. the priming of the pump was also interrupted. successfully switched to a second pump and another aic console. no harm done to the patient. | [5 empty] | The investigation into the priming issue has been completed. the impella pump was returned for investigation. the root cause of the pump unable to start was use related because the red lumen was not removed prior to pump activation indicating a premature pump activation. this report is being filed as part of a retrospective review of historical records.
18047 | 1220648-2024-19807 | 20403446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403446&pc=OZD | 12/03/2021 | 04/03/2025 |  | 10/08/2024 |  |  | 01/27/2022 | 10/08/2024 |  |  | 02/28/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022064765 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon removal of the impella the patient experienced ventricular fibrillation requiring defibrillation. the patient¿s condition stabilized, and the patient had been successfully weaned. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18048 | 1220648-2024-19809 | 20403453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403453&pc=OZD | 10/08/2021 | 05/28/2025 |  | 10/08/2024 |  |  | 02/01/2022 | 10/08/2024 |  |  | 01/31/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022033025 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient developed hemolysis and thrombocytopenia. due to facility restrictions, no further information was able to be obtained. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the hemolysis and thrombocytopenia was not determined due to insufficient clinical details. this report is being filed as part of a retrospective review of historical records.
18049 | 1220648-2024-19810 | 20403464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403464&pc=OZD | 10/27/2021 | 09/23/2025 |  | 10/08/2024 |  |  | 02/22/2022 | 10/08/2024 |  |  | 07/31/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108779 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. approximately seven (7) days post impella explant the patient endured an ischemic cva. per discussion with medical affairs, the physician could not determine if the reported issue was related to the impella cp device, because open heart surgery was also performed. no further information was provided. | [5 empty] | The investigation for the reported stroke/cva issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the cva/stroke was determined to be unrelated to impella since it occurred 7 days after removal. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the cause of the cva/stroke was determined to be unrelated to impella since it occurred 7 days after removal.
18050 | 1220648-2024-19812 | 20403478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403478&pc=OZD | 10/27/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 10/27/2022 | 10/08/2024 |  |  | 08/31/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023249172 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella device setup, the medical staff was unable to prime the device. the impella aic was changed and the issue did not resolve. the impella implant was aborted. | [5 empty] | The investigation into the priming issue has been completed. the impella device was not returned for evaluation. the cause of the case start issue was not determined since neither the product nor the data logs were returned. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18051 | 1220648-2024-19820 | 20403550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403550&pc=OZD | 12/25/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 12/25/2022 | 10/08/2024 |  |  | 08/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023271084 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a pump stop during use of the device. the pump was restarted, but appropriate flows could not be achieved. the pump was removed, and site closed with no harm to the patient. | [5 empty] | The investigation for the reported pump stop has been completed. the product was returned and the pump stop was confirmed. the root cause of the pump stop was determined to be bearing wear. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
18052 | 1220648-2024-19816 | 20403569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403569&pc=OZD | 08/26/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 08/26/2022 | 10/08/2024 |  |  | 03/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213090 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had cp support for a protect iv study patient for just over an hour. the pump was explanted, and support deemed successful. eight months later, the case report form noted a femoral hematoma that was not treated at the access site. unspecified interventions were used. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the access site bleeding cannot be determined due to no product returned and lack of sufficient clinical information. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18053 | 1220648-2024-19811 | 20403570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403570&pc=OZD | 02/04/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 03/01/2022 | 10/08/2024 |  |  | 11/30/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157963 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that three days post explant, the patient experienced an ischemic stroke. no further information was provided. | [5 empty] | The investigation for the reported ischemic stroke has been completed. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the ischemic stroke could not be determined. strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: - patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation - timing of the stroke in relation to impella support (during or post-implant) - detailed description of the symptoms experienced by the patient - neurological examination findings and any focal deficits observed - diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic) - laboratory test results, including coagulation profiles and cardiac markers - any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications - response to any previous anticoagulation or antiplatelet therapies - evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology - as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
18054 | 1220648-2024-19818 | 20403600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403600&pc=OZD | 03/23/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 03/24/2022 | 10/08/2024 |  |  | 07/31/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108936 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced signs of hemolysis with changes in urine when the p levels were raised. lowering the p-levels resolved the issue. the patient was successfully weaned and the device was explanted. the patient survived the explant. | [5 empty] | A.5 ethnicity is unknown. the investigation for the reported hemolysis has been completed. the product nor the data logs were returned. after reviewing the clinical information, it was determined that the cause of the hemolysis was most likely patient condition related since lowering the p-level resolved the issue. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient condition including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
18055 | 1220648-2024-19813 | 20403633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403633&pc=OZD | 01/17/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 01/17/2023 | 10/08/2024 |  |  | 08/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269890 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported a delivery issue involving an impella pump. during attempted placement, it was noted that the impella could not advance into the left ventricle due to the motor housing getting stuck in the patient's calcified stent area. multiple attempts were made using various methods, however they each failed, and the implant was aborted. no further information is available. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. clinical details confirmed unsuccessful multiple attempted with different sheath sizes to advance the impella because of calcification built up. root cause was patient condition due to calcification issue. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18056 | 1220648-2024-19815 | 20403640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403640&pc=OZD | 03/22/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 03/22/2022 | 10/08/2024 |  |  | 02/28/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022047078 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support post cardiac surgery. after five days of support, the team made the choice to place the pump axillary instead of femoral due to cool foot. the patient survived the procedure. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the limb ischemia was most likely patient condition related. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18057 | 1220648-2024-19819 | 20403754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403754&pc=OZD | 01/17/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 01/17/2023 | 10/08/2024 |  |  | 06/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235021 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the device would not go past the valve, 0.18 wire also had slack and not right to aortic arch, the impella cannula doubled over itself. they already had radial access, so a snare was taken from left radial and snared pigtail, to straighten impella and pull it back through the oscar sheath. the device was inspected by physicians and there was a clear kink and now a week point above the motor, decision made to utilize a new impella cp device. no further information was provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely use related as the wire had slack making delivery difficult. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18058 | 1220648-2024-19822 | 20403850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403850&pc=OZD | 01/30/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 01/30/2023 | 10/08/2024 |  |  | 06/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247274 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. there was a failure to deliver the device due to patient's stenotic vessels. insertion was aborted with no reported harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. root cause of the issue was patient condition related because the patient had stenotic iliac artery bilaterally and was unable to pass the impella due to a narrowing in the iliac. this report is being filed as part of a retrospective review of historical records.
18059 | 1220648-2024-19821 | 20403970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403970&pc=OZD | 01/18/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 01/18/2023 | 10/08/2024 |  |  | 11/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that it was not possible to de-air the purge cassette and pump. there was no purge fluid at the outlet. they changed the pump and purge cassette, because they could not be sure the pump was properly primed. another impella cp device was placed without reported any issues. no further information was provided. | [5 empty] | The investigation for the reported issue has been completed. the pump and purge cassette were returned for analysis and were able to prime without issue. purge fluid could be seen exiting the purge gap. no data logs were returned for this investigation. the root cause of the priming issue was most likely use related as no issues were found with the returned product. this report is being filed as part of a retrospective review of historical records.
18060 | 1220648-2024-19823 | 20403978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20403978&pc=OZD | 02/01/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 02/01/2023 | 10/08/2024 |  |  | 07/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023246518 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an acute patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hemolysis. the support level was adjusted following discovery of the condition, and the pump was eventually removed due to the patient undergoing an unrelated operation. no subsequent intervention was required for the hemolysis. there was no reported harm to the patient. | [5 empty] | The investigation for the reported hemolysis has been completed. the cause of the hemolysis and its relation to the impella were not determined since the product, data logs, and sufficient clinical information was not provided. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
18061 | 1220648-2024-19824 | 20404015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20404015&pc=OZD | 01/30/2023 | 10/08/2024 |  | 10/08/2024 |  |  | 01/30/2023 | 10/08/2024 |  |  | 09/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278997 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, physician attempted to reposition the impella cp, the physician pulled the device out of the left ventricle and into the aorta. against recommendations, the surgeon tried to force the device across the valve wirelessly. during this attempt the surgeon causes the pump to kink on itself and in the ascending aorta. it was reported that the images were not clear. the device had to be removed, and the surgeon decided to place an intra-aortic balloon pump. no reported harm to the patient. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not completed. the cause of the positioning issue was wireless re-access. the impella device is not indicated for wireless re-access. this report is being filed as part of a retrospective review of historical records.
18062 | 1220648-2024-19825 | 20404134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20404134&pc=OZD | 07/18/2021 | 10/08/2024 |  | 10/08/2024 |  |  | 03/29/2022 | 10/08/2024 |  |  | 02/28/2023 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022070814 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding. no further information was provided as to the treatment for the reported event. the patient was successfully weaned and the device explanted. the patient survived the explant. | [5 empty] | A.5 is unknown. the investigation for the reported access site bleeding has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical details was provided. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18063 | 1220648-2024-19826 | 20404341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20404341&pc=OZD | 12/15/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 12/15/2022 | 10/08/2024 |  |  | 06/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248961 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, a hematoma developed at the access site. the pump was then explanted, and the patient survived the procedure. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the access site bleeding could not be determined. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18064 | 1220648-2024-19827 | 20404376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20404376&pc=OZD | 12/16/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 12/16/2022 | 10/08/2024 |  |  | 09/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273148 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during use, the physician received continuous alarms based on waveforms despite trying to reposition the device multiple times. the pump was then removed and replaced, and no patient harm was reported. | [5 empty] | The investigation for the reported positioning issues has been completed. the product was returned, and the positioning issue was confirmed. the root cause of the positioning issue was determined to be use related. this report is being filed as part of a retrospective review of historical records.
18065 | 1220648-2024-19828 | 20404378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20404378&pc=OZD | 12/17/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 12/17/2022 | 10/08/2024 |  |  | 09/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274361 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, a purge lumen alarm blocked alarm was presented, and tissue plasminogen activator (tpa) protocol was initiated. the impella pump was unable to be salvaged and a pump failure occurred. the device was explanted, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the pump stop could not be determined. as noted in the impella ifu: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18066 | 1220648-2024-19829 | 20404397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20404397&pc=OZD | 12/25/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 12/25/2022 | 10/08/2024 |  |  | 06/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023246506 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were delivery issues, unable to insert impella cp due to tortuos iliacs and aorta. removed cp and continued with intervention. patient remained stable; intervention successfully completed. no harm was done to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
18067 | 1220648-2024-19830 | 20404404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20404404&pc=OZD | 11/21/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 11/21/2022 | 10/08/2024 |  |  | 03/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204896 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella device experienced suction alarms with limited pump flow. echocardiogram was performed and revealed that the impella pigtail was near the posterior mitral valve. the device was explanted, and the procedural outcome was the patient survived explant. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause of the low pump flow was improper positioning of the device with the pigtail was near to posterior mitral valve sail. repositioning caused the pump to come out of the lv and the pump was presumably reinserted and support continued. this report is being filed as part of a retrospective review of historical records.
18068 | 1220648-2024-19831 | 20404410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20404410&pc=OZD | 11/23/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 11/23/2022 | 10/08/2024 |  |  | 08/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269553 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported a failure to deliver the impella cp device due to the patient's anatomy. insertion was aborted and an intra aortic balloon pump (iabp) was utilized for support. no patient harm or injury noted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the cause of the issue was patient condition related. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18069 | 1220648-2024-19832 | 20404436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20404436&pc=OZD | 11/29/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 11/29/2022 | 10/08/2024 |  |  | 09/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023259713 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after successful implantation of the impella, there was an error message, motor current too high. the pump was started manually, however the pump stopped in auto mode. the pump was removed and replaced with a new impella cp. it was reported that the replacement device ran without problems during the entire procedure. no further information is available. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was completed. the product was returned, and the red lumen was found fractured and wrapped around the impeller. the root cause of the pump being unable to start was red lumen issues due to improper technique. impella cp with smartassist for use during cardiogenic shock section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18070 | 1220648-2024-19833 | 20404440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20404440&pc=OZD | 11/30/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 11/30/2022 | 10/08/2024 |  |  | 09/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023264990 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were suction alarms that could not be resolved. the pump was explanted and exchanged successfully for another impella cp. it was reported that the replacement device ran without suction issues. no further information is available. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was completed. the pump was returned, and no abnormalities were found. the pump was run in a bucket of water and ran at p-9. the pump produced adequate flows. the root cause of the low flow was most likely patient condition as it was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
18071 | 1220648-2024-19834 | 20404476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20404476&pc=OZD | 11/30/2022 | 10/07/2025 |  | 10/08/2024 |  |  | 12/01/2022 | 10/08/2024 |  |  | 03/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207620 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella was placed with venous-arterial extracorporeal membrane oxygenation (va ecmo) support. the patient developed a gastrointestinal bleed which persisted after impella explant. it was noted that the patient was on anticoagulant therapy for a pulmonary embolism. a blood transfusion was required to manage this complication. weaning was successful and the procedural outcome was the patient survived. | [5 empty] | A.5 ethnicity is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18072 | 1220648-2024-19835 | 20404487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20404487&pc=OZD | 12/02/2021 | 10/08/2024 |  | 10/08/2024 |  |  | 12/22/2021 | 10/08/2024 |  |  | 06/30/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023271697 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump was inadvertently pulled back across the aortic valve during patient support. attempts to reposition the pump were unsuccessful and the pump was subsequently removed and replaced with another impella pump. the patient survived the procedure. | [5 empty] | The investigation into the positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issues was patient movement because while patient transported to cardiovascular intensive care unit (cvicu) and pressures dropped suddenly, upon arrival continuation of low pressure and decoupling of waveforms. this report is being filed as part of a retrospective review of historical records.
18073 | 1220648-2024-19836 | 20404494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20404494&pc=OZD | 12/12/2022 | 10/07/2025 |  | 10/08/2024 |  |  | 12/12/2022 | 10/08/2024 |  |  | 03/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204037 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was possible limb ischemia, however, no additional details were provided regarding the treatment or management of this condition. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: limb ischemia impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | Upon clarification from the field rep, there were no flow issues to the patient's limb and hemolysis was not confirmed. there is currently no allegation of a malfunction or serious injury associated with the implanted pump. a regulatory report is no longer necessary
18074 | 1220648-2024-19837 | 20404497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20404497&pc=OZD | 12/11/2022 | 10/08/2024 |  | 10/08/2024 |  |  | 12/11/2022 | 10/08/2024 |  |  | 08/31/2024 |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251524 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during the weaning process, the pump stopped unexpectedly. the pump was then explanted, and no patient harm reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported pump stop has been completed. an electrical resistance test revealed three open lines. the open lines were isolated to the red handle. upon inspection with the keyence the open lines were found to be on the catheter side inside the red handle. the cause of the pump stop was open electrical lines on the catheter side in the red handle. the cause of the open lines was most likely a manufacturing error. as noted in the impella ifu: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18075 | MW5160853 | 20407468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20407468&pc=OZD | 09/28/2024 | 10/03/2024 |  | 10/08/2024 |  |  |  | 10/08/2024 |  |  |  |  | 10/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP PUMP SET |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED INC. |  |  |  |  | 2025465529 |  |  |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Other Health Care Professional | [6 empty] | NA |  | Unknown |  |  |  | Other | No | No | [7 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | During a angiogram, the impella device was deployed. upon repositioning, it was reported that the catheter folded back on itself and there became a separation of the distal impella pigtail from the catheter.
18076 | 1220648-2024-19608 | 20380392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380392&pc=OZD | 10/05/2021 | 10/06/2024 |  | 10/07/2024 |  |  | 10/05/2021 | 10/07/2024 |  |  | 06/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094213 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implant, it was noted that the split plug in which the impella pump hangs from, was left on the device during production. the plug was removed without issue. no patient harm was reported, and the procedure was successfully completed with impella support. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the packaging issue was a brown split plug that impella hangs from during production was left on catheter. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18077 | 1220648-2024-19609 | 20380393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380393&pc=OZD | 10/07/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 10/07/2021 | 10/07/2024 |  |  | 04/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022082043 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a placement signal failure approximate 10 minutes into support. the pump was removed and replaced with no adverse impact to the patient. | [5 empty] | The investigation for the reported issue has been completed. optical spikes were seen on the returned data logs. analysis of the fringe pattern and 5s parameters in the lab was inconclusive. thus, the cause for the spikes is unknown. this report is being filed as part of a retrospective review of historical records.
18078 | 1220648-2024-19610 | 20382949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20382949&pc=PYX | 09/17/2024 | 10/07/2024 |  | 10/07/2024 |  |  | 09/17/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025514557 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | Migration or Expulsion of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5, ecmo and impella rp support ongoing for patient win pccs and suffering from right ventricle (rv) failure. the patient had the rp running for days when the patient had a bowel movement and was rolled in the bed for cleaning and the flows and mc dropped. the low flows were prompting the team to replace the pump but, instead discovered the pump had become mal-positioned. | [5 empty] | The investigation into the reported "positioning issues" has been completed since the original report was filed. product was returned for investigation. the device was visually inspected and found a catheter kink observed close to motor housing side. borescope view of cannula was obtained and significant biomaterial was found inside and blocked around and beneath the impeller at inflow cage. pump was ran with biomaterial and higher/saturated flows reproduced. the data logs were reviewed and motor current spike seen on (b)(6) around 4 am with drop in flows and no change in p-level. placement signal(ps) shifting upwards with rise post the motor current spike. additional motor current spike was seen relevant to ps shift resolving. flow issue did not resolve with repositioning of the pump. on 09/13 afternoon at 2 pm ps shift resolves after second mc spike just prior to increase in p-level to p-9. slight saturated flows seen later which is higher than normal flow range at p-9. rising motor current with slightly higher than normal flows found with decreasing ps. due to likely change of clot position at inflow cage post the mc spike, additional resistance to impeller/motor is observed indicating rising motor current with decrease in ps alongwith flow saturation. saturated and variable in flows with additional resistance to mc continued for next 87 hours until explant on (b)(6). cvp and pa pressures were found to be elevated during this time. the cause of the positioning issue was patient movement related with pump migrating during bedding change. the cause of the low pump flow issue was due to biomaterial ingestion.
18079 | 1220648-2024-19611 | 20383097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383097&pc=OZD | 06/18/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/19/2022 | 10/07/2024 |  |  | 04/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022075961 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during the first day of impella cp use, a tear was noted in the sterile sleeve while adjusting the position. the physician stated that they did not tape or twist the sleeve. the pump was explanted and replaced due to risk of contamination. there was no reported patient harm. | [5 empty] | The investigation for the reported product damage issue has been completed. the impella device was discarded by the customer, therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the sterile sleeve damage was most likely use related as the damage occurred during support. this report is being filed as part of a retrospective review of historical records.
18080 | 1220648-2024-19612 | 20383335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383335&pc=OZD | 09/17/2024 | 11/12/2024 |  | 10/07/2024 |  |  | 09/17/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531213 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 46-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella cp was explanted after 47minutes of support due to limb ischemia. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
18081 | 1220648-2024-19613 | 20383404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383404&pc=OZD | 06/20/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/20/2022 | 10/07/2024 |  |  | 09/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022132243 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Prime | Failure to Pump | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during implantation, the impella cp cassette was not able to de-air. after replacing the purge cassette, de-airing was possible, but the pump did not start. the medical team performed a full replacement of pump and cassette, no with further issues. there was no reported patient harm. | [5 empty] | The investigation for the reported priming and pump did not start issues has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump was visually inspected, and a fractured red lumen was observed. two purge cassettes were returned, and both were run through de-air successfully without issue. the cause of the priming issue was use related since it was unable to be reproduced and the returned product successfully ran de-air. the cause of the pump unable to start was a fractured red lumen removal issue. this report is being filed as part of a retrospective review of historical records.
18082 | 1220648-2024-19614 | 20383423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383423&pc=OZD | 09/13/2024 | 11/25/2024 |  | 10/07/2024 |  |  | 09/18/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1367614 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The patient was successfully weaned off the pump and automated impella controller (aic) support without incident which was one day earlier than planned. there were no complications related to the aic. per medical safety officer review use of the aic cannot conclusively be associated with the outcome. the patient did expire from a ventricular tachycardia episode post support and assessed as related to the impella pump which is reported in manufacturer device report 1220648-2024-19275. | The complainant reported the patient was on impella cp support and there were low purge pressure alarms. no visible leak was seen. the staff requested the automated impella controller (aic) be investigated as well for this issue. |  |  |  |  | The investigation into automated impella controller (aic) - purge issue ¿ other has been completed. the console log confirmed the reported failure mode given there was a purge pressure low - alarm. the failure mode was not reproduced throughout testing. the root cause of the low purge pressure during case could not be determined. b5 updated with patient outcome and mso statement on manufacturer device report 1220648-2024-19614. companion case manufacturer device report 1220648-2024-19275 reporting the impella pump and death. | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
18083 | 1220648-2024-19616 | 20383480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383480&pc=OZD | 03/29/2022 | 10/22/2024 |  | 10/07/2024 |  |  | 03/29/2022 | 10/07/2024 |  |  | 06/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022092827 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Female |  |  | Asian | Malfunction | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the complainant reported hemolysis during use of the device. actions taken to resolve are unknown. there was no additional information provided. | [5 empty] | The investigation for the reported issue has been completed. the cause of the hemolysis was not determined since product was not returned. unknown relation due to insufficient clinical details. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: :¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | Initial mfr was reported for the adverse event of hemolysis. additional information received on (b)(6) 2022 indicated there was no hemolysis that occurred to the patient while on impella cp support.
18084 | 1220648-2024-19615 | 20383481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383481&pc=OZD | 10/01/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 10/01/2021 | 10/07/2024 |  |  | 06/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094477 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a patient with known coronary artery disease was presented to the medical facility for impella placement. it was reported that the pump started, however flows were noted to be low and suction alarms were received. it was reported that the impella catheter was evaluated, and a kink was observed. the medical staff attempted repositioning, but the kink in the cannula remained until the impella pump was not positioned across the aortic valve. it was reported that the device was explanted. no kink was observed when the cannula was outside of the patient¿s body. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported issue has been completed. evidence of a kink was confirmed in the lab. thus, the root cause of the low pump flow was due to the cannula kink. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18085 | 1220648-2024-19617 | 20383507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383507&pc=OZD | 10/09/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 10/09/2021 | 10/07/2024 |  |  | 08/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121957 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. complainant reported that the patient was being placed on impella cp for hemodynamic support. while on support, access site bleeding was noted. hemoglobin drop was noted. device was explanted and patient's condition was stable. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the access site bleeding and its relation to the impella were not determined since neither the product nor sufficient clinical information was provided. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18086 | 1220648-2024-19618 | 20383523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383523&pc=OZD | 10/09/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 10/09/2021 | 10/07/2024 |  |  | 06/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095991 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. the complainant reported a placement signal failure during support. the pump was removed and replaced with no adverse impact to the patient. | [5 empty] | The investigation for the reported issue has been completed. logs were not provided. the product was return and was able to start in the lab and there were no issues with the placement signal. thus, the root cause could not be determined as it was not reproduced. this report is being filed as part of a retrospective review of historical records.
18087 | 1220648-2024-19619 | 20383536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383536&pc=OZD | 09/15/2024 | 11/11/2024 |  | 10/07/2024 |  |  | 09/15/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507897 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 placement was aborted due to the patient¿s anatomy. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella couldn't pass through the vasculature as per the clinical description provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
18088 | 1220648-2024-19620 | 20383540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383540&pc=OZD | 09/15/2024 | 11/07/2024 |  | 10/07/2024 |  |  | 09/15/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524113 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and they had oozing and a hematoma at the access site. pressure was held and the hematoma was expressed. a femstop was placed, the site was redressed and support continued. additionally, the patient had red urine with less than 30 ml/hr. the impella was turned down to p7 due to the suspected hemolysis. the impella was explanted the same day and upgraded to an impella 5.5. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the access site bleed and hemolysis has been completed. the impella was not returned for evaluation. data logs were reviewed which found no motor current (mc) spikes observed. no positioning issue alarms indicated to confirm the rc for reported issue, a few positioning alarms occurred that solved with quick troubleshooting. no suction alarms. no placement signal (ps) trend observed. the cause of the hemolysis cannot be determined since the complaint devise not returned for analyzing and insufficient clinical data provided. the cause of the access site bleeding - major cannot be determined since the complaint devise not returned for analyzing and insufficient clinical data provided.
18089 | 1220648-2024-19621 | 20383545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383545&pc=OZD | 09/18/2024 | 11/13/2024 |  | 10/07/2024 |  |  | 09/18/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557447 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the 71-year-old male patient had an impella 5.5 implant and the pump flipped in the mitral valve after it was successfully turned on. the doctor attempted to reposition the pump, but they were unsuccessful. the pump was explanted and then implanted into the left ventricle. the nurse tried to restart the pump 3 times and each time an impella stopped alarm appeared. the pump was explanted and replaced. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the positioning issue and the pump stop has been completed. the data logs were reviewed which show that they experience "impella in the aorta" alarms and retrograde flow and have trouble achieving flows for 5 minutes. the pump was removed, sat outside the body, and then a guidewire was used to reinsert it, and then there is a high motor current pump stop immediately and then the device was removed. these high motor current pump stop alarms are likely related to the positioning of the pump. the pump was returned and there were was a kink in the cannula seen but no other abnormalities. this kink is likely related to the positioning of the pump. the pump was able to run in a bucket for 10 minutes without reproducing a pump stop. the root cause of the pump did not start is most likely due to use. the root cause of the positioning issue is most likely due to use.
18090 | 1220648-2024-19622 | 20383547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383547&pc=OZD | 12/31/2021 | 03/27/2025 |  | 10/07/2024 |  |  | 12/31/2021 | 10/07/2024 |  |  | 06/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095003 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported when repositioning the sheath during placement, it was noted the patient was bleeding at the access site and a hematoma was noted. manual pressure was held to stop the bleeding. additionally, the physician stated there was a kink in the pump. the decision was made to abort the impella placement and replace the impella with an intra-aortic balloon pump (iabp). patient survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported product damage issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the product had several kinks in the cannula as well as a bent inflow cage. the root cause of the delivery issue was a cannula kink. no additional clinical information was provided regarding the patient's vasculature condition; therefore, the root cause of the cannula kink was unable to be determined. this report is being filed as part of a retrospective review of historical records.
18091 | 1220648-2024-19624 | 20383553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383553&pc=OZD | 09/18/2024 | 07/16/2025 |  | 10/07/2024 |  |  | 09/18/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025496269 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant reported the 73-year-old male patient was on impella cp support. during removal, the physician attempted to access the guidewire re-access port on impella cp and was unable to advance the wire. |  |  |  |  | The investigation for the removal issues has been completed. no product was returned. the reported failure was on physician unable to advance the guidewire when attempted using guidewire re-access port . no other information is available. the cause of the removal issue was not determined since product was not returned and due to insufficient clinical information. corrections to the initial mfr report: b1-selected adverse event. b2-selected death and added date of death. b5-mso review. added d6b. h1-type of reportable event changed to death. h6 impact code added 1802. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of removal issues has been completed. the pump was returned and inspected. a guidewire was able to be inserted through guidewire re-access port and removed without issue. the root cause of the removal issue was use related as no issue reproduced with the returned pump. d9 device returned to manufacturer date added. h6 updated with investigation results of the received device e4 should have been left blank on the initial manufacturer device report number 1220648-2024-19624 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-19624 in accordance with updated procedures.
18092 | 1220648-2024-19623 | 20383554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383554&pc=OZD | 09/18/2024 | 03/27/2025 |  | 10/07/2024 |  |  | 09/18/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1454040 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, the automated impella controller (aic) displayed a yellow controller failure message. | [5 empty] | The investigation for the controller error (yellow alarm) has been completed. the console log confirmed the reported failure mode given there was a pmd has detected a pump speed deviation - alarm issued triggered during the clinical procedure. there were multiple instances of the "impella stopped (motor current high)" and "impella stopped (retrograde flow)" alarms before the pump was unplugged from the console. real time logs confirmed that there are multiple instances of the motor current fluctuations, and which ultimately triggered pmd has detected a pump speed deviation alarm. the console was returned and was visually inspected. the internal circuitry of the console did not reveal any anomaly. the console was run with a known good pump and purge system with no observable anomaly was noted. the root cause for the controller error (yellow alarm) i.e. pmd has detected a pump speed deviation was not related to any problems with the automated impella controller (aic). no issues were found within the aic. b.7 information was added that was inadvertently omitted from the initial manufacturer device report number 1220648-2024-19623. d.1 revised brand name as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-19623. d.2 revised common device name since the initial manufacturer device report number 1220648-2024-19623 was submitted in accordance with updated procedures. d.4 revised model number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-19623. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-19623 was submitted in accordance with updated procedures. | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
18093 | 1220648-2024-19626 | 20383583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383583&pc=OZD | 01/03/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 01/03/2022 | 10/07/2024 |  |  | 08/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022123265 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 21 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported the impella was in place for eleven days when the pump stopped. the pump was able to be restarted, however on day 14 the pump stopped again and was unable to be restarted. the patients activated clotting time was reported to be 160 seconds or more. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation into the pump stop issue has been completed. the impella pump was returned for investigation. a visual inspection of the returned pump revealed large amount of biomaterial was seen. the cause of the pump stop was ingested biomaterial due to patient condition. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18094 | 1220648-2024-19627 | 20383592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383592&pc=OZD | 08/02/2024 | 10/07/2024 |  | 10/07/2024 |  |  | 09/16/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494594 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a 37-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, there was "impella position in aorta" alarms with pulsatile mc, normal flows, and echo-confirmed placement/positioning of the cp in the left ventricle (lv). support continued following the events. the investigation found that this was not a placement signal issue, it was a positioning issue where the impella in the aorta and the pump was repositioned successfully. the investigation did not find any issues with the placement signal, just that the pump became out of position this caused the alarms. | [5 empty] | The investigation into the reported 'positioning issues" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis. the evaluation revealed that the pump was used between (b)(6) 2024. data logs were provided up to 8/1/2024, 9:51 pm. the data logs show no signs of improper positioning or motor current spikes or optical signal issue during the case run till (b)(6) 2024 9:51 am. according to the clinical notes, patient was turn and pump was pulled back in the aorta, confirmed by echo, with an active impella positioned in the aorta alarm. the pump was successfully repositioned. the cause of the positioning issue was most likely patient movement since pump was found in aorta during patient turn.
18095 | 1220648-2024-19629 | 20383631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383631&pc=OZD | 09/21/2024 | 11/25/2024 |  | 10/07/2024 |  |  | 09/21/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533747 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support went for a heart transplant, and it was noted that when the patient woke up the left side was flaccid. computed tomography found large middle cerebral artery stroke. the patient required craniotomy to evacuate the clot. the patient survived the event. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
18096 | 1220648-2024-19630 | 20383642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383642&pc=OZD | 06/20/2022 | 05/14/2025 |  | 10/07/2024 |  |  | 09/26/2022 | 10/07/2024 |  |  | 01/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183310 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombocytopenia | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported a patient experienced lower limb ischemia and thrombocytopenia throughout the duration of impella support. there was no reported intervention performed for either event. no additional information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
18097 | 1220648-2024-19631 | 20383683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383683&pc=OZD | 01/11/2022 | 10/28/2025 |  | 10/07/2024 |  |  | 01/11/2022 | 10/07/2024 |  |  | 04/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022072528 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there were several periods of low purge flow and high purge pressure issues. reportedly, bivalirudin and dextrose 5% in water were only used in the purge due to gastrointestinal bleeding (gi). the clinician recommended removing bivalirudin and using heparin, the physician decided to use only dextrose 5% in water. reportedly an endoscopy was being done to address the gi bleeding, no results were noted. subsequently, physician added heparin to the purge due to possibility of pump stop. high-pressure issues were reportedly resolved with the addition of the heparin. the patient received three units of packed red blood cells throughout the case. in addition, it was reported there was an abrupt pump stop, the team was able to restart the pump however the motor current was reported to be high. the impella was replaced with a new impella pump and started without issues. patient survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The investigation of temporary pump stop, high purge pressure, and vascular damage - peripheral vessel has been completed. data logs showed the pump stopped upon the occurrence of the "impella stopped - motor current high" alarm as the motor current suddenly spiked above 1500. the pump was successfully restarted after several attempts. following the restart, the motor current was temporarily elevated, but no further pump stop alarms occurred until the pump was disconnected. visual inspection of the returned pump revealed a large purge gap with a significant amount biomaterial on and around the impeller. additionally, the pump was returned with the catheter completely severed. in conclusion, it was determined that the cause of the temporary pump stop was bearing wear after 15 days of use, which is beyond design life per design trace matrix. hpp / vascular damage: this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-19631 in accordance with updated procedures.
18098 | 1220648-2024-19625 | 20383706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383706&pc=OZD | 09/20/2024 | 11/06/2024 |  | 10/07/2024 |  |  | 09/20/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025518488 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the impella was in the left ventricle in good position and the automated impella controller (aic) was alarming impella in aorta and the waveforms were not ventricular. the placement signal was not resolved. the physician decided to remove the pump and place a new one. the patient remains on impella support. | [5 empty] | The investigation of placement signal issue has been completed. the data logs were evaluated and there were no problems found regarding the reported issues. d6b explant date added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-19625. e4 should have been left blank. g1 reporting contact fax number missing. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
18099 | 1220648-2024-19632 | 20383710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383710&pc=OZD | 01/11/2022 | 03/06/2025 |  | 10/07/2024 |  |  | 01/11/2022 | 10/07/2024 |  |  | 10/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022150764 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during placement of the impella the physician noted kinking at the cannula, right below the outlet chamber, when he would try to advance the catheter deeper into the left ventricle. the physician stated the patient had a small left ventricle and when trying to position the catheter it would kink. the physician removed the impella after multiple unsuccessful attempts to position the impella. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The product was returned and confirmed a kink in the cannula. imaging was returned from the field showing the cannula kinking as the physician attempted to reposition. the root cause of the low flow and suction was determined to be a kink in the cannula. the root cause of the cannula kink was most likely patient condition as the patient had a small left ventricle. this report is being filed as part of a retrospective review of historical records.
18100 | 1220648-2024-19628 | 20383734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383734&pc=OZD | 01/07/2022 | 03/27/2025 |  | 10/07/2024 |  |  | 01/07/2022 | 10/07/2024 |  |  | 10/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022137998 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the physician contacted the clinician due to not seeing any purge liquid. reportedly, the nurse accidentally started the pump prior to placement. the clinician advised the physician to place a new impella, which was implanted with no issues and the patient reported to be stable. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported pump did not start issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a visual inspection of the returned pump revealed that the red lumen was still present within the outflow cage, obstructing the impeller. the red lumen flag/label was also present. furthermore, purge residue was observed within the purge gap and on the motor housing upon return. the pump was connected to a purge cassette and purged normally. the root cause of the reported case start issue was the pump being started prematurely. this report is being filed as part of a retrospective review of historical records.
18101 | 1220648-2024-19633 | 20383756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383756&pc=OZD | 01/12/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 01/12/2022 | 10/07/2024 |  |  | 06/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095089 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. low flows of 1.7l/min were reported within an hour of insertion of impella. consequently, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was received for evaluation. the root cause of the low pump flow was the observed cannula kink; however, a root cause for the kink could not be determined. this report is being filed as part of a retrospective review of historical records.
18102 | 1220648-2024-19634 | 20383830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383830&pc=OZD | 02/07/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 02/07/2022 | 10/07/2024 |  |  | 07/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022103824 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient had severe tortuosity of the blood vessels and since a long sheath was used it was confirmed there was not a visible thrombus in the gap at the tip of the sheath. reportedly there was a suspected obstruction at the connection part of the medikit hemostatic valve, when the part was disconnected and blood was sucked from the sheath, a thrombus-like substance was removed. the medikit sheath was removed, and the wound was closed, the patients¿ vital signs did not change throughout the removal. it was confirmed that there was no problem with blood flow in the left lower limb. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the thrombus issue has been completed. the cause of the failure mode was misuse since an incorrect sheath was used resulting in the clot ingestion into the sheath resulting from improper technique by the end user. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves as noted in the impella ifu ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ as noted in the impella ifu "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18103 | 1220648-2024-19635 | 20383978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20383978&pc=OZD | 01/28/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 01/28/2022 | 10/07/2024 |  |  | 07/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022122109 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the case was delayed due to purge system failure with a cp and aic. the patient survived the procedure. | [5 empty] | The investigation into the purge leak - pump has been completed. the impella device was returned for evaluation. the cause of the purge leak - pump issue was not determined as the cassette was not returned. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18104 | 1220648-2024-19638 | 20384004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20384004&pc=OZD | 01/14/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 01/20/2022 | 10/07/2024 |  |  | 06/30/2022 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1491383 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the impella was placed, and flows were lower than normal, with suction alarms during the case and a flattened motor current. reportedly following the percutaneous coronary intervention, the physician scanned the arch under fluoroscopy and observed a significant kink in the cannula right below the outlet sitting in the aorta. the surgeon reportedly made unsuccessful attempts to straighten out the cannula, the decision was made to replace this impella with a new impella. there was no patient harm reported. | [5 empty] | The investigation into the low pump flow issue has been completed. the impella pump was returned for investigation. visual inspection of the returned pump confirmed a cannula kink near the motor housing. although the root cause of the low pump flow was determined to be the observed cannula kink, a specific cause for the observed cannula kink could not be determined. the root cause of the low pump flow was determined to be the observed cannula kink, resulting from improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
18105 | 1220648-2024-19636 | 20384032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20384032&pc=OZD | 01/17/2022 | 09/24/2025 |  | 10/07/2024 |  |  | 01/18/2022 | 10/07/2024 |  |  | 08/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022129084 |  | 0048-0047 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Unknown |  |  | Asian | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported during impella placement the cannula was kinking. reportedly the insertion of the impella went smoothly, however a fluoroscopy indicated kinking in the cannula proximal to the outlet. the impella flow dropped, and the patient¿s blood pressure dropped significantly from 100 to 50. the team tried to adjust the position with no improvement resulting in the team suspecting the issue was the patient¿s anatomy. the impella placement was aborted, and an intra-aortic balloon pump placed for support. the impella was replaced. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation of low pump flow and product damage (cannula kink) has been completed. data logs were not received. visual inspection of the returned pump revealed a minor cannula kink proximal to the outflow cage. no other damage or abnormalities were found. when the pump was run, its performance was within specification and no relevant alarms occurred. product damage (cannula kink) : after reviewing the clinical information, it was determined that the cause of the cannula kink was most likely patient anatomy. low pump flow: this complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-19636. g1 reporting contact email revised on manufacturer device report 1220648-2024-19636 in accordance with updated procedures.
18106 | 1220648-2024-19637 | 20384040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20384040&pc=OZD | 06/22/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/22/2022 | 10/07/2024 |  |  | 03/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023208571 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was presented for electrophysiology (ep)/ablation and implanted with an impella cp device for mechanical circulatory support. while on support, there was bleeding coming from the blue butterfly of the guidewire reposition sheath of the impella cp. upon arrival to the icu, the dressing to the right femoral artery was saturated with blood and a large hematoma had formed. after multiple attempts to reposition the repositioning sheath and light pressure with a fem-stop, the physician elected to explant the device. it was reported that the patient was hemodynamically stable after impella cp explant. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that no damage could be seen. the guidewire repo unit was tested to specification and passed. the root cause of the access site bleed was most likely use related as leaks can occur in the area of the blue butterfly when the tip of the butterfly is inserted into the access site. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18107 | 1220648-2024-19641 | 20384275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20384275&pc=OZD | 06/22/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/22/2022 | 10/07/2024 |  |  | 03/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206107 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | MEDTRONIC COREVALVE | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the pump would not cross through the medtronic corevalve in the aorta. the physician tried multiple wires and many attempts with the pump. the pigtail would cross the aorta, but not the inlet cage. after 3 attempts, the pump was aborted, and an intra-aortic balloon pump device was used. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was an interaction with another device. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18108 | 1220648-2024-19639 | 20384306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20384306&pc=OZD | 06/22/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/22/2022 | 10/07/2024 |  |  | 12/31/2022 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022024103 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 9 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella stopped suddenly on the 9th day of support. the pump stop was accompanied with the "impella stopped - high motor current" alarm. the clinicians made multiple attempts to restart the pump, but their efforts were unsuccessful. the decision was made to explant the device after successfully weaning the patient. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that data logs showed the pump stopped immediately upon the sudden occurrence of the "impella stopped - motor current high" alarm; the alarm was accompanied with a motor spike to 1500. the pump was unable to be restarted. there were no abnormal pump performance metrics prior to the pump stop. visual inspection of the returned pump revealed a large impeller purge gap. also, biomaterial was found on and around the impeller. no other damage or abnormalities were found. in conclusion, it was determined that the root cause of the pump stop was bearing wear after over 8 days of use, which is beyond design life. impella cp with smart assist system instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18109 | 1220648-2024-19640 | 20384319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20384319&pc=OZD | 09/18/2024 | 11/19/2024 |  | 10/07/2024 |  |  | 09/18/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557447 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump delivered to support a 71-year-old male patient with coronary artery bypass graft surgery (cabg) and mitral valve repair (mvr). the pump was placed but, positioning was poor in location within the left ventricle (lv). the pump was pulled back out and repositioned. the pump was then again placed within the lv and stopped. the pump stop lead to pump explant and replacement. | [5 empty] | The investigation for the pump stop and the positioning issue has been completed. no product was returned. data logs were reviewed and no particular trend was seen with pump stop failure. the cause of the pump stop issue was not determined since product was not returned and due to insufficient evidence per logs/clinical. the cause of the positioning issue was most likely use related. h6 impact code 4628. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
18110 | 1220648-2024-19642 | 20384464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20384464&pc=OZD | 06/23/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/23/2022 | 10/07/2024 |  |  | 03/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203670 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during impella insertion, the physician stated that it felt as though the impella was sticking on the guidewire. it was noted that several best practices were not followed during insertion: the insertion angle of the pump was slightly steeper than recommended (45 degrees), excessive force was used, and the repositioning unit was advanced into the introducer. ultimately, the decision was to remove the pump and guidewire and insert a new impella cp over a new guidewire. there was no report of patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. no product was returned. after reviewing the clinical information, it was determined that the root cause of the delivery issue was most likely use related since several best practices were not followed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action. this report is being filed as part of a retrospective review of historical records.
18111 | 1220648-2024-19643 | 20384627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20384627&pc=OZD | 09/18/2024 | 11/19/2024 |  | 10/07/2024 |  |  | 09/18/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025550529 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 54-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, the patient complained of pain at the insertion site. support continued. additionally, there were ¿impella in aorta¿ alarms. echo confirmed that the impella was out of the left ventricle. the surgeon attempted to recross without success. the pump was replaced. | [5 empty] | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
18112 | 1220648-2024-19494 | 20384753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20384753&pc=OZD | 07/17/2023 | 10/07/2024 |  | 10/07/2024 |  |  | 07/17/2023 | 10/07/2024 |  |  | 03/31/2025 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343892 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported during placement of the impella the physician was unable to pass the impella past the bifurcation, due to severe abdominal aortic stenosis. the impella placement was aborted, and the procedure plan was to re-attempt placement the following day. there was no patient harm reported, and this device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. per clinical description, the root cause of delivery issue was most likely due to patient condition. the cause of the delivery issue was patient condition related to diseased small vessels. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluation" this report is being filed as part of a retrospective review of historical records.
18113 | 1220648-2024-19644 | 20384824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20384824&pc=OZD | 06/24/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/24/2022 | 10/07/2024 |  |  | 10/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151460 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp had low flows. the decision was made to remove the pump and reposition. obstruction in cannula was noted when removed. the impella was replaced and proceeded with the procedure. | [5 empty] | The investigation for the reported low pump flow issue has been completed. from the data log, suction alarm and low flow were confirmed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18114 | 1220648-2024-19645 | 20384952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20384952&pc=OZD | 06/24/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/24/2022 | 10/07/2024 |  |  | 09/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142335 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on the 17th day of impella cp support, a high motor current pump stop occurred and the decision was made to exchange the pump. | [5 empty] | The investigation for the reported pump stop issue has been completed. the root cause of the pump stop was unable to be determined as no product was returned for evaluation. the pump stop occurred on the 17th day of support which is beyond the 8 day design life per design trace matrix. impella cp with smart assist system instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18115 | 1220648-2024-19646 | 20385175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385175&pc=OZD | 09/19/2024 | 11/18/2024 |  | 10/07/2024 |  |  | 09/19/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025496346 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and the patient had limb ischemia. the patient's left leg was cold and pulses could not be dopplered. an external femoral artery to femoral artery bypass was performed. support continued. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
18116 | 1220648-2024-19647 | 20385376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385376&pc=OZD | 06/24/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/24/2022 | 10/07/2024 |  |  | 03/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211703 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. during implantation, there were delivery issues with the impella cp. due to patient anatomy, the device could not be placed in final positioning and was removed. a second device was attempted to be placed by the femoral artery. there was no report of patient injury. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the issue was patient condition related. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18117 | 1220648-2024-19648 | 20385447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385447&pc=OZD | 05/14/2023 | 05/30/2025 |  | 10/07/2024 |  |  | 05/14/2023 | 10/07/2024 |  |  | 09/30/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-EU |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2023261620 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient, demographics unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the inlet was close to papillary muscle. the physician unsuccessfully attempted to reposition the device. the physician made the decision to remove this impella and replace it with a new impella. there was no patient harm reported, and this device was replaced. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning issues issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the light test was performed. light was seen coming out of the optical fiber at a kink at the 100 cm mark on the catheter. the cause of the issue was a damaged optical fiber line as a result of a kink in the catheter. this report is being filed as part of a retrospective review of historical records.
18118 | 1220648-2024-19649 | 20385495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385495&pc=OZD | 09/19/2024 | 11/20/2024 |  | 10/07/2024 |  |  | 09/19/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025543812 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported left ventricle thrombus was noted on trans-thoracic echocardiography with patient on impella 5.5 support. heparin was given. the patient survived the event. | [5 empty] | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
18119 | 1220648-2024-19650 | 20385638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385638&pc=OZD | 08/08/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 08/08/2021 | 10/07/2024 |  |  | 06/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100364 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported the impella was successfully placed and approximately twenty minutes after start-up the pump stopped with a flat placement signal. the decision was made to remove this pump and replace the pump with a new impella. there was no patient harm reported. | [5 empty] | The investigation for the reported saturated flow issues has been completed. the impella device has been returned for evaluation. a review of the data logs revealed that, upon pump activation, there were rapid increases in motor current (little to no pulsatility), pump flow, and purge pressure. placement signal metrics were inconsistent but were generally increasing. purge heparin levels were at 0 for the entirety of support. visual inspection of the returned pump revealed considerable biomaterial occluding the purge gap. no other abnormalities were found, and the pump appeared to be cleaned prior to return. when the pump was connected to the aic, there were no relevant alarms that occurred. in conclusion, it was determined that the cause of the pump flow issues was ingested biomaterial most likely due to the lack of heparin in the purge solution. this report is being filed as part of a retrospective review of historical records.
18120 | 1220648-2024-19651 | 20385655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385655&pc=OZD | 09/11/2024 | 11/20/2024 |  | 10/07/2024 |  |  | 09/19/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532278 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump support ongoing for a patient admitted in for a high risk percutaneous coronary intervention. the pump was placed and allowed for support and explanted after 46 hours. abiomed¿s long-term outcome and quality indicator (loqi) impella registry database later noted the team had documented an access site bleed with a definite relationship to the pump. the team used suturing. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was determined to be use issue because the additional suture placed at the bleeding site ceased the bleeding and achieved hemostasis. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18121 | 1220648-2024-19652 | 20385679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385679&pc=OZD | 08/16/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 08/19/2021 | 10/07/2024 |  |  | 03/31/2022 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1472488 |  | 0048-0032 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient suffered one or more strokes post impella explant. actions taken are unknown. no further information was provided. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was most likely due to patient condition. as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18122 | 1220648-2024-19653 | 20385718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385718&pc=OZD | 09/09/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 09/09/2021 | 10/07/2024 |  |  | 07/31/2022 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1497440 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Infusion or Flow Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported flow saturations, thrombus at the tip of the placement sheath, and suction with this pump during support. the decision was made to replace the pump. | [5 empty] | The investigation for the reported issues has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. a root cause could not be determined through this evaluation as no product or data logs were provided. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). this report is being filed as part of a retrospective review of historical records.
18123 | 1220648-2024-19654 | 20385725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385725&pc=OZD | 04/29/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 02/08/2022 | 10/07/2024 |  |  | 01/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022043019 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient¿s body movements were so severe the catheter was removed from the left ventricle and was unable to be repositioned. there was no patient harm reported, and this device was replaced with a new impella cp. | [5 empty] | The impella device was not received from the customer, therefore, evaluation of the device was not possible. the investigation into the positioning issue has been completed. the cause of the positioning issue was use related as patient movement caused the device to come out of position. this report is being filed as part of a retrospective review of historical records.
18124 | 1220648-2024-19655 | 20385751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385751&pc=OZD | 09/10/2024 | 11/18/2024 |  | 10/07/2024 |  |  | 09/10/2024 | 10/07/2024 |  |  |  |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025504236 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support experienced an increase in drainage to the jp drain and oozing from the impella access site. sandbag was placed to the right side of the chest. the plan is to transfuse one unit of platelets and fresh frozen plasma. ten days later, the automated impella controller alarmed "high purge pressure" followed by "purge system blocked" alarm. anticoagulation had been held for four days due to pulmonary hemorrhage and hemoptysis. the decision was made to exchange the pump. the patient survived the explant. | [5 empty] | The investigation for the high purge pressure and access site bleed has been completed. there were no data logs returned. the pump was returned and there was biomaterial noted in the purge gap as well as wrapped around the impeller. the pump was soaked and the biomaterial was removed but high purge pressure alarms were still able to be reproduced. near the motor housing, there was a kink in the catheter. additionally, the clinical information stated that there was original trouble trying to make it past the graft anastomosis. the catheter was opened and there was a kink and blood in the purge line. when the pump was cut near the kink, the blockage could be isolated to the motor housing as purge fluid could be expelled. the root cause of the high purge pressure is most likely due to a kinked purge line that lead to the blood ingress. there were no damages seen on the pump repositioning unit and no abnormalities were found. due to limited clinical details provided, the root cause of the access site bleeding - major could not be determined. a.1 revised as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-19655. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
18125 | 1220648-2024-19656 | 20385765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385765&pc=OZD | 08/08/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 02/15/2022 | 10/07/2024 |  |  | 06/30/2022 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1503409 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 91 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. reportedly ten days after the impella had been removed it was reported the patient developed an infection at the insertion site. actions taken to resolve the issue are unknown. | [5 empty] | The impella device was not received from the customer, therefore, evaluation of the device was not possible. the investigation into the infection has been completed. the cause of the infection was most likely patient condition since it occurred 10 days after impella removal and the impella device was confirmed to have been sterilized under validated conditions. this report is being filed as part of a retrospective review of historical records.
18126 | 1220648-2024-19657 | 20385792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385792&pc=OZD | 03/10/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 03/10/2022 | 10/07/2024 |  |  | 12/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176555 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the cp failed to purge/prime correctly. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation into the reported purge leak - pump has been completed. the impella device was returned for evaluation. there was no issue found during the case. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
18127 | 1220648-2024-19658 | 20385812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385812&pc=OZD | 03/10/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 03/10/2022 | 10/07/2024 |  |  | 11/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163642 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a thrombus was observed on the impella device upon removal of the device. actions taken are unknown. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of thrombus. the device was returned for evaluation and confirmed thrombus in the cannula near the impeller. the root cause of the thrombus was determined to be biomaterial. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18128 | 1220648-2024-19659 | 20385826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385826&pc=OZD | 03/08/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 03/08/2022 | 10/07/2024 |  |  | 08/31/2022 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1496671 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the patient developed ischemia during impella support. actions taken to resolve are unknown. no further information was provided. | [5 empty] | The investigation for the reported limb ischemia issues has been completed. the impella device has not been returned for evaluation.the cause of the limb ischemia was not determined due to insufficient clinical details. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
18129 | 1220648-2024-19660 | 20385933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385933&pc=OZD | 03/14/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 03/14/2022 | 10/07/2024 |  |  | 09/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022143035 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that continuous suction alarms occurred during use of the device. the pump was removed and a large clot was noted in the inlet and outlet. the pump was replaced without any harm to the patient. | [5 empty] | The investigation for the reported thrombus issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the issue was biomaterial. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18130 | 1220648-2024-19661 | 20385959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20385959&pc=OZD | 06/26/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/26/2022 | 10/07/2024 |  |  | 06/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105987 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp was leaking near the red hub. a purge bypass was completed successfully and there were no reported issues after. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation for the reported purge leak issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the purge reservoir and filter were returned and inspected. leak reproduction testing exhibited a leak at the yellow luer connection. a crack in the yellow luer was visualized under a microscope. the yellow luer crack was most likely caused by bicarbonate usage and lead to the purge leak reported by the field. the root cause of the purge leak was sidearm yellow luer damage. impella cp with smart assist system instructions for use for the related event are as follows: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." this report is being filed as part of a retrospective review of historical records.
18131 | 1220648-2024-19662 | 20386033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20386033&pc=OZD | 03/20/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 03/20/2022 | 10/07/2024 |  |  | 12/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177758 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after inserting the repositioning sheath a delivery issue where the catheter got caught behind the papillary muscle occurred. reportedly during attempts to reposition the device the physician had difficulty passing the aortic valve, resulting in the catheter getting kinked and damage to the sensor. this impella was removed and replace with a new impella pump. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18132 | 1220648-2024-19663 | 20386076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20386076&pc=OZD | 06/27/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/27/2022 | 10/07/2024 |  |  | 01/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172164 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on the third day of impella support high purge pressure was reported. the next day a high motor current pump stop occurred. despite several attempts, the pump would no longer function. as a last attempt, the pump was manually flushed, but this effort was unsuccessful. the pump was explanted as a result. there was no allegation of patient harm. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that biomaterial was found in the purge gap. the catheter was cut and blood was confirmed in the purge line. the root cause of the high purge pressure was determined to be biomaterial. impella cp with smart assist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18133 | 1220648-2024-19664 | 20387482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20387482&pc=OZD | 06/27/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/27/2022 | 10/07/2024 |  |  | 06/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022103342 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Unknown |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the clinical team reported that "microbubbles" were observed in the left ventricle via echocardiography (echo) imaging during impella support. the bubbles appeared and disappeared intermittently. the clinicians were unable to identify the source of the bubbles, however they ruled out the presence of a ventricular septal defect or patent foramen ovale in the heart. additionally, it was reported that the impella introducer had been flushed. the patient was on extracorporeal membrane oxygenation support at the time. as a result of the impella was explanted. | [5 empty] | The investigation for the reported embolism issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that fluoroscopic imaging sent from the field confirmed the appearance of white dots or "bubbles" on echo. data logs showed continuous "impella flow low" and "suction" alarms upon pump activation; the motor current was pulsatile, however it gradually decreased over the next 30 minutes. the suction conditions eventually resolved. there were no issues with the purge metrics. no damage or abnormalities were found on the returned pump or the purge cassette. when the impella was run with the purge cassette, all pump performance and purge metrics were within specification. therefore, it was determined the root cause of the embolism was most likely patient condition related since the pump performed within specification and there were no purge issues. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18134 | 1220648-2024-19665 | 20388000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20388000&pc=OZD | 06/28/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/29/2022 | 10/07/2024 |  |  | 03/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206500 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during setup, there was a priming issue with the impella cp pump. the clinical team reported that there was no purge fluid observed to be exiting out of the purge gap during case setup. the decision was made to replace the pump with another, which primed with no issues. | [5 empty] | The investigation for the reported priming issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the priming issue was not determined since the product and the data logs were not returned. this report is being filed as part of a retrospective review of historical records.
18135 | 1220648-2024-19672 | 20388389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20388389&pc=OZD | 08/10/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 08/10/2021 | 10/07/2024 |  |  | 05/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022091821 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported a few minutes after completion of balloon aortic valvuloplasty, the impella flows, and motor current dropped with repeated suction alarms. the decision was made to remove this impella due to the low pump flow issue. there was no patient harm reported. | [5 empty] | The investigation into the low pump flow issue has been completed. the impella pump was returned for investigation. the visual inspection revealed a large clot was observed in the cannula wrapped around the impeller. the cause of the low pump flow was ingested biomaterial due to patient condition. this report is being filed as part of a retrospective review of historical records.
18136 | 1220648-2024-19667 | 20388850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20388850&pc=OZD | 06/29/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/29/2022 | 10/07/2024 |  |  | 03/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209050 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced bleeding at the impella access site after the reposition sheath was placed. manual pressure was held at the impella access site, and sutures were deployed. an emergency computed topography scan confirmed a subcutaneous hematoma and there was no retroperitoneal hematoma. the patient's activated clotting time at the time of the bleed was 270 seconds. as a result of the bleeding, the impella cp was explanted. the patient was successfully weaned and explanted. | [5 empty] | The investigation for the reported bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the bleeding was act values (270 seconds) above the recommended 180 second maximum. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18137 | 1220648-2024-19666 | 20389117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20389117&pc=OZD | 06/28/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 03/23/2023 | 10/07/2024 |  |  | 03/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209346 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. one day post impella cp explant, the patient experienced an embolic stroke. there was no information provided on patient outcome or any treatments provided due to the embolic stroke. it was unknown if the embolic stroke was attributed to the impella cp. | [5 empty] | The investigation for the reported embolic stroke issue has been completed. the impella device was not received from the customer, therefore, an evaluation of the device was not possible. the field lt log was reviewed and no spikes in motor current were observed that would indicate a clot passing through the impella pump. the root cause of the embolic stroke was not determined due to insufficient clinical details. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18138 | 1220648-2024-19680 | 20389389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20389389&pc=OZD | 10/05/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 10/05/2021 | 10/07/2024 |  |  | 03/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022061520 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported the impella pump stopped working without prior indication. the device was explanted, and the procedural outcome is unknown. | [5 empty] | The investigation for the reported issue has been completed. the data logs were returned and confirm a high motor current pump stop at the end of the case. the product was returned and a large purge gap could be seen indicating bearing wear. the root cause of the pump stop was determined to be bearing wear due to overuse. the pump stop occurred on day twelve which is four days over the eight day design life of the impella cp optical per design trace matrix 0046-9910. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18139 | 1220648-2024-19669 | 20389476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20389476&pc=OZD | 03/27/2022 | 02/13/2025 |  | 10/07/2024 |  |  | 03/27/2022 | 10/07/2024 |  |  | 11/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168942 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella stopped suddenly after 21 days of support, but was able to be restarted. after pump restarted, the motor current was elevated and showed frequent spikes until the planned impella removal. the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | Data logs showed the pump stopped immediately upon the sudden occurrence of the "impella stopped - motor current high" alarm. the pump stop was brief and the impella was able to restart. following the pump pause, the motor current was elevated and erratic and the flow was saturated for the remainder of support. visual inspection of the returned pump revealed a significantly large impeller purge gap. no other external damage or abnormalities were found. therefore, it was determined that the cause of the pump stop was bearing wear after 21 days of support, which is beyond design life.
18140 | 1220648-2024-19671 | 20390067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20390067&pc=OZD | 03/18/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 03/18/2022 | 10/07/2024 |  |  | 05/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2022086466 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 49 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that when connecting the impella cp optical with aic, the message "connect gray cable" appeared. the team attempted to troubleshoot by aic replacement, but found that they needed to replace the pump. the second pump was successfully placed. no patient harm or injury reported. | [5 empty] | The device was received for evaluation. the returned pump was connected to a console and pump sn (b)(6) initiated without issue. case start was prompted to begin. the pump performed as indicated. the root cause of the pump not detected could not be determined.
18141 | 1220648-2024-19684 | 20390469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20390469&pc=OZD | 10/25/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 10/25/2021 | 10/07/2024 |  |  | 04/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022075432 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the complainant reported high purge pressure. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation into the high purge pressure has been completed. the impella device was not returned for evaluation. the root cause of the high purge pressure was unable to be determined as no data or product was returned for analysis. this report is being filed as part of a retrospective review of historical records.
18142 | 1220648-2024-19674 | 20390618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20390618&pc=OZD | 08/16/2021 | 02/27/2025 |  | 10/07/2024 |  |  | 08/16/2021 | 10/07/2024 |  |  | 11/30/2022 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021014847 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support,the pump stop after 10 days and unable to restart. the patient was escalated to veno-arterial extracorporeal membrane oxygenation (va ecmo) with plans was to escalate to an impella 5.5 but waiting for sepsis to clear and central lines exchanged. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the device was found to have an enlarged purge gap. this indicates the bearing had worn out and stopped the pump. the cause of the pump stop was bearing wear. the pump stop occurred beyond the 8 day impella cp indication for use per design trace matrix 0046-9910. this report is being filed as part of a retrospective review of historical records.
18143 | 1220648-2024-19668 | 20390619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20390619&pc=OZD | 06/29/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/29/2022 | 10/07/2024 |  |  | 01/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023175785 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella was slipping out the ventricle and being unable to re-cross the valve. a new impella device was used with no harm to the patient. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the root cause of the positioning issue was determined to be use related. this report is being filed as part of a retrospective review of historical records.
18144 | 1220648-2024-19687 | 20390996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20390996&pc=OZD | 06/13/2023 | 10/07/2024 |  | 10/07/2024 |  |  | 06/13/2023 | 10/07/2024 |  |  | 12/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316008 |  | 1000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was going to be implanted with an impella cp device for mechanical circulatory support. it was reported that during implantation, the physician had trouble getting past the distal descending aorta. multiple attempts to advance were made with no success and the procedure was aborted. there was no reported harm to the patient as a result of this incident. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was most likely patient condition related. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18145 | 1220648-2024-19677 | 20391056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391056&pc=OZD | 10/05/2021 | 10/17/2025 |  | 10/07/2024 |  |  | 10/05/2021 | 10/07/2024 |  |  | 07/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108810 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that when the patient¿s cardiac function was stable, the impella was explanted. during removal, it was reported that thrombus was observed near the pump discharge. weaning was successful, the procedural outcome was the patient survived surgery. | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that when the patient¿s cardiac function was stable, the impella was explanted. during removal, it was reported that thrombus was observed near the pump discharge. weaning was successful, the procedural outcome was the patient survived surgery. |  |  |  |  | The investigation for the reported issue has been completed. no product was returned for evaluation and no data logs were provided for review; therefore, an evaluation of the device was not possible. it was reported that biomaterial was observed in the outflow upon explant of the pump. the root cause of the biomaterial was unable to be determined. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | B2 selections on initial manufacturer device report number were erroneously selected. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h1 type of report was erroneously selected on initial manufacturer device report number. h6 health effect - impact code 2199 was erroneously entered on initial manufacturer device report number.
18146 | 1220648-2024-19695 | 20391118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391118&pc=OZD | 06/07/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/07/2022 | 10/07/2024 |  |  | 02/29/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197520 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during repositioning of the impella, the pump got caught in the mitral valve structure resulting in the device entering the aorta and the surgeon being unable to cross the valve. the decision was made to replace this impella, during placement of the new pump the patient reportedly required multiple rounds of cardio-pulmonary resuscitation. no further patient-related issues were reported, and this pump was successfully replaced with a new impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the positioning issues has been completed. the investigation of the data logs showed the sudden occurrence of the "impella position in aorta" alarm, followed by erratic placement signal metrics and non-pulsatile motor current. after reviewing the clinical information, it was determined that the cause of the positioning issue was wireless re-access. the impella device is not indicated for wireless re-access per ifu 0048-9007. the cause of the positioning issue was wireless re-access resulting from improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
18147 | 1220648-2024-19692 | 20391129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391129&pc=OZD | 07/10/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 07/10/2022 | 10/07/2024 |  |  | 02/28/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022049844 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump stopped due to high motor current. the device was removed, and no patient harm was reported. | [5 empty] | The device was returned, and an investigation was completed. during testing, the device ran for at least 5 minutes after activation and resulted in a stoppage. the pump was then torn down and biomaterial was found in the rear bearing. this was found to be the most likely cause of the pump stoppage. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18148 | 1220648-2024-19693 | 20391239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391239&pc=OZD | 06/13/2023 | 10/07/2024 |  | 10/07/2024 |  |  | 06/14/2023 | 10/07/2024 |  |  | 03/31/2025 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328071 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient's hemoglobin decreased to the 9.5 level, and retroperitoneal hemorrhage was confirmed by cat scan. hemoglobin improved to the 10.0 level when the medical staff administered two units of red blood cells. although there was no problem with impella drive, weaning was proceeded and it was confirmed that the patient was hemodynamically stable. the device was surgically removed. the source of the bleeding is unknown. it was noted that the impella cp worked without issue. the patient survived the explant. | [5 empty] | The investigation for the reported issue has been completed. the impella device was discarded by the customer. the root cause of the access site bleeding was unable to be determined as no product was returned for review. this report is being filed as part of a retrospective review of historical records.
18149 | 1220648-2024-19673 | 20391278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391278&pc=OZD | 08/15/2021 | 02/13/2025 |  | 10/07/2024 |  |  | 08/15/2021 | 10/07/2024 |  |  | 05/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089973 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the impella cp device stopped after 12 days of use. the team attempted to troubleshoot, and finally removed and replaced the pump. no harm or injury was noted. | [5 empty] | The investigation for the reported pump stop issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause of the pump stop was blood ingress as a result of high purge pressure. the cause of the high purge pressure was most likely a kink in the purge line. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into the reported pump stop has been completed. the data logs confirmed the high purge pressure and purge system blocked alarms followed by a high motor current pump stop. the spike in motor current is instant suggesting a kink in the purge line. the pump was returned and the was unable to start. purge was run and the high purge pressure alarms were reproduced. the pump was torn down and blood was found in the motor. the root cause of the pump stop was determined to be blood ingress due to high purge pressure. this device was also used beyond its design life of 8 days per the design trace matrix. corrected information provided in h10.
18150 | 1220648-2024-19675 | 20391282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391282&pc=OZD | 08/21/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 08/21/2021 | 10/07/2024 |  |  | 05/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022090468 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a kinked cannula was noted. the device was removed and replaced without further issue or adverse impact to the patient. | [5 empty] | The investigation for the reported product damage (cannula kink) has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the low pump flow or suction or non-pulsatile motor current from a kinked cannula. this report is being filed as part of a retrospective review of historical records.
18151 | 1220648-2024-19678 | 20391328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391328&pc=OZD | 10/05/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 10/05/2021 | 10/07/2024 |  |  | 11/30/2022 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022068914 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella device was implanted through the left femoral artery (lfa) using ultrasound and a guided micro-puncture access technique. as the pump was initiated, high flow saturation was noted. the device was explanted, and the medical staff placed a new impella pump to manage this complication. | [5 empty] | The investigation for the reported issue has been completed. data logs were returned and confirm the high motor current/flows. the pump was returned and a fiber was found wrapped around the impeller. the pump was run in a bucket of water with the fiber and the saturated flows were reproduced. the fiber was removed and the flow reading and the motor current reading were within normal specification. the root cause of the saturated flows/high motor current was determined to be an ingested fiber wrapped around the impeller. the fiber is most likely a suture. this report is being filed as part of a retrospective review of historical records.
18152 | 1220648-2024-19704 | 20391365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391365&pc=OZD | 07/07/2022 | 03/28/2025 |  | 10/07/2024 |  |  | 07/07/2022 | 10/07/2024 |  |  | 03/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212146 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the device became kinked and was unable to deliver adequate flows after repositioning once the repositioning sheath was removed. the decision was made to place a new pump. no patient harm was reported. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The cause of the low pump flow was a cannula kink resulting from improper positioning upon sheath removal. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation positioning issues, and product damage (cannula kink) is not determined.
18153 | 1220648-2024-19682 | 20391371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391371&pc=OZD | 04/29/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 04/29/2022 | 10/07/2024 |  |  | 02/29/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192845 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, a kinked cannula with low pump flow after placement of the device. the pump was removed and replaced with no adverse impact to the patient. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
18154 | 1220648-2024-19701 | 20391629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391629&pc=OZD | 07/05/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 07/05/2022 | 10/07/2024 |  |  | 03/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211188 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during implantation, the physician was unable to advance the pump through the aortic valve. impella placement was aborted. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer, and therefore an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the reported delivery issues were most likely due to the patient's calcified vessels and aortic valve stenosis. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18155 | 1220648-2024-19688 | 20391641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391641&pc=OZD | 06/27/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/27/2022 | 10/07/2024 |  |  | 03/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204122 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during impella cp support, echocardiology was used and visualized the formation of micro-bubbles in the ventricle. the medical team was unable to identify if bubbles were from air or cavitation. the decision was made to explant the pump due to the microbubbles. | [5 empty] | The investigation for the reported embolism, has been completed. the returned pump was successfully run through de-air and then connected to the returned pump and purge exited the purge gap. purge values were within specification and no air in purge alarms reproduced. the cause of the embolism was most likely patient condition related since the returned product did not produce an air in purge alarm and purge operated to specification. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18156 | 1220648-2024-19715 | 20391665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391665&pc=OZD | 02/12/2023 | 10/07/2024 |  | 10/07/2024 |  |  | 02/12/2023 | 10/07/2024 |  |  | 09/30/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023284574 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the impella cp cannula was kinked due to severe aortic stenosis. successfully exchanged pump for a new impella cp. no patient harm was reported. | [5 empty] | The investigation into the reported delivery issue has been completed. the device was not returned for evaluation. clinical information was reviewed, and the root cause of the delivery issue was determined to be patient condition as the patient had severe aortic stenosis. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18157 | 1220648-2024-19679 | 20391669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391669&pc=OZD | 04/10/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 04/10/2022 | 10/07/2024 |  |  | 12/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163277 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that a pump stopped 12 hours before planned explant. several attempts were made to restart the device, but no success. there was no patient harm. | [5 empty] | The investigation for the reported pump stop issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the pump stop was bearing wear. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18158 | 1220648-2024-19685 | 20391677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391677&pc=OZD | 04/16/2023 | 10/07/2024 |  | 10/07/2024 |  |  | 04/16/2023 | 10/07/2024 |  |  | 12/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024319620 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp pump could not be fully advanced into the patient because of challenging anatomy. therefore, the insertion was aborted. the patient survived the procedure. | [5 empty] | A.5 is unknown. the impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18159 | 1220648-2024-19691 | 20391738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391738&pc=OZD | 06/27/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/29/2022 | 10/07/2024 |  |  | 04/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167973 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that that the patient was on impella cp support as a bridge to transplant. there was a suspicion that the patient had pulmonary bleeding due to hemoptysis and compatible computed tomography scan images. as a result, it was decided to explant the device and a centrimag implant was scheduled a few days later. | [5 empty] | The investigation for the reported patient injury has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the patient injury was determined to be patient condition and the relationship to impella is unknown as it was noted in the complaint that the lung bleed was not attributed to impella. this report is being filed as part of a retrospective review of historical records.
18160 | 1220648-2024-19696 | 20391801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391801&pc=OZD | 06/21/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/21/2022 | 10/07/2024 |  |  | 02/29/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202670 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was an impella cp pump stop. the staff (without warning) pushed the stretcher with some force into the ambulance and were not paying attention to the console. the white power cable to the cp was rammed into the door of the ambulance and snapped it out the console breaking the pins in the connector cable to automated impella controller (aic). the pump stopped and immediately alarmed. the connector cable was broken. the cp was replaced. | [5 empty] | The investigation for the reported pump stop issue has been completed. data logs showed the pump stopped almost immediately upon the occurrence of the "impella stopped (rpm deviations" alarm (#115). all metrics disappeared following the pump stop. visual inspection of the returned pump revealed the front half of the patient cable plug had been broken off with several connector pins completely missing. the remaining few connector pins were mangled. no other damage or abnormalities were found. after reviewing the clinical information, it was concluded that the cause of the pump stop was damaged connector pins due to excessive force on the patient cable. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18161 | 1220648-2024-19670 | 20391816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391816&pc=OZD | 10/27/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 10/27/2021 | 10/07/2024 |  |  | 06/30/2022 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1494424 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis and hematuria. fluids and repositioning were attempted to address the issue, but did not resolve the hemolysis. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation into the hemolysis has been completed. the impella device was not returned for evaluation. the most likely cause of the hemolysis was not determined. this report is being filed as part of a retrospective review of historical records.
18162 | 1220648-2024-19686 | 20391839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391839&pc=OZD | 04/30/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 04/30/2022 | 10/07/2024 |  |  | 01/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183102 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the physician was unable to get the impella cp device across the axillary artery. a new impella 5.5 was used without any adverse impact to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was possible. however, clinical details provided were to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
18163 | 1220648-2024-19681 | 20391885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391885&pc=OZD | 10/05/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 10/05/2021 | 10/07/2024 |  |  | 06/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022107201 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock and chest pain was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump flows dropped and motor current ¿mc¿ went flat. device was explanted to manage this complication. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported issue has been completed. the returned product was visually inspected and biomaterial was observed entraining the impeller. this report is being filed as part of a retrospective review of historical records.
18164 | 1220648-2024-19711 | 20391937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391937&pc=OZD | 09/21/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 09/21/2022 | 10/07/2024 |  |  | 04/30/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220120 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. following placement of the pump, the patient sat up resulting in positioning issue of the pump being pulled back across the aortic valve (av) regardless of the touch-borst valve being locked. this device was removed, and a new device placed. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the positioning issue has been completed. no product was returned. the data logs showed a sudden "impella position in aorta" alarm accompanied by an immediate drop in flow and motor current pulsatility. the pump performance metrics were unable to be recovered. after reviewing the clinical information, it was determined that the cause of the positioning issues was patient movement. the failure was confirmed, root cause patient movement. this report is being filed as part of a retrospective review of historical records.
18165 | 1220648-2024-19699 | 20391942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20391942&pc=OZD | 07/05/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 07/05/2022 | 10/07/2024 |  |  | 11/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166868 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a positioning issue with the impella cp when the patient was moved and the impella migrated out of the ventricle. the doctor was unsuccessful in repositioning. the decision was made to explant the pump and replace it. no patient harm was reported. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. after reviewing the provided data logs and clinical information, it was determined that the cause of the positioning issues was patient movement. this report is being filed as part of a retrospective review of historical records.
18166 | 1220648-2024-19705 | 20392125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392125&pc=OZD | 06/14/2023 | 10/07/2024 |  | 10/07/2024 |  |  | 06/14/2023 | 10/07/2024 |  |  | 01/31/2025 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328676 |  | 1000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella cp pump registered high purge pressure. the purge cassette was initially replaced during troubleshooting. however, the issue remained. due to the persistent alarm, the decision was made to explant the device and replace with another impella cp pump. the patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was discarded by the customer. data confirmed the failure mode. data showed motor current and placement signal spiked when the pump started andt the purge pressure (pp) increased gradually from 400 mmhg to 1100 mmhg with high pp alarms on. the high pp remained to the rest of the case. data also showed the pump started with flow 4.0 liters per minute (l/min) then dropped to 2.0 l/min and remained to the end of support. the product was not returned for review. the root cause of the high purge pressure was unable to be determined as limited clinical details were provided and no product was returned for review. this report is being filed as part of a retrospective review of historical records.
18167 | 1220648-2024-19689 | 20392139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392139&pc=OZD | 04/24/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 04/24/2022 | 10/07/2024 |  |  | 01/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181743 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, hemolysis was observed in the patient. it was stated that it was possible that the impella pigtail was tangled in the chordae, causing the catheter to move, or that suction interfered. it is unknown how hemostasis was achieved regarding hemolysis. weaning was successful, the device was removed, and the procedural outcome was the patient survived explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
18168 | 1220648-2024-19683 | 20392150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392150&pc=OZD | 10/10/2021 | 10/07/2024 |  | 10/07/2024 |  |  | 10/10/2021 | 10/07/2024 |  |  | 03/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022063218 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported immediately after changing purge cassette and bag, that there was a pump stop alarm. efforts to restart the impella pump were unsuccessful. the impella was explanted and replaced with a new pump. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported issue has been completed. review of the provided data logs confirmed an 'impella stopped - motor current high' alarm was triggered. multiple attempts were made to restart the pump, but were unsuccessful. of note, prior to the pump stop, a purge cassette exchange was performed. the new purge cassette was deaired for over 5 minutes before the yellow luer was reconnected to the pump. evaluation of the returned product found that the purge gap was large. the pump was torn down and bearing wear was observed. the root cause of the pump stop was bearing wear. as noted in the impella ifu: impella cp with smart assist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18169 | 1220648-2024-19698 | 20392217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392217&pc=OZD | 07/02/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 07/02/2022 | 10/07/2024 |  |  | 03/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023208577 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp pump did not start. all attempts to troubleshoot failed and the pump was replaced. no patient harm was reported. | [5 empty] | The investigation for the reported pump did not start issue has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was use related. this report is being filed as part of a retrospective review of historical records.
18170 | 1220648-2024-19700 | 20392233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392233&pc=OZD | 07/11/2022 | 03/13/2025 |  | 10/07/2024 |  |  | 07/11/2022 | 10/07/2024 |  |  | 03/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203650 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the pump exhibited purge pressure alarms and troubleshooting was attempted. the purge cassette was replaced, and eventually the pump was as well. the patient was taken off all anticoagulants due to bleed at all sites. additionally, it was reported that there was no systemic anticoagulants due to gastrointestinal bleed concerns. no further intervention was specified, but it was reported that the patient survived. | [5 empty] | This complaint has been identified as part of a retrospective review. the data logs confirm the increase in purge pressure and decrease in purge flow. the root cause of the high purge pressure was unable to be determined as no product was returned for analysis. the root cause of the access site bleeding - major and vascular damage - peripheral vessel could not be determined as insufficient clinical details were provided. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18171 | 1220648-2024-19697 | 20392234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392234&pc=OZD | 06/28/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/28/2022 | 10/07/2024 |  |  | 03/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200185 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there was a thrombus present upon explant of the impella cp. when the impella was removed, tissue-like material had adhered to the inhalation part and pigtail part. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the observed biomaterial was most likely patient condition. the pump ran without issue. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18172 | 1220648-2024-19676 | 20392384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392384&pc=OZD | 03/29/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 03/29/2022 | 10/07/2024 |  |  | 06/30/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094969 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, patient had hemolysis. there was no treatment and no lasting harm or injury to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. hemolysis was reported as possibly related to impella. the cause could not be determined, and relationship could not be confirmed due to lack of clinical details. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
18173 | 1220648-2024-19690 | 20392427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392427&pc=OZD | 06/13/2023 | 10/13/2025 |  | 10/07/2024 |  |  | 06/13/2023 | 10/07/2024 |  |  | 07/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255611 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  |  | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella pump was inserted with no complications. when trying to exchange the peel away with the repositioning sheath, the impella was moved above the aorta. the impella prolapsed and was folded in half above the aorta. the doctor was unable to reinsert the pump and decided to explant the impella cp and implant another impella cp. it was reported that the patient survived the procedure. | [5 empty] | A.5 ethnicity is unknown. the impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records. | The root cause of the product damage cannula kink could not be determined as insufficient clinical details were provided. the root cause of the positioning issue was use related since pump was found to be half folded during start of the case and later another cp pump was placed successfully. g1 reporting contact email revised on manufacturer device report 1220648-2024-19690 in accordance with updated procedures.
18174 | 1220648-2024-19703 | 20392478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392478&pc=OZD | 06/14/2023 | 10/07/2024 |  | 10/07/2024 |  |  | 06/14/2023 | 10/07/2024 |  |  | 03/31/2025 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342341 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that an echocardiogram verified a thrombus in the right ventricular outflow tract and left ventricle. no medical intervention was reported. the medical staff noted that there is no concern for mobility or interference with extracorporeal membrane oxygenation and the impella. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the product was not returned for investigation. data logs were returned. the root cause of the thrombus issue was most likely patient condition related to the clot found in the right ventricle outflow track and the left ventricle. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18175 | 1220648-2024-19707 | 20392484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392484&pc=OZD | 11/16/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 11/23/2022 | 10/07/2024 |  |  | 01/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185918 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced a cerebral hemorrhage. the patient was kept on impella support throughout the procedure and survived. | [5 empty] | The investigation for the reported stroke/cva has been completed. the product was returned and stroke/cva was confirmed. the root cause of the stroke/cva was determined to be sidearm filter damage due to excessive force. as noted in the impella ifu: cva/stroke: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18176 | 1220648-2024-19709 | 20392511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392511&pc=OZD | 07/07/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 07/07/2022 | 10/07/2024 |  |  | 12/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173743 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that when the pump was plugged into the aic console, it displayed a yellow "pump failure" message. a new pump was used to continue the case, and no patient harm was reported. | [5 empty] | The device was returned, and the investigation has been completed. during testing, the device went through case start without issue. the imc logs from the field were reviewed and confirm the impella defective alarm. the root cause of the impella defective alarm/ pump not detected was most likely use related, as the pump had no issue going through case start when returned. this report is being filed as part of a retrospective review of historical records.
18177 | 1220648-2024-19694 | 20392540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392540&pc=OZD | 06/01/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 06/01/2022 | 10/07/2024 |  |  | 12/31/2023 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170105 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the team noted signs of hemolysis in the rise of pfhg. the issue was resolved by repositioning the pump. | [5 empty] | The investigation for the reported hemolysis issue has been completed. the impella device has not been returned for evaluation. the cause of the issue was not determined since product was not returned. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
18178 | 1220648-2024-19708 | 20392611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392611&pc=OZD | 07/07/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 07/07/2022 | 10/07/2024 |  |  | 02/29/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195604 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. during insertion, the pump failed to deliver when the easy guide lumen was removed in error. the physician could not load and advance the pump. the doctor decided to open another pump and discard the first. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was most likely the incorrect removal of the red lumen. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18179 | 1220648-2024-19702 | 20392659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392659&pc=OZD | 06/14/2023 | 10/07/2024 |  | 10/07/2024 |  |  | 06/14/2023 | 10/07/2024 |  |  | 03/31/2025 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024345291 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that there was low pump flow with the impella cp. upon start up, impella flows dropped immediately. the physician attempted re-starting the pump without success. continuous suction alarms were noted while the pump was running. the patient survived the procedure. | [5 empty] | A.5 is unknown. the investigation for the reported issue has been completed. the pump was sent back for investigation. a significant amount of biomaterial was found on the impeller and cannula. the root cause of the issue was biomaterial. this report is being filed as part of a retrospective review of historical records.
18180 | 1220648-2024-19718 | 20392785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392785&pc=OZD | 06/22/2023 | 10/07/2024 |  | 10/07/2024 |  |  | 06/22/2023 | 10/07/2024 |  |  | 03/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212176 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported placement of the impella was aborted as the patient had a previous iliac stent and the impella would not pass even while using the long 14 french oscor sheath. there was no patient harm reported and this device was removed. | [5 empty] | The investigation into the delivery issue has been completed. the product was discarded. the cause of the delivery issue- anatomy was patient condition related because the patient had a previous iliac stent and impella would not pass even while using the long 14 french oscor sheath. the complaint device was not returned for investigation. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18181 | 1220648-2024-19706 | 20392861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392861&pc=OZD | 07/07/2022 | 10/07/2024 |  | 10/07/2024 |  |  | 07/07/2022 | 10/07/2024 |  |  | 02/29/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202060 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. the team was unable to deliver the pump when the wire became entangled in the easy guide lumen. due to the chance of pump damage the case support was started with a new pump. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the reported issue could not be determined. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18182 | 1220648-2024-19717 | 20392867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392867&pc=OZD | 06/20/2023 | 10/07/2024 |  | 10/07/2024 |  |  | 06/20/2023 | 10/07/2024 |  |  | 01/31/2025 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024314817 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella placement was aborted due to the physician being unable to advance the impella through the patient¿s vasculature. there was no patient harm reported and this device was removed. | [5 empty] | The investigation into the delivery issues has been completed. the product was not returned. the cause of the delivery issue was patient condition related to abnormal anatomy. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18183 | 1220648-2024-19710 | 20392869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392869&pc=OZD | 07/08/2022 | 09/19/2025 |  | 10/07/2024 |  |  | 07/08/2022 | 10/07/2024 |  |  | 02/29/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196945 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | Defective Device | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during implantation, there were delivery issues and introducer issues with the impella cp. when the inlet was entering the aorta, both the pump and the sheath kinked. both were removed and a new sheath and device was placed. there was no reported patient harm. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records. | No product was returned. after reviewing the clinical information it was determined that the cause of both the delivery issue and the introducer issue was most likely patient condition related, specifically the patient's tortuous vessels. the cause of the cannula kink was most likely patient condition related, specifically the patient's tortuous vessels.
18184 | 1220648-2024-19712 | 20392911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392911&pc=OZD | 02/11/2023 | 10/07/2024 |  | 10/07/2024 |  |  | 02/11/2023 | 10/07/2024 |  |  | 10/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290121 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during impella placement when attempting to cross the aortic valve with the impella the cannula kinked and buckled, resulting in the delivery issue and not being able to cross into the left ventricle. reportedly two attempts were made before the physician made the decision to abort the placement of this impella and placed a new wire and new impella pump. this impella pump was replaced with another impella pump. | [5 empty] | The investigation into the delivery issue issue has been completed. the impella pump was returned for investigation. a visual inspection of the returned pump noted a cannula kink was observed. the root cause of the issue was use related to improper technique resulting in the kinked cannula, which obstructed the delivery process. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18185 | 1220648-2024-19713 | 20392915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392915&pc=OZD | 02/12/2023 | 10/07/2024 |  | 10/07/2024 |  |  | 02/12/2023 | 10/07/2024 |  |  | 12/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024294919 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the physician stated the pump did not prime correctly. reportedly the console jumped over the purge prime wizard and went straight to pressure zero. a new automated impella controller (aic) and purge cassette were tried and had the same issue. the next pump that was placed was damaged by the doctor, however a implant with 3rd pump was successful. there was no patient harm reported. | [5 empty] | The investigation into the priming issue has been completed. the impella pump was returned for investigation. the investigator reviewed the logs and noted the purge cassette and pump were connected to the console prior to case start, leading to the software skipping over steps 2-3 since the cassette and disc were already plugged in. no abnormalities were found on the returned pump. the pump was reprogrammed and successfully ran through case start. the pump was then successfully run in a bucket of water. the product operated to specification. the cause of the priming issue was use and technique related, specifically plugging in the purge cassette and pump prior to initiating case start. this report is being filed as part of a retrospective review of historical records.
18186 | 1220648-2024-19714 | 20392916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392916&pc=OZD | 04/16/2023 | 02/27/2025 |  | 10/07/2024 |  |  | 04/16/2023 | 10/07/2024 |  |  | 11/30/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024301301 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported a failure to prime the device. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported priming issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that pump was reprogrammed to enable to run the case start. the pump went through the case start without issue. no issues were found on the pump under visual inspection. therefore, the root cause of the priming issue was most likely due use issue. this report is being filed as part of a retrospective review of historical records.
18187 | 1220648-2024-19716 | 20392931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20392931&pc=OZD | 02/12/2023 | 10/07/2024 |  | 10/07/2024 |  |  | 02/13/2023 | 10/07/2024 |  |  | 12/31/2024 |  | 10/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024289254 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the physician was placing a second pump, after not priming the first pump correctly, and damaged the pump during placement. the physician had to use a third pump to complete the intervention. this pump was damaged by the physician and was replaced with another pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the reported product damage has been completed. the cause of the product damage was use related, specifically improper technique during case setup when backloading the guidewire. this report is being filed as part of a retrospective review of historical records.
18188 | 1220648-2024-19500 | 20379543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379543&pc=OZD | 04/21/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 04/21/2023 | 10/06/2024 |  |  | 12/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322297 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with high-risk percutaneous intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. it was reported that the device was unable to be delivered due to patient¿s anatomy. the patient's condition caused the long sheath to kink on insertion. the pump insertion was aborted and a new impella cp device was used without further issue. there was no harm to the patient. | [5 empty] | The investigation for the reported product damage (cannula kink) issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the product damage (cannula kink) issue was patient condition related. the patient had excess adipose tissue, and a deep femoral artery, causing the long sheath to kink on insertion. this report is being filed as part of a retrospective review of historical records.
18189 | 1220648-2024-19501 | 20379803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379803&pc=OZD | 03/27/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 03/27/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024300842 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the physician decided to escalate the patient from the impella cp to an impella 5.5. upon removal, it was noted that there was extensive clot on the impella cp. there was no further reported harm to the patient. | [5 empty] | The device was returned, and the investigation for the reported thrombus has been completed. the cp was returned and thrombus was confirmed around the outflow. the root cause of the thrombus was most likely patient condition. as noted in the impella ifu: embolism/thrombus: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18190 | 1220648-2024-19502 | 20379807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379807&pc=OZD | 03/29/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 03/29/2023 | 10/06/2024 |  |  | 01/31/2025 |  | 10/06/2024 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309355 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with unknown race/ethnicity and unknown primary indication was implanted with an impella cp device for mechanical circulatory support. after starting the pump, there was a sudden increase of purge pressure. the impella flow was reduced, but subsequently the pump was explanted and replaced, which resolved the issue. there was no reported patient harm. | [5 empty] | The investigation for the reported high purge pressure has been completed. the product was returned and the high purge pressure was confirmed. the returned pump was visually inspected and biomaterial in the purge gap was observed. no other abnormalities were observed. the cause of the issue was biomaterial in the purge gap. this report is being filed as part of a retrospective review of historical records.
18191 | 1220648-2024-19503 | 20379864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379864&pc=OZD | 09/10/2024 | 11/21/2024 |  | 10/06/2024 |  |  | 09/10/2024 | 10/06/2024 |  |  |  |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494525 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient on impella cp support along with extra corporeal membrane oxygenation for 5 days when the patient suffered a cerebrovascular accident (cva) and expired. the cva was hemorrhagic in nature. the patient had cardiac arrest outside the hospital and the cause of the cerebral hemorrhage was unknown because the cerebral hemorrhage was widespread. therefore, it was determined that the death was not related to the impella. | [5 empty] | The investigation stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
18192 | 1220648-2024-19504 | 20379874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379874&pc=OZD | 09/10/2024 | 11/21/2024 |  | 10/06/2024 |  |  | 09/10/2024 | 10/06/2024 |  |  |  |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528980 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support the patient had limb ischemia with damage. support was continued. the patient survived the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the low pump flow and myocardial ischemia has been completed. data logs were reviewed which did not show motor current (mc) spikes. many suction alarms observed. there were no positioning issues. placement signal (ps) no reliable alarms were observed. case started with low mc and low pump flow at p-9 and flow went back to required flow range at p-6. the cause of the low pump flow issue cannot be determined since the complaint device not returned for analysis and insufficient clinical data provided. the root cause of the ischemia could not be determined without additional clinical details. added-h6 med dev prob code 1248 corrections to the initial mfr report: b5-it was initially reported that the patient had limb ischemia with damage. this should have stated that the patient had myocardial ischemia with damage. additionally, it was reported that there were existing suction alarm during support.
18193 | 1220648-2024-19505 | 20379875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379875&pc=OZD | 03/29/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 03/29/2023 | 10/06/2024 |  |  | 09/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278080 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there were continuous suction alarms. the pump was repositioned several times, but the problem was not resolved. no additional informano was provided related to the resolution. no patient harm was reported. | [5 empty] | The device was returned, and the investigation for the reported low pump flow has been completed. the cause of the low pump flow was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
18194 | 1220648-2024-19506 | 20379877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379877&pc=OZD | 09/10/2024 | 11/12/2024 |  | 10/06/2024 |  |  | 09/10/2024 | 10/06/2024 |  |  |  |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525012 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump delivered for an acute myocardial infarction/cardiogenic shock patient. the pump was placed but during the support there was persistent suction and the team questioned a bleed at the retroperitoneum. there was no visible access site hematoma. the team gave eight units of blood and explanted the pump after just over three hours and the patient expired. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18195 | 1220648-2024-19507 | 20379880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379880&pc=OZD | 04/01/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 04/01/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024313369 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the physician was not satisfied with the flow rate. the impella was removed, flushed, and reinserted, but the issue persisted. the impella was removed, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia could not be determined. this report is being filed as part of a retrospective review of historical records.
18196 | 1220648-2024-19508 | 20379882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379882&pc=OZD | 04/05/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 04/06/2023 | 10/06/2024 |  |  | 05/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225186 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, flows were low and suction alarms occurred. upon further investigation it was also noted that the cannula kinked. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the product was returned, and a kink was confirmed in the cannula. no other damage or abnormalities were found. in conclusion, it was determined that the cause of the low pump flow was a kinked cannula. this report is being filed as part of a retrospective review of historical records.
18197 | 1220648-2024-19509 | 20379887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379887&pc=OZD | 04/05/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 04/05/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311687 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient of unknown demographics noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the doctor attempted both sides to place the impella. the doctor attempted a long sheath, but too stenosed to passed. then the doctor attempted short sheath but was unable to pass the impella past lesions; therefore, the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of difficulty inserting/removing introducer. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the difficulty inserting/removing introducer was due to patient condition related with patient having stenosis. this report is being filed as part of a retrospective review of historical records.
18198 | 1220648-2024-19510 | 20379890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379890&pc=OZD | 09/10/2024 | 11/19/2024 |  | 10/06/2024 |  |  | 09/10/2024 | 10/06/2024 |  |  |  |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447461 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support developed a hematoma and oozing from the right femoral artery. there is a venous line along with the impella arterial site in the right femoral. manual pressure was applied, and the dressing was cleaned. the tuohy was loose but the nurse tightened and two units of blood were given. the impella was removed with no complications. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
18199 | 1220648-2024-19511 | 20379907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379907&pc=OZD | 09/10/2024 | 11/18/2024 |  | 10/06/2024 |  |  | 09/10/2024 | 10/06/2024 |  |  |  |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544155 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Defective Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after an impella cp was placed and during high risk percutaneous coronary intervention procedure, a 7 french destination sheath was attempted for single access but was unsuccessful. a companion sheath was then used successfully. initial attempt of single access noted there was oozing coming out of hemostatic valve. there was no more oozing with the companion sheath and the percutaneous coronary intervention proceeded. there was no reported harm to the patient. | [5 empty] | The introducer was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation into access site bleeding has been completed. the product was changed to the 14 french sheath as that was what the costumer was complaining about. both the 14 french sheath and the 7 french were returned but damages were noted on only the 14 french. a leak test was set up using a syringe luer, manometer luer, and long introducer flush valve luer all connected together via 3 way luer tee. the syringe was filled with colored water and pressure was slowly ramped up to reproduce a leak. a leak initially reproduced at 170 mmhg on one side of sheath score line. a leakage rate of less than 2.5g per minute at a pressure of 180 mmhg is acceptable when a dilator is in the 14 french sheath. the returned introducer did not meet its 180 mmhg pressure specification during reproduction testing. a leak was reproduced on the returned introducer at pressures under the 180mmhg specification. the root cause of the access site bleeding - major is due to a sheath score line split. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ during review of the investigation results, the failure mode was updated to access site bleeding. the following fields were corrected since the initial manufacturer device report number 1220648-2024-19511 was submitted according to the investigation results. b.1 changed from product malfunction to adverse event. b.2 added required intervention. h.1 changed from malfunction to serious injury. h.6 changed health effect - clinical code and medical device problem code. h.10 added instructions for use for the new failure mode detected in the investigation results. b.7 added information as it was inadvertently omitted on the initial manufacturer device report number 1220648-2024-19511.
18200 | 1220648-2024-19512 | 20379908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379908&pc=OZD | 05/16/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/16/2023 | 10/06/2024 |  |  | 12/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024317949 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the technician was unable to unlock the tuohy during catheter preparation. the technician reportedly attempted to unlock the tuohy by using hemostats and ended up ripping the sterile sleeve. this impella was unable to be used due to concerns of sterility and was replaced with another impella. there was no patient harm reported, and this device was not used. | [5 empty] | The investigation into the product damage issue has been completed. the impella pump was returned for investigation. the visual inspection confirmed the clinical details indicating the sterile sleeve was ripped by the technician. the root cause of this issue was due to improper technique by the end user resulting in the sterile sleeve being ripped. this report is being filed as part of a retrospective review of historical records.
18201 | 1220648-2024-19513 | 20379910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379910&pc=OZD | 05/17/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/17/2023 | 10/06/2024 |  |  | 02/28/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325260 |  | 1000080 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the impella was successfully placed, however when the pump was stated a severe kink just distal to motor developed that was flow limiting. reportedly the pump was only able to flow at 1.0 liters per minute and had to be maintained on pressure level two with continuous suction during the case. the physician unsuccessfully attempted multiple times to reposition the impella to resolve the kink, however as soon as the guide wire was removed the pump would kink. the patient¿s ascending aorta was noted to have a very sharp 90 degree turn around the arch that seemed to be causing kink. additionally, the clinician provided the physician with the option of trying a new impella, however the physician declined and chose to proceed with the case indicating if the patient decompensated a new impella would be used. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the low pump flow has been completed. per the clinical summary the low flow occurred due to a cannula kink, the patient had a small aortic arch. the data review confirmed the low flow when the pump was started and remained for the rest of the case with auto suction response and suction alarms. based on the clinical description and the data analysis, the root cause of low flow was due to a cannula kink and the root cause of cannula kink was due to patient condition. this report is being filed as part of a retrospective review of historical records.
18202 | 1220648-2024-19514 | 20379912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379912&pc=OZD | 05/17/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/17/2023 | 10/06/2024 |  |  | 01/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2024309692 |  | 0048-0024-JP |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was experiencing cardiopulmonary arrest was to have an impella cp placed for mechanical circulatory support. it was reported that the impella could not be placed into the left ventricle due to the narrowness of the left chamber of the heart. the impella placement was aborted due to the anatomy of the heart and the device removed. there was no patient harm reported, and this device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the difficulty inserting the introducer due to the patient¿s condition has been completed. per clinical details the cause of the issue was patient condition related. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18203 | 1220648-2024-19515 | 20379915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379915&pc=OZD | 08/30/2024 | 11/06/2024 |  | 10/06/2024 |  |  | 09/10/2024 | 10/06/2024 |  |  |  |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete | Other |  | Yes | Initial submission | Followup | [8 empty] | The hospital reported to the united states food and drug admisinstration a user medwatch for a patient who suffered a stroke after coming out of a triple bypass procedure and never recovered. the patient was takent to recovery and seven hours later, the patient had not awoke yet. the intesnstive care unit staff noticed signs of a stroke. the patient did not recover and expired. the team noted concerns about the influence of the impella during the triple bypass. due to the redacted user medwatch, the pump type and patient/hospital demographics are unknown. | [5 empty] | A.1 is unknown. d.4 serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1, e.2 and e.3 are unknown. h.4 device manufacture date is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
18204 | 1220648-2024-19516 | 20379918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379918&pc=OZD | 05/08/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/08/2023 | 10/06/2024 |  |  | 12/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024326184 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. a transesophageal echocardiogram (tee) showed impella shallow with pigtail behind posterior papillary. the impella cp was decreased to p2 for repositioning. interaction with mitral valve, patient's rhythm worsening and severe hypotension noted. repositioning attempts were unsuccessful. a new impella cp was placed. | [5 empty] | The investigation for the reported positioning failure issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the positioning failure issue was use related. this report is being filed as part of a retrospective review of historical records.
18205 | 1220648-2024-19517 | 20379919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379919&pc=OZD | 05/09/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/09/2023 | 10/06/2024 |  |  | 01/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024313542 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the impella cp device came out of position in the left ventricle (lv). when the team was unable to reposition, the pump was explanted and replaced. there was no patient harm reported. | [5 empty] | The investigation for the reported access site bleeding issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the positioning issue was patient movement. this report is being filed as part of a retrospective review of historical records.
18206 | 1220648-2024-19518 | 20379920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379920&pc=OZD | 08/17/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 08/17/2021 | 10/06/2024 |  |  | 10/31/2022 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021003190 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a kink in the cannula. the impella was removed. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the product damage has been completed. no data logs were provided for review. the clinical details provided were insufficient to determine a root cause. the cause of the cannula kink was not determined due to insufficient information, no data logs, and the product not being returned. this report is being filed as part of a retrospective review of historical records.
18207 | 1220648-2024-19520 | 20379936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379936&pc=OZD | 08/10/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 08/10/2021 | 10/06/2024 |  |  | 05/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022085926 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported there was a blood clot/thrombus on the shaft of the impella cp when it was removed. the patient survived the procedure. | [5 empty] | A.5 ethnicity is unknown. the investigation for the reported thrombus issues has been completed. the impella device has been returned for evaluation. visual inspection of the returned pump revealed significant amounts of biomaterial wrapped around the cannula and ingested into the inflow and outflow cages. therefore, it was confirmed that biomaterial was observed on the pump at the time of pump removal. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18208 | 1220648-2024-19519 | 20379937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379937&pc=OZD | 05/09/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/09/2023 | 10/06/2024 |  |  | 02/28/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024324708 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that after implanting the device the patient was transferred to the intensive care unit (icu) bed when the access site started bleeding. a decision was made to remove the device and the access site was closed. the patient required one unit of blood, and tolerated device removal well. | [5 empty] | The investigation for the reported access site bleeding issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the issue was patient movement due to improper technique. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18209 | 1220648-2024-19521 | 20379938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379938&pc=OZD | 05/27/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/27/2023 | 10/06/2024 |  |  | 02/28/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321891 |  | 1000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant had a cp support that was ongoing when the placement signal was unreliable. the pulmonic stenosis issues did not prompt the team to explant the pump. the cp support was in for two days when then explanted. the patient survived and there was no lasting harm or injury, | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. data analysis: data showed all 5s parameter was in specification. there was no spike or trend for placement signal during the case. data also showed low flow occurred when pump started that triggered auto suction response & suction alarms along with ps low. all these alarms were off when pump was run on p-6 to the rest of the case. product was not returned for review. based on clinical description & data analysis, the root cause of optical signal issue was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
18210 | 1220648-2024-19522 | 20379940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379940&pc=OZD | 03/08/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 03/08/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024299551 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the staff noticed that the right leg color was pale, pulses could not be found with doppler and skin was mottled. pulses did not improve with weaning of pressors and volume. the patient was then brought back to cath where left femoral artery (lfa) to right femoral artery (rfa) reperfusion sheath was inserted. flow was restored, color improved and doppler pulses to right foot. the patient had an episode of atrial fibrillation (afib) with rapid ventricular rate (rvr). the physician was unable to reposition shallower as the catheter would jump out of the lv. the pump was explanted. | [5 empty] | The investigation for the reported limb ischemia and ventricular fibrillation issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the limb ischemia issue was patient condition related to diseased/ small vessels. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
18211 | 1220648-2024-19523 | 20379945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379945&pc=OZD | 03/09/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 03/09/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024302462 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the medical staff attempted several times but was unable to push the impella past the aortic valve. the decision was made to proceed with the percutaneous coronary intervention (pci) and abort impella support. there was no harm to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the delivery issue was patient condition related. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18212 | 1220648-2024-19525 | 20379948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379948&pc=OZD | 05/28/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/28/2023 | 10/06/2024 |  |  | 01/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024315565 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the device exhibited suction alarms, the device was deep in the ventricle. the physician repositioned the device however this did not correct the problem. the physician decided to explant the device. no further information is available. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. complaint purge cassette visually inspected. purge cassette cut open. performed purge leak testing. leak observed from piston seal. root cause of the purge cassette leak was piston failure. complaint cp pump was visually inspected. no broken blade and no cannula kink observed. biomaterial observed around impella and inside the cannula that obstructed the pump flow. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18213 | 1220648-2024-19524 | 20379949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379949&pc=OZD | 08/27/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 08/27/2021 | 10/06/2024 |  |  | 05/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080573 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during insertion, the pump encountered force while trying to cross the aortic valve and became kinked. device was replaced with another impella cp pump. no patient harm was reported. | [5 empty] | The investigation for the reported product damage (cannula kink) has been completed. the product was returned and the cannula kink was confirmed. the cause of the delivery issue cannula kink was most likely use related since excessive force was reported during insertion. this report is being filed as part of a retrospective review of historical records.
18214 | 1220648-2024-19527 | 20379950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379950&pc=OZD | 05/15/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/15/2023 | 10/06/2024 |  |  | 01/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183101 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia. actions taken to resolve are unknown. additionally, chronic superficial femoral artery occlusion was discovered by angiogram upon removal of the sheath. the physician unsuccessfully attempted intervention, however the surgeon was unable to cross the chronic superficial femoral artery occlusion. the patient was transferred to the intensive care unit. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the limb ischemia reversable has been completed. the cause of the limb ischemia- reversible was chronic occlusion distal to the sheath; the reported issue is patient condition related and unrelated to impella. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
18215 | 1220648-2024-19526 | 20379951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379951&pc=OZD | 07/25/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 07/25/2023 | 10/06/2024 |  |  | 04/30/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351283 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported due to the patient¿s vasculature, vessels would not accommodate the impella, the physician was unable to place the impella. there was no patient harm reported, and this device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. the cause of the delivery issue was patient condition related due to diseased right femoral vessels per clinical information. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluation" this report is being filed as part of a retrospective review of historical records.
18216 | 1220648-2024-19528 | 20379952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379952&pc=OZD | 05/12/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/12/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024314633 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. low purge flow throughout case. dropped to <2ml/hour at the end of case and caused a high purge/blocked purge alarm. impella cp was explanted and replaced with another impella cp. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the log confirms high purge pressure at the end of the case and shows a motor speed drop at the time that the purge pressure becomes high. biomaterial and dried dextrose were observed in and around the purge gap. the pump was rinsed with water and biomaterial was still seen in the purge gap. the pump was attached to an aic and a purge cassette, and the high purge pressure was reproduced. the catheter was cut near the motor housing and fluid easily exited the catheter. the cause of the high purge pressure was biomaterial. this report is being filed as part of a retrospective review of historical records.
18217 | 1220648-2024-19529 | 20379953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379953&pc=OZD | 09/02/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 09/02/2021 | 10/06/2024 |  |  | 10/31/2022 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021017619 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported that there was an alarm showing that the position detection signal was not stable and that the position waveform could not be drawn. device was explanted. no patient harm reported. | [5 empty] | The investigation for the reported placement signal issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. this report is being filed as part of a retrospective review of historical records.
18218 | 1220648-2024-19530 | 20379954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379954&pc=OZD | 05/29/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/29/2023 | 10/06/2024 |  |  | 12/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325682 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited alarms and a flat motor current with suspected clot entrainment. the decision was made to remove the pump. it was pulled back across the aortic valve while they were preparing to remove. a new pump was not inserted as the patient has already been hemodynamically stable. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. no data logs were returned and there is no data available on constellation. the cause of the saturated flow was most likely biomaterial based on the provided clinical information. this report is being filed as part of a retrospective review of historical records.
18219 | 1220648-2024-19531 | 20379957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379957&pc=OZD | 05/19/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/19/2023 | 10/06/2024 |  |  | 03/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024327455 |  | 1000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a high motor current followed by a pump stop alarm occurred during impella support. the impella was able to restart, however, flows notably decreased. the pump was subsequently replaced to resolve the reported issue. it was noted following explant that a thrombus was discovered within the left ventricle. there was no patient harm reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the logs show a high motor current pump stop when the first pump start is attempted. the pump is then able to be started and was run successfully. once the pump is running there are suction alarms and motor current spikes. both of these events indicate that the pump is likely contacting biomaterial. the cause of the pump not starting initially was most likely biomaterial based on the log analysis and the observation of a left ventricle thrombus. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18220 | 1220648-2024-19532 | 20379958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379958&pc=OZD | 05/29/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/29/2023 | 10/06/2024 |  |  | 03/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024339505 |  | 1000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported red urine approximately one hour into impella support. repositioning resolved the issue and patient's urine cleared. the patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. no product was returned for investigation. no abnormalities were observed in the returned log. there are no logs from the first hospital before the patient was transferred. the cause of the hemolysis was most likely improper positioning of the device based on the provided clinical information that it was confirmed through imaging that the pump was not in a good position and urine color improved after repositioning. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
18221 | 1220648-2024-19533 | 20379959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379959&pc=OZD | 05/30/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/30/2023 | 10/06/2024 |  |  | 12/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295688 |  | 0048-0024-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that during attempted delivery of an impella cp pump, the guidewire met resistance and subsequently bent. the guidewire was removed, and a new guidewire was utilized to deliver the pump. upon the later completion of support, the pump was explanted, and it was discovered that the shaft of the cannula was kinked. no report of patient harm. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the introducer issue was the use of a sheath that was not indicated for use with impella. small vessel diameter reported. the cause of the issue was most likely a kinked cannula from improper delivery technique and small patient vessel. product not returned to verify the kink. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18222 | 1220648-2024-19534 | 20379960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379960&pc=OZD | 05/19/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/19/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311984 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported access site bleeding after removal of introducer and advancement of the repositioning sheath. hemostasis could not be controlled; therefore, the pump was removed. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on clinical description, the root cause of access site bleeding was most likely due to use issue, as perclose had failed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18223 | 1220648-2024-19535 | 20379961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379961&pc=OZD | 05/30/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/30/2023 | 10/06/2024 |  |  | 03/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331328 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there was an impella cp pump stop. despite multiple attempts, the pump would not start. the pump was successfully explanted. no further information provided. | [5 empty] | The investigation for the reported issue has been completed. the pump stopped on day eighteen of support. the impeller purge gap was observed to be very large on the returned pump. the cause of the pump stop was bearing wear. the pump stop occurred beyond the eight-day impella cp indication for use as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18224 | 1220648-2024-19536 | 20379962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379962&pc=OZD | 05/30/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/30/2023 | 10/06/2024 |  |  | 09/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276154 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Implant Pain | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported a vasovagal reaction during removal of the peel away introducer. neo-synephrine spray was used to reverse the reaction. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the patient injury, specifically the vasovagal reaction, was most likely patient condition related. this report is being filed as part of a retrospective review of historical records.
18225 | 1220648-2024-19537 | 20379963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379963&pc=OZD | 08/03/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 08/03/2021 | 10/06/2024 |  |  | 06/30/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022093377 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped. the issue could not be resolved with normal troubleshooting attempts. therefore, the pump was removed. no patient harm was reported and the patient survived the explant. | [5 empty] | The investigation into the pump stop issue has been completed. the returned data logs showed the pump running without issue until an "impella disconnected while running" alarm occurred. the electrical resistances of the motor cable lines were within specification. dried blood was found in the red pump handle and a small hole was found in the catheter which was likely covered during support. thus, the root cause of the impella defective alarm was due to the hole in the catheter, causing blood to leak into the red pump handle and corrode the electrically erasbale programmable read-only memory. this was made known to the field after an inadvertent disconnect of the pump from the console. the cause of the pump stop was an impella defective alarm, which occurred due to a hole in the catheter which caused blood to ingress into the red pump plug. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18226 | 1220648-2024-19538 | 20379964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379964&pc=OZD | 05/30/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/30/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311186 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient was being placed on impella cp for hemodynamic support. it was reported that the complainant reported low pump flow during use of the device, with a clot identified on the inflow upon removal. the pump was exchanged with no reported harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. no product was returned for evaluation. the cause of the saturated flow issue was most likely due to biomaterial with motor current spike observed in data log review and clot formation on inflow cage per clinical review. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18227 | 1220648-2024-19539 | 20379965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379965&pc=OZD | 06/22/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 06/22/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283070 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that impella did not go through normal set up mode, instead, it jumped straight to implant screen missing out red side arm flush screen. as advised, staff changed the catheter and discarded the current one. no reported harm to patient. | [5 empty] | The investigation for the reported issue has been completed. the pump was returned to the lab. the pump was able to go through the case start steps, and the sn was recognized. no logs were returned from the field to confirm that priming steps were skipped. the root cause of the priming issue cannot be determined. this report is being filed as part of a retrospective review of historical records.
18228 | 1220648-2024-19540 | 20379966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379966&pc=OZD | 08/28/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 09/07/2021 | 10/06/2024 |  |  | 01/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022045979 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon removal of the impella, lower limb ischemia was discovered. no additional measures were taken based on the ischemia. the patient had remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the limb ischemia has been completed. no data logs were provided for review. the clinical details provided were insufficient to determine a root cause. the cause of the limb ischemia was not determined due to insufficient information, no data logs, and the product not being returned. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
18229 | 1220648-2024-19541 | 20379967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379967&pc=OZD | 06/01/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 06/01/2023 | 10/06/2024 |  |  | 09/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278090 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the complainant reported the impella insertion was aborted due to inadequate iliac anatomy. no further information is available. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. root cause of the delivery issue was patient condition related because unable to deliver catheter through small iliac caliber as the patient had peripheral artery disease (pad)/peripheral vascular disease (pvd). as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18230 | 1220648-2024-19542 | 20379968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379968&pc=OZD | 06/01/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 06/01/2023 | 10/06/2024 |  |  | 03/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331666 |  | 0048-0014 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the complainant reported a failure to deliver the impella device due to patient's vasculature. insertion was aborted with no reported harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was patient condition related, because cp insertion aborted due to patient anatomy and impella resistance was in the aortic valve. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18231 | 1220648-2024-19543 | 20379969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379969&pc=OZD | 06/02/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 06/02/2023 | 10/06/2024 |  |  | 01/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333611 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there were low pump flows with the impella cp. after placement, flows suddenly went from 3.5 to 1.5 liters per minute. placement was good and there was plenty of volume, but a sudden change occurred which limited the flows. the device was explanted and replaced with another impella cp. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. data logs were returned and investigated. per the clinical information provided, the root cause of the low pump flow issue was most likely a kinked cannula. this report is being filed as part of a retrospective review of historical records.
18232 | 1220648-2024-19544 | 20379970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379970&pc=OZD | 06/05/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 06/05/2023 | 10/06/2024 |  |  | 03/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024335873 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with unknow race and ethnicity with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while of support the pump stopped, an attempt to restart was attempted but was unsuccessful. the pump was weaned. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. according to the clinical, on the tenth day of impella cp support a pump stop occurred. in an attempt to resolve the issue, the restart procedure was performed. however, it was unsuccessful, so the decision was made to wean the patient off of support and remove the pump. product evaluation confirmed the presence of bearing wear. data logs were not returned for analysis. the cause of the pump stop was bearing wear due to overuse of the pump longer than the specified duration. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18233 | 1220648-2024-19545 | 20379971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379971&pc=OZD | 08/11/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 08/11/2021 | 10/06/2024 |  |  | 06/30/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095101 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was placed and initial flows were 3.6 liters. reportedly, during the attempt to advance the impella cp, the impella kinked resulting in flows unable to reach past 2.5 liters. the physician made the decision to leave the impella cp in place and complete the percutaneous coronary intervention regardless of the kink. the patient reportedly remained stable throughout the case. the patient remained on impella cp support and was successfully weaned. there was no patient harm reported and the patient survived the explant. | [5 empty] | A.5 ethnicity is unknown. the investigation into the low pump flow issue has been completed. the impella pump was returned for investigation. the visual inspection of the returned product revealed a kink in the cannula. the low pump flow or suction or non-pulsatile motor current from a kinked cannula that resulted from improper pump positioning. this report is being filed as part of a retrospective review of historical records.
18234 | 1220648-2024-19546 | 20379972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379972&pc=OZD | 06/03/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 06/05/2023 | 10/06/2024 |  |  | 03/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024335858 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient if unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported there were low pump flows with the impella cp. the pump did start initially. there was a reduction of flow and performance after approximately thirty seconds to one minute. the decision was made to exchange impella cp with another impella cp. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. data analysis: data confirmed the low flow when pump started & remained to the rest of the case that triggered auto suction response & suction alarms. device analysis: visual inspection found no issues on the pump. pump ran without issue under bench test. the root cause of the low flow was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
18235 | 1220648-2024-19547 | 20379973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379973&pc=OZD | 06/11/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 06/11/2023 | 10/06/2024 |  |  | 02/28/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342004 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there were low pump flows with the impella cp. an acute suction event occurred and was not able to be resolved with decreasing performance levels and repositioning of the device. there was concern for clot. the physician replaced the device by out wiring the old and deploying brand new device in the same access site. no patient harm reported due to the issue. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. pump was not returned for investigation. data log shows the drop in pump flow after motor current spike and suction event. the cause of the low pump flow issue was ingested biomaterial based on data logs and clinical information. this report is being filed as part of a retrospective review of historical records.
18236 | 1220648-2024-19548 | 20379974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379974&pc=OZD | 09/18/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 09/18/2021 | 10/06/2024 |  |  | 09/30/2022 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1504947 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | During the preparation of the catheter, the medical professional was unable to pass the guidewire through the red cheater. when the effort was made, the cheater folded and kinked the wire. a new impella device was used and placed without issue. there was no direct patient involvement. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a visual inspection of the returned product revealed the red lumen folded in on itself within the end of the cannula proximal to the motor housing. significant kinking was noted along the red lumen. the guidewire was not returned. in conclusion, it was determined the cause of the root cause was improper insertion technique when using the old red lumen. this report is being filed as part of a retrospective review of historical records.
18237 | 1220648-2024-19549 | 20379975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379975&pc=OZD | 06/10/2023 | 02/27/2025 |  | 10/06/2024 |  |  | 06/11/2023 | 10/06/2024 |  |  | 01/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024334627 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. it was reported that the impella cp had high purge pressure and motor stopped. a new impella cp was successfully inserted. no patient harm was reported due to the issue, | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | Product was returned for investigation. no physical abnormalities observed. during the first test, high purge pressure was observed and pump was soaked in warm water for some time. pump ran successfully during the second startup without any high purge pressure event. data log did not reported any high purge pressure event. there was no issue found during the case based on the returned log and product. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
18238 | 1220648-2024-19550 | 20379976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379976&pc=OZD | 08/17/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 08/17/2021 | 10/06/2024 |  |  | 03/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022065689 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. reportedly, a fluoroscopy revealed a kink in the cannula. the physician unsuccessfully attempted to reposition the impella. there was no patient harm reported and the device was removed and replaced with a new impella cp. | [5 empty] | A.5 is unknown. the investigation into the product damage issue has been completed. the impella pump was returned for investigation. the field log was reviewed and low pump flows throughout the case were confirmed. visual inspection of the returned pump confirmed a cannula kink near the motor housing. although the root cause of the low pump flow was determined to be the observed cannula kink, a specific cause for the observed cannula kink could not be determined. this report is being filed as part of a retrospective review of historical records.
18239 | 1220648-2024-19551 | 20379977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379977&pc=OZD | 02/24/2023 | 03/10/2025 |  | 10/06/2024 |  |  | 02/24/2023 | 10/06/2024 |  |  | 10/13/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024287037 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp had an optical sensor failure at setup. they replaced the automated impella controller and found this did not remedy the reported issue. ultimately, the cardiac surgeon decided to replace the pump with another impella cp. | [5 empty] | The investigation for the reported optical signal issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records. | The impella device investigation is underway. upon completion of our analysis, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records.
18240 | 1220648-2024-19552 | 20379978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379978&pc=OZD | 09/25/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 09/25/2021 | 10/06/2024 |  |  | 07/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110838 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted an impella cp device for mechanical circulatory support. the clinical team reported that the pump flow was only at approximately 1 l/min despite the pump being in proper position. it was stated that there was evidence of flow saturation and the impella appeared slightly kinked under fluoroscopy. the clinicians suspected was due to the patient's tortuous aortic anatomy. due to unresolved low flow, the pump was explanted and replaced with a second impella cp. no patient harm or injury noted. | [5 empty] | The investigation for the reported product damage (cannula kink) issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the low pump flow was most likely cannula kinking due to the patient's tortuous vasculature. this report is being filed as part of a retrospective review of historical records.
18241 | 1220648-2024-19553 | 20379979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379979&pc=OZD | 02/26/2023 | 09/24/2025 |  | 10/06/2024 |  |  | 02/26/2023 | 10/06/2024 |  |  | 07/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253656 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited continuous placement signal low alarms throughout impella support. several attempts were made to reposition the device despite multiple attempts. the resolution for this issue is unknown. the device was normally explanted and it was reported that the patient survived the procedure. | [5 empty] | A.5 is unknown. the impella device investigation is underway. upon completion of our analysis, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs showed that the "placement signal low" alarm persisted throughout support. however, there were no abnormal placement signal metrics and all 5s parameters were within specification. visual inspection of the returned pump revealed a significant catheter kink proximal to the kink protector. the optical sensor 5s parameters were out of specification. optical continuity testing revealed optical fiber damage at the kink. however, based on the logs, the catheter kink was determined to be unrelated to the optical issue. in conclusion, it was determined that the cause of the optical signal issue was patient condition related. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b.7. updated. d.7a. updated. d.10. updated. h.1., h.2., h.3., h.6., and h.10. updated.
18242 | 1220648-2024-19554 | 20379982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379982&pc=OZD | 09/26/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 09/26/2021 | 10/06/2024 |  |  | 05/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089752 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that after the impella started, flows were observed to be low, and a suction alarm was received. fluoroscopy identified a bend in the impella cannula below the outlet. the initial pump was explanted, and a new device was placed to manage this complication. it was noted that no damage to the cannula was noted after inspection following removal of the impella device. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported product damage (cannula kink) issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the root cause of the suction and low flow was a cannula kink. the root cause of the cannula kink was most likely due to the patient's calcified aortic stenosis (as). this report is being filed as part of a retrospective review of historical records.
18243 | 1220648-2024-19555 | 20379983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379983&pc=OZD | 11/16/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 11/16/2021 | 10/06/2024 |  |  | 09/30/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2022133301 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during an emergency case the nurse started the impella too early as the guidewire was still in place. the impella was removed and a new impella was successfully placed. there was no patient harm reported. | [5 empty] | The investigation into the pump did not start issue has been completed. the impella pump was returned for investigation. the clinical team reported that the impella did not start during the emergent placement of the pump. it was admitted that the impella was activated prematurely after the pump was implanted while the guidewire was still in place. visual inspection of the returned pump revealed a long piece of guidewire coating remaining in the cannula, indicating clear evidence of premature pump activation. no other abnormalities were found. therefore, the cause of the case start issue was improper technique by the end user as the pump was started prematurely before guidewire removal. this report is being filed as part of a retrospective review of historical records.
18244 | 1220648-2024-19556 | 20379985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379985&pc=OZD | 04/07/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 04/07/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311574 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella access site. the patient was transfused with multiple blood products and the impella cp was explanted. the patient survived the procedure. | [5 empty] | The investigation has been completed. no product was returned. as per clinical information, the impella sheath was inserted and the introducer sheath was peeled away. the access site started to bleed around the sheath and then the device was explanted and replaced with a new pump. the cause of the access site bleeding issue was not determined since the product was not returned and with insufficient clinical information. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18245 | 1220648-2024-19557 | 20379988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379988&pc=OZD | 04/22/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 04/22/2023 | 10/06/2024 |  |  | 04/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221559 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, pump had continuous suction alarms indicating low pump flow. device repositioning did not resolve the issue. the pump was removed and replaced with a new impella cp device. no harm to the patient was reported. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the issue was patient condition related. this report is being filed as part of a retrospective review of historical records.
18246 | 1220648-2024-19558 | 20379989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379989&pc=OZD | 04/26/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 04/26/2023 | 10/06/2024 |  |  | 12/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024302870 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed oozing from the access site coinciding with fluid collection around the groin. manual pressure was applied and the pump was subsequently explanted as a result of the noted condition. an impella 5.5 pump was then implanted via axillary access for ongoing support. it was noted the patient survived the procedure. | [5 empty] | The investigation for the reported access site bleeding issues has been completed. the impella device has not been returned for evaluation. the cause of the issue was not determined since product was not returned. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18247 | 1220648-2024-19559 | 20379991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379991&pc=OZD | 04/27/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 04/27/2023 | 10/06/2024 |  |  | 01/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306655 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the impella cp was unable to deliver the appropriate flow. the pump was explanted and replaced. the new impella cp was implanted without complication, and immediately achieved expected flow and performance. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has not been returned for evaluation. the cause of the issue was not determined since product was not returned. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18248 | 1220648-2024-19561 | 20379999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379999&pc=OZD | 04/27/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 04/27/2023 | 10/06/2024 |  |  | 12/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024318137 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that due to the patient's anatomy the impella was unable to cross the aortic valve. the physician made the decision to abort impella support. there was no patient harm the patient was reported to be stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. the investigator reported that due to the patient anatomy, the impella didn't cross the aortic valve. the cause of the issue was patient condition related to diseased valve. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluation" this report is being filed as part of a retrospective review of historical records.
18249 | 1220648-2024-19560 | 20380000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380000&pc=OZD | 04/27/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 04/27/2023 | 10/06/2024 |  |  | 02/28/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320373 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella was placed and when started the flow were lower than expected, despite repositioning the device the flows remained below 2 liters per minute. the physician made the decision to replace this device with another impella, which reportedly functioned normally. there was no patient harm reported, and this device was replaced. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the low pump flow has been completed. the product was not returned, and no data logs were uploaded for review. based on the clinical data the patient had vessel disease and there were an immediate suction alarm after insertion. the root cause most likely was biomaterial ingestion. this report is being filed as part of a retrospective review of historical records.
18250 | 1220648-2024-19562 | 20380001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380001&pc=OZD | 04/28/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 04/28/2023 | 10/06/2024 |  |  | 09/09/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281134 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the physician requested replacement of the impella due to wire and catheter kinked and was unable to be implanted. unable to pass impella 018 wire through guide. impella was replaced and a new impella was successfully placed. there was no patient harm reported. | [5 empty] | The investigation into the product damage has been completed. the impella pump was returned for investigation. the returned pump was investigated and no damage to the pump was observed. the guidewire was not returned for investigation. a 0.18" guidewire from the lab was successfully inserted into the pigtail and exited the outflow cage. a kink in the red lumen was observed near the outflow cage. no other abnormalities were observed. the cause of the product damage was use related due to improper technique. this report is being filed as part of a retrospective review of historical records.
18251 | 1220648-2024-19563 | 20380002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380002&pc=OZD | 03/02/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 03/02/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024293876 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient, demographics unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient had prior stenting of the iliac artery and after the 14 french peel away was inserted, the physician was unable to advance the pump. the physician aborted placement of the impella as they were unable to advance the pump due to the patient¿s peripheral anatomy. reportedly an intra-aortic balloon pump was placed. there was no patient harm reported, and placement of the impella was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. per clinical review the investigator reported that due to the patient anatomy, the impella didn't cross the aortic valve. the cause of the issue was patient condition related to diseased valve. this report is being filed as part of a retrospective review of historical records.
18252 | 1220648-2024-19564 | 20380003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380003&pc=OZD | 05/01/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/01/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312025 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the impella was placed and when started on auto the flows were low at 0.6 liters per minute (lpm) with constant suction. the pressure was increased to none however the flow would still not increase above 1.2 lpm with constant suction. the clinician spoke with a staff member and was told they had given volume the clinician advised the staff person to decrease the pressure level without suction breaking and that the problem could be a thrombus since flows never increased since startup. impella was removed and a new impella placed. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the low pump flow has been completed. the cause of the issue was not determined since the product was not returned. this report is being filed as part of a retrospective review of historical records.
18253 | 1220648-2024-19565 | 20380004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380004&pc=OZD | 05/01/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/02/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298238 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a retroperitoneal bleed, which caused a drop in hemoglobin. the patient received one unit of packed red blood cells. no further information regarding the bleed or treatment was available. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the access site bleeding has been completed. per clinical details the root cause of the retroperitoneal bleed was most likely patient condition and the relationship with impella is unknown. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18254 | 1220648-2024-19566 | 20380005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380005&pc=OZD | 05/01/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/01/2023 | 10/06/2024 |  |  | 12/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322705 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during impella placement that resistance was met at the aortic bifurcation. the physician unsuccessfully attempted to advance using a dilater. the impella placement was aborted due to the patient¿s anatomy as the physician was unable to advance the impella for placement. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. the root cause of the delivery issue was patient condition related, because the patient's vessels were severely calcified and had peripheral artery disease. impella cp with smart assist system instructions for use for the related event are as follows: this report is being filed as part of a retrospective review of historical records.
18255 | 1220648-2024-19568 | 20380006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380006&pc=OZD | 05/02/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/02/2023 | 10/06/2024 |  |  | 01/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311083 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during placement of the impella the fellow pulled the impella back into the aorta and they were unable to recross the aortic valve. this device was removed and a new impella was successfully placed. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the positioning issue has been completed. per clinical review the cause of delivery/positioning issue was due to improper use by the end user. the cause of the issue was wireless re-access. the impella device is not indicated for wireless re-access. the root cause of the issue was improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
18256 | 1220648-2024-19567 | 20380007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380007&pc=OZD | 03/03/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 03/03/2023 | 10/06/2024 |  |  | 09/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278509 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the impella was successfully placed however a kink was noted in the ascending aorta causing suction alarms and low flow of 1.5 liters per minutes. reportedly the impella was adjusted under fluoroscopy, but the kink remained. the physician attempted placement with a second impella and encountered the same kinking issue at the same location. the decision was made to abort impella placement and proceed with the procedure. there was no patient harm reported. | [5 empty] | The investigation into the delivery and product damage issues has been completed. the impella pump was returned for investigation. the returned complaint pump was visually inspected. a cannula kink was observed below the outlet. the root cause of the issue was patient related as the patient had a small aortic arch (aortic anatomy) that kinked the cannula and caused the low pump flow and delivery issues. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluation" this report is being filed as part of a retrospective review of historical records.
18257 | 1220648-2024-19569 | 20380008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380008&pc=OZD | 05/02/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/02/2023 | 10/06/2024 |  |  | 10/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024294435 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the sheath was split in the laboratory and was covered with tegaderm to maintain sterility. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the product damage has been completed. the cause of the issue was not determined since the product was not returned. this report is being filed as part of a retrospective review of historical records.
18258 | 1220648-2024-19570 | 20380015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380015&pc=OZD | 01/30/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 01/30/2023 | 10/06/2024 |  |  | 10/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285845 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the touhy burst lock was not tight and was disconnected from the sheath. reportedly attempts to tighten the lock were unsuccessful, however the device measured at 89 centimeters, which is baseline. the impella was successfully placed, and the procedure completed without complications. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the introducer damage has been completed. the cause of the issue was not determined since the product was not returned. this report is being filed as part of a retrospective review of historical records.
18259 | 1220648-2024-19571 | 20380016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380016&pc=OZD | 05/03/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/03/2023 | 10/06/2024 |  |  | 01/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2024303911 |  | 0048-0014 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  |  | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during impella placement the physician was unable to pull the 0.18 guide wire back across the aortic valve to start the pump. the physician made the decision to replace this impella and place a new impella. there was no patient harm reported, and this device was replaced. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. the cause of the issue was not determined since the product was not returned and clinical details provided were insufficient to determine a root cause. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18260 | 1220648-2024-19572 | 20380017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380017&pc=OZD | 11/17/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 11/17/2021 | 10/06/2024 |  |  | 07/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2022109307 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 20 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, race and ethnicity unknown, in heart failure was implanted with an impella cp device for mechanical circulatory support. it was reported the pump stopped overnight and was restarted. the failure was suspected to be due to a clot, the impella had been in use for 20 days. the decision was made to replace this impella. there was no patient harm reported, and this device was replaced. | [5 empty] | The investigation into the pump stop has been completed. the impella pump was returned for investigation. the field data logs were reviewed and multiple impella stopped ¿ motor current high alarms were confirmed on 17nov2021. prior to these events, the pump appears to have functioned normally and provided over 19 to 20 days of support. the returned pump was visually inspected, and a large purge gap was observed. the pump was plugged into a console and would not start. therefore, the root cause of the pump stop was determined to be bearing wear due to the duration of usage. this is due to the pump being used for 19 days, which is beyond the intended design life of 8 days (design trace matrix 0046-9910). impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi & impella rp with smart assist system section: using the automated impella controller with the impella catheter as noted in the impella ifu ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smart assist system catheter as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
18261 | 1220648-2024-19573 | 20380018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380018&pc=OZD | 05/03/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/03/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024313327 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding. the clinician reported upon arrival to the unit that the nurse reported the impella was not well secured, and the groin site had significant oozing. the physician made the decision to remove the impella. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the access site bleeding has been completed. the cause of the issue was determined to most likely be use related based on the clinical description. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18262 | 1220648-2024-19574 | 20380019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380019&pc=OZD | 05/03/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/03/2023 | 10/06/2024 |  |  | 10/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024293483 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while physician attempted to insert impella cp, they were unable to advance due to severe calcification at the bifurcation. the impella was removed. shockwave therapy attempted to break up calcium. they made another attempt to advance the impella with no success. the decision was made to abort impella placement and use intra aortic balloon pump. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was patient condition related due to severe patient calcification at the bifurcation. this report is being filed as part of a retrospective review of historical records.
18263 | 1220648-2024-19576 | 20380020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380020&pc=OZD | 05/08/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/08/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023278956 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient of unknow age and gender in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a decrease in purge flow coinciding with an increased purge pressure during impella support. the pump was removed and replaced to address the issue. no patient harm was reported. | [5 empty] | The investigation for the reported high purge pressure issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was possible. clinical details provided were sufficient to determine a root cause. the cause of the high purge pressure was most likely biomaterial buildup in the motor. this report is being filed as part of a retrospective review of historical records.
18264 | 1220648-2024-19575 | 20380021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380021&pc=OZD | 05/05/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/05/2023 | 10/06/2024 |  |  | 01/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024310127 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support, pump cannot be advanced, successful implantation after changing to new pump. no harm done to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was possible. clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was patient condition due to patient vasculature. this report is being filed as part of a retrospective review of historical records.
18265 | 1220648-2024-19577 | 20380022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380022&pc=OZD | 03/06/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 03/06/2023 | 10/06/2024 |  |  | 11/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024292555 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the purge filter is leaking. no harm done to the patient. no intervention. pump running afterwards with normal motor current and purge flow. | [5 empty] | The investigation for the reported pump purge leak issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device confirmed a luer leak. the cause of the purge leak was sidearm filter damage. this report is being filed as part of a retrospective review of historical records.
18266 | 1220648-2024-19578 | 20380023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380023&pc=OZD | 03/05/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 03/05/2023 | 10/06/2024 |  |  | 10/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290626 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, there was a fluid leak from the purge cassette. in the process of replacing the purge cassette, it was noted that the purge disc holder had been broken off and that the purge disc was punctured. there was no patient harm as a result of the issues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. no clinical details provided were sufficient to determine a root cause. this report is being filed as part of a retrospective review of historical records.
18267 | 1220648-2024-19579 | 20380059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380059&pc=OZD | 11/17/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 11/17/2021 | 10/06/2024 |  |  | 09/30/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022137584 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, ethnicity and race unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the pump did not start, the pump position was confirmed by trans-esophageal echocardiogram which appeared to have no flows and no movement at the inlet cage. the team suspected a possible thrombus on the inlet cage. this pump was removed and replaced with another impella which worked appropriately. there was no patient harm reported, and this device was replaced. | [5 empty] | The investigation into the pump did not start has been completed. the impella pump was returned for investigation. the product was returned, and a clot was found in the cannula. the pump was run in a bucket of water with the clot and the low flow conditions were reproduced. the clot was then removed, and the pump was run again, this time the pump produced flows that were within specification. the root cause of the low flows was determined to be ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
18268 | 1220648-2024-19580 | 20380127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380127&pc=OZD | 11/18/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 11/18/2021 | 10/06/2024 |  |  | 09/30/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022133835 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported upon startup that the impella flows were low at under 2 liters per minute with continuous suction events, despite repositioning under transthoracic echocardiogram. it was recommended to remove the impella due to a suspected thrombus ingestion. the patient had coded, and cardio-pulmonary resuscitation and numerous defibrillations took place for almost two hours during the procedure. this device was removed and replaced with another impella which functioned normally. there was no patient harm reported. | [5 empty] | The investigation into the low pump flow has been completed. the impella pump was returned for investigation. the field log was reviewed and suction with low pump flow occurred upon activation. position changes were confirmed 3 minutes into the run as troubleshooting measures. the returned pump was visually inspected, and biomaterial was observed entraining the impeller. the cause of the low pump flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
18269 | 1220648-2024-19581 | 20380128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380128&pc=OZD | 12/07/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 01/25/2022 | 10/06/2024 |  |  | 01/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022045986 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the cp was explanted, the patient experienced a cerebrovascular accident. actions taken are unknown. | [5 empty] | The investigation into the reported stroke/cva has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine a root cause. the cause of the cva/stroke was determined to be unknown relation to impella since the stroke was due to cerebrovascular event type was bleeding. purge solution type was not reported. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18270 | 1220648-2024-19582 | 20380141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380141&pc=OZD | 11/21/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 11/21/2021 | 10/06/2024 |  |  | 03/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022090616 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported fluoroscopy revealed the impella appeared kinked near the outflow area. reportedly the flows were less than 2 liters per minute. the decision was made to remove this impella and replace it with a new impella. there was no patient harm reported. | [5 empty] | The investigation into the product damage has been completed. the impella pump was returned for investigation. the returned pump was visually inspected, and the cannula kink was confirmed. the investigator determined improper technique by the end user as the impella is not indicated for wireless re-access. this report is being filed as part of a retrospective review of historical records.
18271 | 1220648-2024-19583 | 20380142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380142&pc=OZD | 12/08/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 12/08/2021 | 10/06/2024 |  |  | 02/28/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022046877 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump stopped. two backup controllers were tried, but the device was still unable to restart. the device was explanted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the product was not returned. the root cause of the pump stop was not determined. the pump stop occurred beyond the 8-day impella cp indication for use per design. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18272 | 1220648-2024-19584 | 20380158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380158&pc=OZD | 11/22/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 11/22/2021 | 10/06/2024 |  |  | 07/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022112948 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported flows were not adequate for the scheduled procedure with flows of 1.2 to 1.5 liters per minute in auto mode. reportedly a fluoroscopy revealed a kink below the outflow cage. the physician made the decision to remove the impella after multiple attempts to reposition the device did not resolve the flow issue. there was no patient harm reported. | [5 empty] | The investigation into the product damage has been completed. the impella pump was returned for investigation. the field data logs were reviewed, and low pump flow was confirmed on 22nov2021. visual inspection of the returned pump confirmed a cannula kink adjacent to the pump outflow. the root cause of the low pump flow was determined to be the observed cannula kink adjacent to the outflow. a specific root cause for the cannula kink could not be determined. this report is being filed as part of a retrospective review of historical records.
18273 | 1220648-2024-19585 | 20380172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380172&pc=OZD | 09/21/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 11/24/2021 | 10/06/2024 |  |  | 05/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022089408 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 22 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported the patient developed hemolysis from insertion to removal. the patients activated clotting time was 160 seconds or less. the positioning of the impella was confirmed under fluoroscopy and the hemolysis was reported to have been resolved, although it was not reported how the hemolysis was resolved. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into hemolysis has been completed. the site reported that the hemolysis was resolved when managing the patient at activated clotting time (act) values above 160. the root cause of the reported hemolysis was most likely related to anticoagulation management as the hemolysis resolved when the patient's act was managed above 160 seconds, concluding improper anticoagulation by the end user. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: mpella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
18274 | 1220648-2024-19586 | 20380176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380176&pc=OZD | 12/15/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 12/15/2021 | 10/06/2024 |  |  | 10/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022141094 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. on day 5 of support, a sudden pump stop occurred. the pump was replaced to address the issue. the patient survived the procedure. | [5 empty] | The investigation into the reported pump stop issue has been completed. analysis of the provided data logs confirmed an 'impella stopped - motor current high' alarm was triggered on (b)(6) 2021. evaluation of the returned pump found biomaterial on the impeller. the root cause of the pump stop was most likely ingested biomaterial. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18275 | 1220648-2024-19587 | 20380177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380177&pc=OZD | 12/15/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 12/15/2021 | 10/06/2024 |  |  | 09/30/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022135739 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. it was reported that there was a delivery issue when the cp pump would not track over the .018 guidewire. as a result, the decision was made to explant the device and replace it with a new impella cp device. no patient harm was reported. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the product was returned. the returned pump was inspected, and no abnormalities were noted. the root cause of the delivery issues is most likely patient condition. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18276 | 1220648-2024-19588 | 20380178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380178&pc=OZD | 12/25/2022 | 09/24/2025 |  | 10/06/2024 |  |  | 12/25/2022 | 10/06/2024 |  |  | 09/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023277109 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that continuous suction alarms were received overnight at low pump levels despite adequate preloading and good positioning. it was stated that the issue was likely thrombus related. the staff decided on a pump exchange. a large hematoma was also noted at the insertion site and there was a concern for a vascular injury. the second pump was aborted. the site was closed with deployment of existing perclose and an additional perclose with good hemostatis. the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the investigation of the low pump flow and access site bleeding has been completed. after reviewing the clinical details clot-related stated and large hematoma noted at insertion site and concern for vascular injury. in reviewing the data logs many suction alarms appeared that confirms the clot. the root cause most likely biomaterial. complaint device not returned for more investigation. the root cause of the access site bleeding - major could not be determined as insufficient clinical details were provided. corrections to the initial mfr report: g1 mdr reporting contact email. h6 investigation findings 4251, investigation conclusion 50 added.
18277 | 1220648-2024-19589 | 20380180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380180&pc=OZD | 12/27/2021 | 10/08/2025 |  | 10/06/2024 |  |  | 12/27/2021 | 10/06/2024 |  |  | 09/30/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022143836 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the impella was placed without issue, however, upon return to the cardiac catheterization unit the impella was noted to be under the papillary muscle. reportedly after attempting to reposition the impella under echocardiogram, the pump was found to be in the aorta. after several attempts to reposition under fluoroscopy, the physician was unable to cross the aortic valve and the impella was exchanged for a new impella, which was successfully placed and positioned. while on venoarterial extracorporeal membrane oxygenation the impella flow was decreased to reduce hemolysis. the device was explanted due to the positioning issue and replaced with another impella cp. | Us complainant reported that the patient (pt) was placed on impella cp for hemodynamic support. it was reported that the flow on the va ecmo/impella was adjusted to reduce hemolysis. additionally, it was also reported that the pump was pulled out of the left ventricle (lv) and was unable to advance back across. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. the root cause of the positioning issue was most likely due to improper technique. the cause of the hemolysis issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18278 | 1220648-2024-19590 | 20380181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380181&pc=OZD | 11/29/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 11/29/2021 | 10/06/2024 |  |  |  |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022105844 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella was placed, and the flow rate did not increase. reportedly a piece of tissue was found, sucked in, which was thought to be from a tissue injury in the left ventricle prior to impella placement. this device was removed and another impella placed. | [5 empty] | The investigation into the low pump flow has been completed. the impella pump was returned for investigation. the investigator reviewed the data logs, which revealed low flow and non-pulsatile motor current throughout impella support, accompanied by recurrent "suction" alarms. visual inspection of the returned pump revealed the presence of moldy biomaterial on and around the impeller. the pump was run on auto and "impella flow reduced" alarm occurred, reproducing the failure mode. in conclusion, it was determined that the cause of the low pump flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
18279 | 1220648-2024-19591 | 20380182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380182&pc=OZD | 01/02/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 02/01/2022 | 10/06/2024 |  |  | 06/30/2022 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1494424 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. three days after impella removal, the patient endured an access site infection, which was treated with administering antibiotics. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist instructions for use for the related event are as follows: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18280 | 1220648-2024-19592 | 20380183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380183&pc=OZD | 05/16/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/16/2023 | 10/06/2024 |  |  | 06/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235265 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a hematoma at the arteriotomy site of the impella device when the patient was moved to bed rest. it was reported pressure was held and femostop placed and hemostasis was achieved after four hours, until the patient moved again resulting in another hematoma developing. the physician made the decision to remove this impella due to concerns of continued bleeding. | [5 empty] | The investigation into the access site bleeding/hematoma has been completed. the product was not returned for review. based on the clinical description, the root cause of the access site bleeding was most likely due to patient movement during support. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18281 | 1220648-2024-19593 | 20380186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380186&pc=OZD | 05/26/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/26/2023 | 10/06/2024 |  |  | 07/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022115153 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the physician was unable to advance the impella past the left ventricle due to the patient¿s anatomy. the physician reportedly made the decision to remove this impella and was able to successfully place a new impella. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. per review of the clinical description, the root cause of the delivery issue was most likely due to the patient¿s anatomy. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18282 | 1220648-2024-19594 | 20380188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380188&pc=OZD | 05/17/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/17/2023 | 10/06/2024 |  |  | 07/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253735 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella sheath was placed in the left femoral artery and a lesion was noted at the bifurcation of the iliac. the physician performed an iliac angioplasty with balloon, however the impella was still unable to pass the lesion. the procedure was aborted due to the patient¿s anatomy and the device removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18283 | 1220648-2024-19595 | 20380189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380189&pc=OZD | 12/01/2021 | 10/06/2024 |  | 10/06/2024 |  |  | 12/01/2021 | 10/06/2024 |  |  | 07/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2022123386 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the purge system failed and the impella pump stopped, the device was explanted. there was no patient harm reported. | [5 empty] | The investigation into the pump stop issue has been completed. the impella pump was returned for investigation. the data logs were not returned. the pump was returned and would not start in the laboratory. the visual inspection noted the catheter was cut open and revealed blood in the purge line. the root cause of the pump stop was due to blood ingress. however, as no kink was found and logs were not returned, the cause of the blood ingress is not known. the root cause of the pump stop was blood ingress into the motor. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi & impella rp with smart assist system section: using the automated impella controller with the impella catheter as noted in the impella ifu ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smart assist system catheter as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
18284 | 1220648-2024-19597 | 20380190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380190&pc=OZD | 12/20/2022 | 10/08/2025 |  | 10/06/2024 |  |  | 12/20/2022 | 10/06/2024 |  |  | 09/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273339 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that upon advancement of the impella device, suction alarms were received with the pump flows decreasing to 1.2 liters per minute with a pump speed level of p2. the positioning was altered numerous times with no resolution to the suction issue. additionally, escalation of care was considered. however, the patient¿s atrium began to tear causing bleeding issues. the impella was explanted. the chest was closed and the patient was taken to the unit to implant a new device. the patient is stable. | [5 empty] | The impella pump, purge cassette and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for low pump flow has been completed. no biomaterial was observed on the returned pump. no cannula kinks, impeller damage or foreign material was observed. no other abnormalities were observed. the rt log shows low pump flow and suction following the cabg procedure. the next rt log (the last rt log in the case) shows continued suction alarms and low flow despite the device being repositioned. the cause of the low pump flow is most likely biomaterial based on the provided clinical information.
18285 | 1220648-2024-19596 | 20380191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380191&pc=OZD | 05/18/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/18/2023 | 10/06/2024 |  |  | 12/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303554 |  | 1000080 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing transcatheter aortic valve replacement was to have an impella cp device implanted for mechanical circulatory support. it was reported the impella was placed and observed to be in the aorta in the catheterization laboratory. the patient was taken to the intensive care unit with the catheter malpositioned. reportedly interventional cardiology made the decision to return the patient to the catheterization laboratory to remove and replace the impella. there was no patient harm reported, and this device was replaced. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the positioning issue has been completed. the data review indicated about 25 minutes into support, the flow decreased gradually from 3.8 liters per minute (l/min) to 2 l/min which triggered auto suction response and suction alarms. approximately 1 hour and 37 minutes later, the pump was in improper position which triggered an alarm impella position unknown and impella position in aorta. the pump was then manually stopped and disconnected from the console. per clinical description and data analysis, the root cause of positioning issue was most likely due to use issue. this report is being filed as part of a retrospective review of historical records.
18286 | 1220648-2024-19598 | 20380198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380198&pc=OZD | 12/18/2022 | 04/11/2025 |  | 10/06/2024 |  |  | 12/18/2022 | 10/06/2024 |  |  | 07/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238445 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella device was unable to maintain a consistent flow and suction alarms were received. the pump placement was confirmed to be in the correct position. however, flows continually decreased. no kinks were identified and the purge cassette was changed with no resolution. the physician decided to explant the initial pump and replace it with a new impella cp. | [5 empty] | A.5 is unknown. the impella pump and purge cassette were returned and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. the returned product was visually inspected and biomaterial was observed on the impeller. the root cause of the issue was biomaterial.
18287 | 1220648-2024-19599 | 20380200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380200&pc=OZD | 02/09/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 02/13/2023 | 10/06/2024 |  |  | 05/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223810 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Pumping Stopped | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was being placed on impella cp for hemodynamic support. it was reported that the physician reported relevant hemolysis during the first 24 hours on impella support. the patient remained on support and the treatment is unknown. it was also reported that there was a pump stop. the pump was explanted and the patient survived the procedure. | [5 empty] | The investigation has been completed. no product was returned. the data logs were reviewed and show suction alarms within the first 24 hours. the suction alarms were resolved when the p-level was lowered. because the hemolysis did not persist past 24 hours it is unlikely to be a malfunction of the product. the root cause of the hemolysis is most likely patient condition. the data logs show multiple motor current spikes before the high motor current pump stop on the 26th day of support. the root cause of the pump stop is unable to be determined as the pump was not returned for analysis. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18288 | 1220648-2024-19600 | 20380202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380202&pc=OZD | 05/18/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/18/2023 | 10/06/2024 |  |  | 12/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024325037 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, race unknown, noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella was placed and when started the flows were low at 1.0 liters per minute. the decision was made to remove this impella and replace it with a new impella, which was successfully placed. there was no patient harm reported. | [5 empty] | The investigation into the low pump flow issue has been completed. the impella pump was returned for investigation. the returned pump was visually inspected, a cannula kink was observed near the outflow cage, which obstructed the impella flow path. no other abnormality were found. the root cause of the low pump flow was a cannula kink due to improper technique used while inserting the impella. this report is being filed as part of a retrospective review of historical records.
18289 | 1220648-2024-19601 | 20380208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380208&pc=OZD | 11/19/2022 | 04/16/2025 |  | 10/06/2024 |  |  | 01/05/2023 | 10/06/2024 |  |  | 08/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249165 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 32 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that a broken air-filter occurred after four days after the impella cp was implanted and the purge pressure decreased. the impella cp was explanted and a new impella cp was implanted. no patient harm or injury noted. | [5 empty] | A.5 is unknown. the impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | Based on the clinical account, the cause of the purge leak was sidearm filter damage. a crack in the air filter was found in the returned device. the cause of the sidearm filter damage was most likely due to ipa interaction.
18290 | 1220648-2024-19602 | 20380212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380212&pc=OZD | 12/05/2022 | 10/06/2024 |  | 10/06/2024 |  |  | 12/05/2022 | 10/06/2024 |  |  | 03/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022059754 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient was experiencing a lot of groin bleeding despite angle matching and other best practices. the patients activated clotting time (act) was 192 seconds, purge was being run without heparin until the patients act came down more given the groin bleeding. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation into the access site bleeding has been completed. the product was not returned. the cause of the access site bleeding was high patient activated clotting time, values (192 seconds). impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18291 | 1220648-2024-19603 | 20380215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380215&pc=OZD | 12/21/2022 | 10/06/2024 |  | 10/06/2024 |  |  | 12/21/2022 | 10/06/2024 |  |  | 09/30/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273395 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that low pump flows and consistent suction even at pump speed level p2. the motor current was flat despite repositioning. the device was explanted and replaced with a new pump. the procedural outcome was the patient survived. | [5 empty] | The investigation for the reported issue has been completed. the returned pump was visually inspected and biomaterial was observed on the impeller and in the inflow. the cause of the issue was biomaterial. this report is being filed as part of a retrospective review of historical records.
18292 | 1220648-2024-19604 | 20380228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380228&pc=OZD | 05/18/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/18/2023 | 10/06/2024 |  |  | 12/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177125 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella. it was reported bubbles of blood was seen leaking from the red impella plug. the impella was removed and replaced with a new impella. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the purge leak issue has been completed. the cause of the blood leaking from the red handle was most likely blood ingress due to a hole in the catheter resulting from improper use. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi section: cleaning as noted in the impella ifu ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18293 | 1220648-2024-19605 | 20380385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380385&pc=OZD | 05/18/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 05/18/2023 | 10/06/2024 |  |  | 01/31/2025 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312227 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during setup, prior to placement, the physician noticed that no purge fluid was exiting by the motor housing. the physician did not use this impella and opened a new impella for the procedure. there was no patient harm reported. | [5 empty] | The investigation into the priming issue has been completed. the impella pump was returned for investigation. the field log was reviewed, and sufficient purge flow and pressure were confirmed indicating that the pump was purging properly. the user likely did not wait long enough for purge to exit the purge gap or could not see the fluid exiting prior to disconnecting the device. there was no issue found during the case. the purge functioned and operated to specification. this report is being filed as part of a retrospective review of historical records.
18294 | 1220648-2024-19606 | 20380388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380388&pc=OZD | 04/29/2023 | 10/06/2024 |  | 10/06/2024 |  |  | 04/29/2023 | 10/06/2024 |  |  | 07/31/2024 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023258927 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella stopped. normal troubleshooting attempts to restart the device were unsuccessful. the impella cp was explanted and replaced to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the product was returned, and the pump stop was confirmed. the root cause of the pump stop was bearing wear. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18295 | 1220648-2024-19607 | 20380391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20380391&pc=OZD | 02/07/2022 | 10/06/2024 |  | 10/06/2024 |  |  | 02/07/2022 | 10/06/2024 |  |  | 10/31/2023 |  | 10/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022150458 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during the weaning process for the impella planned removal the pump stopped. there was no patient harm reported, and the patient was successfully weaned. | [5 empty] | The investigation into the pump stop issue has been completed. the impella pump was returned for investigation. data logs showed the pump stopped immediately upon the sudden occurrence of alarm 119. the pump was unable to be restarted. the "impella stopped" alarm reproduced when attempting to run the returned pump. electrical testing was performed and the resistances between all three motor phases were found to be within specification. however, when manipulating the cables, an open line was reproduced. also, the electrical power (v+/v-) was beyond specification at around 5 megaohms and there was an open line in the optical sensor values. additionally, there was a kink on the catheter at the 85-cm mark. when opening the pump handle, the components were found to be loose. in conclusion, it was determined that the cause of the pump stop was an open electrical line in the power cable, most likely due to excessive force on the catheter by the end user. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18296 | 1220648-2024-19367 | 20378826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378826&pc=OZD | 11/14/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 11/14/2022 | 10/05/2024 |  |  | 06/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238849 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. during implantation, the device could not be advanced from the left iliac artery into the aorta. the device was removed with the intention of using a longer sheath for placement, but the pump was found to have biomaterial in the outflow. the team chose to use a new impella device, and there was no reported patient harm. | [5 empty] | The investigation for the reported difficult insertion has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues is most likely due to patient condition. as noted in the impella ifu: difficult to insert/remove pump through introducer: impella 5.5 & cp impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during impella section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18297 | 1220648-2024-19368 | 20378827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378827&pc=OZD | 11/15/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 11/15/2022 | 10/05/2024 |  |  | 03/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204516 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was found there was a kink in the cannula near the outflow, as well as numerous suction alarms. the pump was also found to be occlusive in the limb. the pump was explanted, and the team chose not to replace it. the patient remained stable on medical therapy, and no harm was reported. | [5 empty] | A visual inspection of the pump confirmed the kink in the cannula. kinks in the introducer sheath were also observed. the lt log confirms suction alarms and low pump flow. the images from the rep show the cannula kinked while the patient is on support. the cause of the low pump flow was most likely small heart patient condition related resulting in a kinked cannula. the anatomy of the patient caused the kink in the cannula which then caused the low pump flow.
18298 | 1220648-2024-19369 | 20378874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378874&pc=OZD | 02/24/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 02/24/2022 | 10/05/2024 |  |  | 11/30/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023164207 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the impella cp stopped on day 8 of support. the patient was hemodynamically stable, and the pump was explanted. no patient harm was reported. | [5 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the investigation for the reported pump stop has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the pump stop was bearing wear. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18299 | 1220648-2024-19370 | 20378875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378875&pc=OZD | 02/25/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 02/25/2022 | 10/05/2024 |  |  | 09/30/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144069 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. a case start issue was reported after prepping the device. the console advanced to the placement screen without turning the pump on. a new console was used; however, with same results. a new pump was opened and prepped without issue. no patient harm was reported. | [5 empty] | The investigation for the reported placement signal issue has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. there was no issue found during the case. the pump was never activated by the user. the returned device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
18300 | 1220648-2024-19371 | 20378876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378876&pc=OZD | 02/27/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 02/27/2022 | 10/05/2024 |  |  | 01/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022038404 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, purge side arm cracked. the pump was removed/replaced without any adverse impact to the patient. | [5 empty] | The investigation for the reported purge leak issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the purge leak was due to sidearm bond damage as a result of excessive force. as noted in the impella ifu: impella cp with smart assist system section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18301 | 1220648-2024-19372 | 20378877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378877&pc=OZD | 02/28/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 02/28/2022 | 10/05/2024 |  |  | 11/30/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169131 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the impella pump would not advance past the bifurcation of the iliac. the physician aborted the impella placement after multiple unsuccessful attempts were made to advance the pump into the descending aorta. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the difficult inserting/removing pump through introducer has been completed. the root cause of the delivery issues was determined to be patient condition as the patient had calcified vessels. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18302 | 1220648-2024-19373 | 20378880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378880&pc=OZD | 03/16/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/16/2023 | 10/05/2024 |  |  | 10/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142129 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. during support, the device exhibited high purge pressure and suction. the team replaced the purge cassette and continued with the same pump, and there was no patient injury reported in relation to this issue. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the low pump flow could not be determined. this report is being filed as part of a retrospective review of historical records.
18303 | 1220648-2024-19374 | 20378884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378884&pc=OZD | 03/16/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/16/2023 | 10/05/2024 |  |  | 09/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023282556 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. however, the catheter was too short to pass through the patient's long and tortuous anatomy, and the procedure was aborted. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was patient anatomy, most likely due to patient condition. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18304 | 1220648-2024-19375 | 20378886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378886&pc=OZD | 02/28/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 02/28/2022 | 10/05/2024 |  |  | 11/30/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167031 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during placement while inserting the cannula the inlet broke. the patients¿ arteries were reported to be tortuous. it was reported the cause of the break was due to the patient¿s vascular anatomy and an additional surgeon was required to assist. the device was replaced, and the procedure aborted. there was no patient harm reported. | [5 empty] | The investigation into the cannula pigtail detachment has been completed. the impella pump was returned for investigation. the visual inspection of the returned device showed the inflow cage detached from the end of the metal cage, indicating the bond remained intact. it is likely the cannula had kinked, and the metal of the inflow cage had cut the cannula and allowed the inflow to detach. the root cause was due to excessive force causing a cannula kink. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings, cautions, and precautions ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ section: warnings & cautions: warnings ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records.
18305 | 1220648-2024-19376 | 20378887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378887&pc=OZD | 03/16/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/16/2023 | 10/05/2024 |  |  | 06/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236958 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed a hematoma in the groin. the physician chose to explant the impella and reschedule the procedure. no further information is available. | [5 empty] | The investigation for the reported hematoma has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the hematoma could not be determined. as noted in the impella ifu: access site bleeding: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records.
18306 | 1220648-2024-19377 | 20378888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378888&pc=OZD | 02/28/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 02/28/2022 | 10/05/2024 |  |  | 10/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152036 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported low flows were present throughout the therapy. the patient remained on impella support and was successfully weaned. no patient harm was reported. | [5 empty] | The investigation into the low pump flow has been completed. the impella pump was returned for investigation. the returned product was visually inspected, and a cannula kink was observed. the cause of the failure mode was a kinked cannula due to the patient¿s condition. this report is being filed as part of a retrospective review of historical records.
18307 | 1220648-2024-19378 | 20378889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378889&pc=OZD | 03/17/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/17/2023 | 10/05/2024 |  |  | 10/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024288276 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. during implantation, the long impella sheath was placed but was unable to advance fully. the team then attempted to place the pump multiple times, but it was unable to move past the calcified iliac artery. the procedure was aborted, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was patient anatomy, most likely due to patient condition. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18308 | 1220648-2024-19379 | 20378890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378890&pc=OZD | 02/12/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 02/12/2023 | 10/05/2024 |  |  | 11/30/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2023161811 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed a hematoma, the resolution for this issue is unknown. it was reported that the patient was weaned off the impella normally and survived the procedure. | [5 empty] | The investigation for the reported hematoma issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding issue was most likely patient condition related. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18309 | 1220648-2024-19380 | 20378891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378891&pc=OZD | 03/19/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/19/2023 | 10/05/2024 |  |  | 10/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290100 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the device suddenly dropped in flows to 0.1-0.4 l/min at setting p5, and suction alarms triggered. the pump was explanted, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the low pump flow could not be determined. this report is being filed as part of a retrospective review of historical records.
18310 | 1220648-2024-19382 | 20378892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378892&pc=OZD | 11/19/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 11/19/2022 | 10/05/2024 |  |  | 08/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262868 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced an access site bleed coinciding with support. heparin was discontinued. the patient was weaned off the device and the device was normally explanted. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported access site bleeding issues has been completed. the impella device has not been returned for evaluation. the root cause of the access site bleeding was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18311 | 1220648-2024-19381 | 20378893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378893&pc=OZD | 03/08/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 03/08/2022 | 10/05/2024 |  |  | 12/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2023171985 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support and after four days of use, the device exhibited an impella stopped alarm. the device stopped, the motor current spiked, then went flat. an attempt was made to resolve this by replacing the controller, however this was unsuccessful. the physician decided to explant the device. it was reported that the patient was hemodynamically stable, did not require an additional device, and the patient survived the procedure. | [5 empty] | The investigation for the reported pump stop has been completed. the product was returned and the problem was confirmed. during testing, the device ran for 5 minutes after activation and the pump stopped. per the results of the clinical review and investigation: the root cause of the pump stop was determined to be bearing wear. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18312 | 1220648-2024-19384 | 20378894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378894&pc=OZD | 03/19/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 03/19/2022 | 10/05/2024 |  |  | 07/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022116506 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported bubbles were noted to the outflow under trans-esophageal echocardiogram approximately three hours after placement. the impella cp was removed, and patient was escalated from an impella cp to an impella 5.5. | [5 empty] | The investigation into the embolism has been completed. the impella pump was returned for investigation. the returned pump and returned purge cassette were run through de-air, placed in a basin of water, and no air bubbles were produced from the purge gap. purge flow was within specification. the cause of the embolism was most likely use related since it was unable to be reproduced on returned product. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smartassist catheter insertion. section: axillary insertion of the impella cp with smartassist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". this report is being filed as part of a retrospective review of historical records.
18313 | 1220648-2024-19383 | 20378895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378895&pc=OZD | 03/21/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/21/2023 | 10/05/2024 |  |  | 01/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024308953 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. due to the anatomy of the patient, the cannula kinked and produced greatly reduced pump flows. no intervention was specified, but there was no reported patient harm. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. returned fluoro imaging showed a kinked cannula. therefore, the root cause of the low pump flow was a cannula kink as a result of patient condition. this report is being filed as part of a retrospective review of historical records.
18314 | 1220648-2024-19385 | 20378898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378898&pc=OZD | 03/21/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/21/2023 | 10/05/2024 |  |  | 11/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309985 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. after multiple attempts, the device was unable to be delivered due to patient anatomy and severe left ventricular hypertrophy. the procedure was aborted, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy/condition. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18315 | 1220648-2024-19386 | 20378908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378908&pc=OZD | 03/21/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 03/21/2022 | 10/05/2024 |  |  | 10/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157114 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a pump stop and multiple attempts to restart the pump failed. the impella was immediately removed from the left ventricle into the descending aorta and the patient reported to remain hemodynamically stable. reportedly the operating room was backed up and the pump was left in until the next day, leaving the pump in the descending aorta. the nurse was questioned by the clinician if they had any concerns about leaving a pump that was not functioning in place that long and she stated that the physician was not concerned as long as patient was systemically anticoagulated. the impella was removed the following day. there was no patient harm reported. | [5 empty] | The investigation into the pump stop has been completed. the impella pump was returned for investigation. the data logs were returned and confirm a high motor current pump stop on day 14 of support. the product was returned, and a large purge gap was observed. the pump stop occurred on day 14 of use which is beyond the indicated design life of 8 days. the root cause of the pump stop was determined to be bearing wear due to the duration of use. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18316 | 1220648-2024-19387 | 20378909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378909&pc=OZD | 03/22/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/22/2023 | 10/05/2024 |  |  | 11/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024308513 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. during insertion, the physician was unable to fully advance the device due to patient anatomy/aortic calcification. the wire had no visible kink when the wire was removed. the procedure was aborted, and there was no reported harm to the patient. | [5 empty] | The device was returned, and an investigation into the reported delivery issue has been completed. the returned pump was visually inspected, and a kink was observed in the cannula. a pigtail kink was observed. a guidewire was attempted to be inserted but resistance was felt at the pigtail kink. the cause of the delivery issue was a tight aortic arch patient condition related that resulted in a cannula kink. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18317 | 1220648-2024-19388 | 20378910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378910&pc=OZD | 03/18/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 03/18/2022 | 10/05/2024 |  |  | 11/30/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023167263 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the impella was showing signs of saturated flow, indicating the possibility of failure which could result in the pump stopping. the pump was replaced with a new impella cp. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the saturated flow has been completed. the pump was used past the indicated duration of use, however without the returned product the cause of the issue could not be determined. the cause of the issue was not determined since the product was not returned. this report is being filed as part of a retrospective review of historical records.
18318 | 1220648-2024-19389 | 20378911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378911&pc=OZD | 03/22/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/22/2023 | 10/05/2024 |  |  | 11/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303412 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed a hematoma at the access site. the device was removed and replaced with another impella cp with a larger sheath. the patient survived the procedure. | [5 empty] | The investigation for the reported hematoma has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the hematoma could not be determined. as noted in the impella ifu: hematoma: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18319 | 1220648-2024-19390 | 20378912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378912&pc=OZD | 03/17/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 03/22/2022 | 10/05/2024 |  |  | 12/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162334 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient¿s urine was wine red in color indicating hemolysis. reportedly, an echocardiography confirmed the positioning of the impella was acceptable at 35 millimeters. it was reported the pressure was lowered and hemolysis improved upon weaning and removal of the impella. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into hemolysis has been completed. the cause of the issue was not determined since the product was not returned and clinical details provided were insufficient to determine a root cause. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions, including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿ section: suction: as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ the clinical details provided were insufficient to determine a root cause.
18320 | 1220648-2024-19391 | 20378913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378913&pc=OZD | 03/22/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 03/22/2022 | 10/05/2024 |  |  | 12/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023175398 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. reportedly, a clot was ingested during placement of the impella and when the impella was removed, it came back across the aortic valve (av) and had to be reinserted. the physician chose to leave the impella in place regardless of the maximum mean flow being 2.0 liters per minute. there was no patient harm reported. | [5 empty] | The investigation into the low pump flow has been completed. the impella pump was returned for investigation. the data logs confirm a period of low flow after the pump was reinserted. the suction condition improved after about 20 minutes. the product was returned and ingested biomaterial was found in the outflow cage area. the root cause of the low pump flow was most likely biomaterial picked up on the redelivery of the pump. the root cause of the issue was biomaterial found in the outlet. this report is being filed as part of a retrospective review of historical records.
18321 | 1220648-2024-19392 | 20378915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378915&pc=OZD | 03/23/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 03/23/2022 | 10/05/2024 |  |  | 01/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174133 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump stopped after 18 days of support. multiple unsuccessful attempts were made to restart the pump. the decision was made to remove the pump. there was no patient harm reported. | [5 empty] | The investigation into the pump stop has been completed. the impella pump was returned for investigation. the data logs showed the pump stopped immediately upon the sudden occurrence of the "impella stopped - motor current high" alarm. the impella was returned with most of the catheter cut off. a large impeller purge gap was observed on the returned product. in conclusion, it was determined that the cause of the pump stop was bearing wear due to duration of 18 days of use, which is beyond design life. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18322 | 1220648-2024-19393 | 20378916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378916&pc=OZD | 12/06/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 12/06/2022 | 10/05/2024 |  |  | 08/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268322 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic |  | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, an access site bleed that occurred after advancing the repositioning sheath. the pump was ultimately removed and replaced with a new impella in the contralateral groin. | [5 empty] | The investigation for access site bleeding has been completed. the device was not returned for evaluation, nor were any clinical notes provided, therefore no root cause could be determined. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18323 | 1220648-2024-19394 | 20378917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378917&pc=OZD | 03/23/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 03/23/2022 | 10/05/2024 |  |  | 12/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172071 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump was placed, and low flows were noted with a kinked cannula. patient remained on impella support and subsequently weaned off. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the low pump flow has been completed. per the clinical information it was determined the most likely root cause of the low pump flows was a kinked cannula due to improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
18324 | 1220648-2024-19395 | 20378935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378935&pc=OZD | 03/22/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/22/2023 | 10/05/2024 |  |  | 12/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311233 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient of unknown race/ethnicity noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the insertion was unsuccessful due to patient anatomy. no intervention or resolution were specified, but there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was patient condition. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18325 | 1220648-2024-19396 | 20378937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378937&pc=OZD | 01/27/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/21/2023 | 10/05/2024 |  |  | 04/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023211551 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The eu complainant had a support ongoing with the cp when after just 6 hours of support the console alerted for retrograde flow and stopped. the team troubleshot by the replacement of the aic (automated impella controller). the aic then would not allow for the pump restart and so the pump additionally was replaced. the cp is not a full support pump. there was no harm or injury from the pump stop. the patient is now supported by a new pump and console. | [5 empty] | The root cause of the pump stop cannot be determined since product was not returned. in the first log it shows the ¿high motor current¿ alarms and the pump stopping. in the second log it shows that there were a lot of suction alarms throughout the case but they adjust the position of the impella and flows began to improve. the third log is when they switched to another console and were not able to start it again. there is no indication of biomaterial and the clinical stated that there was no thrombus seen. this pump was used for 303 hours which is beyond design life indication. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
18326 | 1220648-2024-19397 | 20378943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378943&pc=OZD | 12/07/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 12/07/2022 | 10/05/2024 |  |  | 09/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263124 |  | 1000302 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during delivery, the physician removed the impella with a 018 wire (and found that wire was kinked). physician re-accessed the patient with a pigtail and 035 wire, they used a new v18 wire for the exchange. this time the impella couldn't cross the iliac as there was calcium. physician had to remove everything and dilate with abiomed 14f(25cm) dilators, then the physician managed to slowly advance the impella into the left ventricle. it was reported that the patient was stable throughout the procedure. no further information is available. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition related. patient condition due to calcification impella couldn't cross iliac and patient had some aortic stenosis. impella cp with smartassist system instructions for use for the related event are as follows: section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
18327 | 1220648-2024-19398 | 20378951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378951&pc=OZD | 03/24/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/24/2023 | 10/05/2024 |  |  | 12/31/2024 |  | 10/05/2024 |  | IMPELLA CP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024307476 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. approximately 15 minutes into the procedure, the automated impella controller (aic) produced a suction alarm. the device was lowered to p-1 setting, but the alarm continued. the pump was then advanced and withdrawn, but this did not resolve the issue. finally, the patient was given fluid, but the alarm remained. the pump was then explanted and replaced, and the alarm did not continue. the procedure was completed successfully, and no patient harm was reported. | [5 empty] | The investigation for the reported low pump flow has been completed. the product was returned, and the low pump was confirmed. per the results of the clinical review and investigation: the log shows suction alarms at the time of the low pump flow. at the time that the low pump flow starts there is a motor current spike and a temporary drop in placement signal. the returned pump was investigated, and no biomaterial was found around the impeller. no cannula kinks were noted. the pump was purged with sterile water and attempted to be run but had high motor current pump stops when attempting to start the pump. there was also high purge pressure noted. the catheter was opened, and no blood was observed in the purge line near the motor housing. the purge line was cut at the 19cm marking on the catheter and purge fluid then traveled easily through the line. the blockage was isolated to the motor. the motor was torn down, and blood was observed on the magnet. the cause of the low pump flow was biomaterial. this report is being filed as part of a retrospective review of historical records.
18328 | 1220648-2024-19399 | 20378955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378955&pc=OZD | 08/09/2022 | 03/10/2025 |  | 10/05/2024 |  |  | 08/09/2022 | 10/05/2024 |  |  | 09/30/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022143849 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, a pump stop was reported, and the patient experienced supraventricular tachycardia. during the arrhythmia, impella flows dropped to 0.0, briefly to p0, but the device was restarted and returned to p6. a second suction event then occurred spontaneously, with a drop in flows and motor current (from 600s to 400s). the purge fluid was changed to tissue plasminogen activator. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported vf/vt/af/at, low pump flow, pump stop and optical signal issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18329 | 1220648-2024-19400 | 20378956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378956&pc=OZD | 12/08/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 12/07/2022 | 10/05/2024 |  |  | 08/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272273 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during delivery the impella cp kinked. the pump was not able to deliver up the aorta and so it was removed and replaced with another imeplla cp. the second pump was placed without report of harm or injury from the delay in support initiation. no further information is available. | [5 empty] | The investigation for the reported product damage (kinked cannula) issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was most likely patient condition related since the patient had a transcatheter aortic valve replacement and cardio-pulmonary resuscitation was performed throughout. this report is being filed as part of a retrospective review of historical records.
18330 | 1220648-2024-19401 | 20378973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378973&pc=OZD | 12/14/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 12/14/2022 | 10/05/2024 |  |  | 06/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247097 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Migration or Expulsion of Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the repositioning sheath process, the pump came out of position and became kinked. there was also oozing reported at the access site, requiring the physician to discontinue heparin. the pump was explanted and replaced, and there was no reported harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the investigation for the reported positioning issue has been completed. clinical details provided were sufficient to determine a root cause. the root cause of the positioning issue was determined to be use related as the pump was advanced during insertion of the repositioning unit. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18331 | 1220648-2024-19402 | 20378976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378976&pc=OZD | 12/27/2022 | 09/24/2025 |  | 10/05/2024 |  |  | 12/27/2022 | 10/05/2024 |  |  | 08/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269897 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, a pump stop issue with the impella device. upon initiation of support, the motor current was observed to be high and the pump subsequently stopped. the pump was restarted, however once p7 was reached, the pump once again stopped. the pump was again restarted and the pump was able to stabilize at support level p3. also, some blood clotting was noted in impella sheath but was cleared prior to pump insertion. the scheduled procedure was completed without further issue. there was no harm to the patient. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the investigation of the temporary pump stop and thrombus has been completed. the pump was visually inspected and biomaterial was found wrapped around the impeller. the pump was able to start in the lab with slightly higher than normal flows at p9 around 4.3 l/min. log review indicates intermittent pump stops at the beginning of the case, with the pump ultimately able to run with saturated flows. based on this evidence, and the fact the patient had issues with clotting, the cause of the intermittent pump stops was due to biomaterial.
18332 | 1220648-2024-19403 | 20378995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378995&pc=OZD | 04/05/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 04/05/2023 | 10/05/2024 |  |  | 10/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285225 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient was found to have two distinct left ventricular thrombus bundles. the impella was explanted. the patient survived the procedure. | [5 empty] | The investigation for the reported thrombus has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the thrombus was due to patient condition related. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
18333 | 1220648-2024-19404 | 20378997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378997&pc=OZD | 04/06/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 04/06/2023 | 10/05/2024 |  |  | 08/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270227 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, saturated flows were observed. as a result, the decision was made to explant the device and replace it with an impella 5.5. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of saturated flow. the device was not received for evaluation. the root cause of the saturated flow was unable to be determined as no product was returned for analysis. this report is being filed as part of a retrospective review of historical records.
18334 | 1220648-2024-19406 | 20378998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378998&pc=OZD | 04/06/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 04/06/2023 | 10/05/2024 |  |  | 01/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306651 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to being implanted with an impella cp device for mechanical circulatory support. the pump damage at the red lumen/guidewire interaction. the impella was explanted and replaced with a new impella cp to address the issue. no patient harm was reported due to the issue. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. visual inspection found kink on the red lumen about 3.5 inches from the flag. no issue was found on the pigtail. there were no evidences indicate that if the red lumen was kinked out of the box or during using, therefore, the root cause of delivery issue was unable to be determined. this report is being filed as part of a retrospective review of historical records.
18335 | 1220648-2024-19407 | 20378999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378999&pc=OZD | 02/27/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 02/27/2023 | 10/05/2024 |  |  | 09/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273431 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the placement signal was lost. the impella was explanted and no patient harm was reported. | [5 empty] | The investigation into the reported placement signal issue has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine a root cause. the most likely cause of the placement signal issue was patient condition related. this report is being filed as part of a retrospective review of historical records.
18336 | 1220648-2024-19405 | 20379000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379000&pc=OZD | 02/19/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 02/19/2023 | 10/05/2024 |  |  | 05/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2023222830 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported there was leakage from the pressure reservoir of the pump catheter body. reportedly the "purge pressure drop" alarm was continuously sounding until the pump catheter body was replaced. despite lowering the impella support level from pressure eight to pressure four or two, no significant changes in the patient¿s hemodynamics were observed. the decision was made to remove impella and place an intra-aortic balloon pump (iabp). upon removal of the impella, the clinician confirmed the liquid leaking was from the back side of the pressure reservoir, however no damage such as cracks were observed at the relevant location. there was no patient harm reported, and this device was explanted. | [5 empty] | The investigation into the purge leak has been completed. the impella pump was returned for investigation. a video received from the field confirmed the purge leak emanated from the pressure reservoir of the purge sidearm. when reproduction testing was performed on the returned pump, the pressure reservoir leaked. further visual inspection revealed the source of the leak was a puncture hole in the blue diaphragm of the pressure reservoir. therefore, it was determined that the cause of the purge leak was sidearm reservoir damage resulting from improper technique by the end user. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi section: cleaning as noted in the impella ifu ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18337 | 1220648-2024-19408 | 20379001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379001&pc=OZD | 04/07/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 04/07/2023 | 10/05/2024 |  |  | 10/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022148061 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was found the impella was across the valve. the team was unable to reposition appropriately due to underlying valve disease. the impella cp was explanted and care was escalated to impella 5.5. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of positioning issue. the device was not received for evaluation. the root cause of the positioning issue was due to patient movement. this report is being filed as part of a retrospective review of historical records.
18338 | 1220648-2024-19410 | 20379006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379006&pc=OZD | 06/16/2023 | 10/01/2025 |  | 10/05/2024 |  |  | 06/16/2023 | 10/05/2024 |  |  | 04/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226503 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, purge alarms were observed. the impella was explanted and replaced with a new impella cp. no patient harm was reported due to the issue, | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation reported for purge leak-pump has been completed. data confirmed the pump was unable to go through the case start on both consoles ic8466 & ic5825 that triggered alarm 407 (purge system open) & 404 (air in purge system). product was not returned for review. the root cause of the priming issue was most likely use to use issue. g1: mdr reporting contact email updated.
18339 | 1220648-2024-19409 | 20379007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379007&pc=OZD | 04/18/2023 | 05/28/2025 |  | 10/05/2024 |  |  | 04/19/2023 | 10/05/2024 |  |  | 11/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278950 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to being implanted with an impella cp device for mechanical circulatory support. the impella failed to delivery due to a kink. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the delivery issue is most likely due to a kink in the catheter. this report is being filed as part of a retrospective review of historical records.
18340 | 1220648-2024-19411 | 20379009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379009&pc=OZD | 04/23/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 04/23/2023 | 10/05/2024 |  |  | 11/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024294973 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a hematoma developed while in cath lab. manual pressure and the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of access site bleeding. the device was not returned for evaluation. the root cause of the access site bleeding was a lack of vessel recoil onto the sheath due to bleeding at arteriotomy. this report is being filed as part of a retrospective review of historical records.
18341 | 1220648-2024-19412 | 20379010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379010&pc=OZD | 05/22/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/22/2023 | 10/05/2024 |  |  | 01/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024329013 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp. after the initial startup, the pump stopped and could not be restarted. the impella cp was explanted and replaced to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the product was returned. there was no issue found with the pump. the device functioned/operated to specification. the pump stop was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
18342 | 1220648-2024-19413 | 20379011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379011&pc=OZD | 02/19/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 02/19/2023 | 10/05/2024 |  |  | 10/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024288712 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, impella pulled back across the aortic valve (av). attempts to reposition the impella back across the left ventricle (lv) were not successful. the physician made the decision to remove the impella and see if the patient was able to tolerate the removal and was prepared to reinsert the impella if needed. there was no patient harm reported, and this device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the positioning issue has been completed. the cause of the issue was wireless re-access. the impella device is not indicated for wireless re-access resulting in improper use by the end user. this report is being filed as part of a retrospective review of historical records.
18343 | 1220648-2024-19414 | 20379012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379012&pc=OZD | 05/19/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/19/2023 | 10/05/2024 |  |  | 02/28/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024319703 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. upon delivery, it was observed the cannula kinked below the outlet and it resulted in low flows during impella support. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was not received for evaluation. the root cause of the low pump flow was due to a kinked cannula due to improper technique use. this report is being filed as part of a retrospective review of historical records.
18344 | 1220648-2024-19415 | 20379014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379014&pc=OZD | 06/16/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/16/2023 | 10/05/2024 |  |  | 12/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303756 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp support, it was suspected a clot entrained into pump during the patient's ablation procedure. the patient was receiving ventricular tachycardia ablation when the pump went into continuous suction. after assessing the pump, it was suspected that the pump entrained a clot, and the left ventricle waveform and motor current were flat, and flows decreased to 0.5 liters per minute. despite not having adequate flow, fluorodcopic imaging showed that the pump was in the proper position as well as echocardiography imaging. it was decided that the md would insert a new impella pump. patient had to receive an increase in pressors during this time as well. the patient was successfully placed on a new impella cp device. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was not received for evaluation. the root cause of the low pump flow was unable to be determined as no product or logs were returned for analysis. this report is being filed as part of a retrospective review of historical records.
18345 | 1220648-2024-19419 | 20379015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379015&pc=OZD | 05/22/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/22/2023 | 10/05/2024 |  |  | 03/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024329702 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. the impella could not be positioned in the ventricle because of the kinking and stenosis in this area. the insertion was aborted due to unexpected non-visible anatomy in bifurcation. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18346 | 1220648-2024-19421 | 20379016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379016&pc=OZD | 05/07/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/21/2023 | 10/05/2024 |  |  | 01/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316442 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced hematoma at the impella access site. the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of access site bleeding. the device was not received for evaluation. the root cause of the access site bleeding was unable to be determined as no product was not returned. this report is being filed as part of a retrospective review of historical records.
18347 | 1220648-2024-19420 | 20379017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379017&pc=OZD | 05/28/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/28/2023 | 10/05/2024 |  |  | 02/29/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200628 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump stopped during use of the device. the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the device was returned for evaluation and confirmed pump stop. the root cause of the pump stop was determined to be biomaterial. this report is being filed as part of a retrospective review of historical records.
18348 | 1220648-2024-19416 | 20379018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379018&pc=OZD | 05/19/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/19/2023 | 10/05/2024 |  |  | 12/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325705 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with unknow race and ethnicity was attempted to being implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp pump could not be fully advanced into the patient because of challenging anatomy; therefore, the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient condition related to diseased aortic valve. this report is being filed as part of a retrospective review of historical records.
18349 | 1220648-2024-19422 | 20379019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379019&pc=OZD | 05/24/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/24/2023 | 10/05/2024 |  |  | 10/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283392 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to being implanted with an impella cp device for mechanical circulatory support. the impella was advanced, but unable to advance around aortic arch. after several attempts with different wires. the physician stated the patient had a lot of calcium and a porcelain aorta which prevented advancement of impella. the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation into the delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the most likely cause is due to the patient¿s condition. this report is being filed as part of a retrospective review of historical records.
18350 | 1220648-2024-19423 | 20379020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379020&pc=OZD | 05/24/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/24/2023 | 10/05/2024 |  |  | 01/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333715 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella had continuous suction alarms. the physician decided to pull out the impella and place a new sheath to maintain artery access for a coronary artery grafting bypass. the procedure was completed successfully and the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was returned for evaluation and confirmed low pump flow. the root cause of the low pump flow was determined to be biomaterial. this report is being filed as part of a retrospective review of historical records.
18351 | 1220648-2024-19424 | 20379021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379021&pc=OZD | 05/24/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/24/2023 | 10/05/2024 |  |  | 12/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024302294 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with unknown ethnicity noted as high-risk percutaneous coronary intervention was attempted to being implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp pump could not be fully advanced into the patient because of challenging anatomy; therefore, the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation into the delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the most likely cause is due to the patient¿s condition since calcification in vessels was observed. this report is being filed as part of a retrospective review of historical records.
18352 | 1220648-2024-19425 | 20379022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379022&pc=OZD | 05/25/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/26/2023 | 10/05/2024 |  |  | 07/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253570 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, optical signal malfunctioned was observed. the decision was made to explant the device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of optical signal issue. the device was not received for evaluation. the root cause of the optical signal issue was unable to be determined as no product was returned for analysis. this report is being filed as part of a retrospective review of historical records.
18353 | 1220648-2024-19426 | 20379023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379023&pc=OZD | 06/17/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/19/2023 | 10/05/2024 |  |  | 05/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328718 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the impella cp was in good amplitude of the aorta, but the optical signal alarm was observed. action taken to resolve the issue are unknown. no patient harm was reported. | [5 empty] | The investigation has been completed for the reported issue of optical signal issues. the device was not received for evaluation. the root cause of the optical signal issue was most likely patient condition related. this report is being filed as part of a retrospective review of historical records.
18354 | 1220648-2024-19417 | 20379024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379024&pc=OZD | 05/05/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/05/2023 | 10/05/2024 |  |  | 12/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309162 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella access site and angle re-matching and skin re-suture sutures were applied to achieve hemostasis. the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of access site bleeding. the device was not received for evaluation. the root cause of the access site bleeding was unable to be determined as product not returned and limited clinical details were provided. this report is being filed as part of a retrospective review of historical records.
18355 | 1220648-2024-19418 | 20379025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379025&pc=OZD | 05/22/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/22/2023 | 10/05/2024 |  |  | 01/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024332220 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the cannula folded upon itself during repositioning of the impella pump. although the pump was able to reobtain proper positioning, the crease in the cannula was still visible just distal to the outlet/motor housing. due to the noted damage, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18356 | 1220648-2024-19428 | 20379026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379026&pc=OZD | 05/26/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/26/2023 | 10/05/2024 |  |  | 12/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024326303 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. after many attempts of repositioning, the axillary/aorta size was not favorable. the patient continued on support until the device was successfully weaned. no patient harm was reported due to issue . | [5 empty] | The investigation for positioning issue was completed and no device was returned. the cause of the issue was most likely patient condition related. this report is being filed as part of a retrospective review of historical records.
18357 | 1220648-2024-19427 | 20379027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379027&pc=OZD | 05/25/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/25/2023 | 10/05/2024 |  |  | 01/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338667 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to being implanted with an impella cp device for mechanical circulatory support. the impella cp pump could not be fully advanced into the patient because of challenging anatomy; therefore, the insertion was aborted. no patient harm reported due to the issue, | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient condition. this report is being filed as part of a retrospective review of historical records.
18358 | 1220648-2024-19430 | 20379029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379029&pc=OZD | 06/06/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/06/2023 | 10/05/2024 |  |  | 01/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318859 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that that the cardiac surgeon attempted three times to insert the pump in the patient heart; however, was unable to pass the aorta due to severe peripheral artery disease. the cardiac surgeon attempted multiple other times different routes; however, was unable to implant the impella cp due to the patient¿s anatomy and ultimately aborted the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was patient condition related because the impella was unable to pass into aorta due to severe peripheral artery disease (pad). as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18359 | 1220648-2024-19429 | 20379030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379030&pc=OZD | 06/05/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/05/2023 | 10/05/2024 |  |  | 01/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024330956 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the complainant reported the patient had extensive peripheral vascular disease and despite multiple attempts to gain access via the left femoral artery the cardiac surgeon was unable to obtain access. the impella cp implant was aborted due to patient anatomy. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was patient condition related to peripheral artery disease (pad)/peripheral vascular disease (pvd) vessel condition based on provided clinical information. this report is being filed as part of a retrospective review of historical records.
18360 | 1220648-2024-19431 | 20379033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379033&pc=OZD | 06/07/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/07/2023 | 10/05/2024 |  |  | 02/28/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325829 |  | 1000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had blood coming from in between blue suture hub and white connection piece. the device was removed shortly after being implanted. patient is stable post explant. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. product was not returned for review. the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18361 | 1220648-2024-19432 | 20379036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379036&pc=OZD | 06/07/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/07/2023 | 10/05/2024 |  |  | 03/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024340374 |  | 1000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the complainant reported that an impella cp optical pump was unable to advance into the left ventricle due to moderate aortic stenosis and a tight aortic arch. as a result of the patient's anatomy, the decision was made to abort the implant. the complainant reported that an impella cp optical pump was unable to advance into the left ventricle due to moderate aortic stenosis and a tight aortic arch. as a result of the patient's anatomy, the decision was made to abort the implant. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. product was not returned for review. based on clinical description, the root cause of delivery issue was due to patient condition. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18362 | 1220648-2024-19433 | 20379037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379037&pc=OZD | 06/07/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/07/2023 | 10/05/2024 |  |  | 03/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328781 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the cp wire bent in iliac and was unable to pass cp. the doctor got a new 0.018 wire and changed to longer sheath. the longer sheath was still unable to cross. the doctor stated the impella and wire both were bent on both devices. the procedure was aborted. no patient harm was reported due to the issue. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. no product was returned. the cause of the delivery issue was due to patient condition with patient having tortuous vessels leading to bent pump during insertion. the cause of the delivery issue was due to patient condition with patient having tortuous vessels leading to bent guidewire during insertion. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18363 | 1220648-2024-19435 | 20379038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379038&pc=OZD | 06/07/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/07/2023 | 10/05/2024 |  |  | 01/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024332329 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported impella cp delivery was unsuccessful due to the patient having calcified vessels. the impella insertion was aborted. no patient harm reported due to the issue. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause right and left illiac were ballooned therefore the sheath was unable to be advanced fully. the impella was unable to pass the distal aorta due to the calcium burden. root cause of the delivery issue was patient condition related. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18364 | 1220648-2024-19434 | 20379039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379039&pc=OZD | 06/07/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/07/2023 | 10/05/2024 |  |  | 03/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024328779 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that that an impella cp pump was unable to advance beyond the iliacs due to the tortuous vasculature. multiple attempts were made to deliver the pump, however, the issue persisted. it was reported that the patient was bleeding badly with perclose, manual pressure and a femstop used is standard care. the second cp kit that was being prepped appeared to be bent. the decision was subsequently made to abort the implant. patient is stable post issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18365 | 1220648-2024-19437 | 20379040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379040&pc=OZD | 06/08/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/08/2023 | 10/05/2024 |  |  | 05/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231095 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for pre-operation placement. the complainant reported bleeding at illiac dissection. the device was explanted, angioseal and femstop were applied. the patient went for coronary artery bypass graft (cabg) the next day and did fine. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. as per clinical, access site was bleeding due to iliac dissection. the cause of the access site bleeding issue was not determined since product was not returned and with insufficient clinical information. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18366 | 1220648-2024-19436 | 20379041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379041&pc=OZD | 06/07/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/07/2023 | 10/05/2024 |  |  | 11/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024314512 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that during transport the patient developed a moderate size hematoma causing bleeding at the access site due to a broken perclose suture that happened in transport. the physician resolved this by advancing the repositioning sheath and tightening the remaining perclose suture. patient is stable post issue. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the access site bleeding was use issue related due to broken perclose suture that happened in transport. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18367 | 1220648-2024-19438 | 20379043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379043&pc=OZD | 02/20/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/20/2023 | 10/05/2024 |  |  | 09/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273458 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. while on support, impella flows were lower than expected. the physician made the decision to remove the impella and replace it with another impella. the physician was pleased with the position and flows of the newly placed impella. there was no patient harm reported, and this device was replaced with another impella. | [5 empty] | The investigation into the low pump flow has been completed. the impella pump was returned for investigation. the visual inspection of the returned pump noted a cannula kink near the outflow. the investigator determined the root cause of the low flow issue was a cannula kink due to the end users improper use technique. this report is being filed as part of a retrospective review of historical records.
18368 | 1220648-2024-19439 | 20379044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379044&pc=OZD | 07/04/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 07/04/2023 | 10/05/2024 |  |  | 02/28/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321063 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, intermittent suction alarms occurred. the team explanted the pump and observed extensive underlying left ventricular clot. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was not returned for evaluation. the root cause of the low pump flow was determined to be biomaterial found on pump based on clinical information. this report is being filed as part of a retrospective review of historical records.
18369 | 1220648-2024-19440 | 20379045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379045&pc=OZD | 07/05/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 07/05/2023 | 10/05/2024 |  |  | 12/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295272 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. shortly after starting the impella, the positioning waveform disappeared. there was no bending of the connection cable or unusual insertion method. the decision was made to explant the device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of optical signal issue. the device was not received for evaluation. the root cause of the optical signal issue was unable to be determined since product was not returned, data logs were not returned and insufficient clinical information provided. this report is being filed as part of a retrospective review of historical records.
18370 | 1220648-2024-19441 | 20379046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379046&pc=OZD | 07/05/2023 | 03/27/2025 |  | 10/05/2024 |  |  | 07/05/2023 | 10/05/2024 |  |  | 03/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024345267 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the pump experienced a drop in flow rate coinciding with suction alarms. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | Returned complaint cp pump with serial number (b)(6) was visually inspected. a big chunk of biomaterial was observed in the pigtails and seen also in the outlet. no other abnormality found. the root cause of the low pump was determined to be ingested biomaterial. this report is being filed as part of a retrospective review of historical records. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. d6 is unknown. this report is being filed as part of a retrospective review of historical records.
18371 | 1220648-2024-19442 | 20379047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379047&pc=OZD | 07/05/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 07/05/2023 | 10/05/2024 |  |  | 12/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024324358 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to being implanted with an impella cp device for mechanical circulatory support. upon initiation of support, flows flows did not exceed 1.0 l/min in auto mode. staff attempted to adjust the flow level, but no increase in flows was observed. the impella device was repositioned, yet there was still no change in flow rates. consequently, the impella was explanted, and a new device was implanted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the low pump flow was most likely was patient condition as patient had severe three-vessel disease at the time with ef 30%. this report is being filed as part of a retrospective review of historical records.
18372 | 1220648-2024-19444 | 20379049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379049&pc=OZD | 02/20/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 02/20/2023 | 10/05/2024 |  |  | 11/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024294944 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella was reported to have been placed without issue. while on support, the patient became hypertensive and experienced significant access site groin bleeding at the puncture site around the repositioning sheath. the physician removed the impella reporting the patients blood pressure was too high and the groin bleeding excessively. reportedly, upon removal of the device, the patient was stable. the impella was removed due to the bleeding and patient's condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the access site bleeding has been completed. per clinical details the investigator determined that the root cause of the bleeding was patient condition related, as the patient became hypertensive, and three stents were placed. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18373 | 1220648-2024-19443 | 20379050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379050&pc=OZD | 07/05/2023 | 03/27/2025 |  | 10/05/2024 |  |  | 07/05/2023 | 10/05/2024 |  |  | 01/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333918 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, flows were low and suction alarms occurred. consequently, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. pump was investigated and no cannula kinks were observed. biomaterial was found around the impeller. no other abnormalities were found. the root cause of the low pump flow was biomaterial related to patient condition. this report is being filed as part of a retrospective review of historical records. | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
18374 | 1220648-2024-19445 | 20379110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379110&pc=OZD | 02/21/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 02/21/2023 | 10/05/2024 |  |  | 10/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024296954 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that after placement, the cannula kinked outside the aortic valve resulting in diminished flows of less than 1.0 liters per minute. reportedly, a fluoroscopy revealed a kink just below the outlet. the physician and clinician attempted to resolve it with repositioning of the impella, which was unsuccessful. the pump was removed and replaced with another impella, and flows were reportedly achieving 3.5 liters per minute. there was no patient harm reported, and this device was replaced. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the product damage has been completed. the review of the product log confirms the reported low pump flow and shows suction alarms before explant. there is also one impella position in aorta alarm before explant. the cause of the low pump flow was most likely a kinked cannula due to the patient anatomy based on the provided clinical information. this report is being filed as part of a retrospective review of historical records.
18375 | 1220648-2024-19446 | 20379175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379175&pc=OZD | 11/02/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 02/09/2023 | 10/05/2024 |  |  | 08/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251519 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, a hematoma at the groin site was noted. no additional information was received related to the intervention or resolution. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the access site bleeding was not determined since neither product nor data logs were returned and there are insufficient clinical details. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18376 | 1220648-2024-19447 | 20379190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379190&pc=OZD | 02/27/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 02/28/2023 | 10/05/2024 |  |  | 11/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024292047 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that device had continuous suction alarms due to kinked cannula. device position was not improvable. device was successfully explanted after recovery. patient was stable post procedure. | [5 empty] | The investigation for the reported issue has been completed. returned complaint pump visually inspected. no biomaterial. no broken blades found. kink in cannula and catheter observed. root cause of the low pump flow was the cannula kink due to patient condition because the physician had very difficult insertion with calcified ostium iliac. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ this report is being filed as part of a retrospective review of historical records.
18377 | 1220648-2024-19448 | 20379205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379205&pc=OZD | 02/28/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 02/28/2023 | 10/05/2024 |  |  | 08/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260079 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after 3 days of support, impella was explanted and escalated to impella 5.5 due to patient's condition. a thrombus was noted at the pump explant. patient survived the procedure. | [5 empty] | The investigation for the reported thrombus issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the thrombus was patient condition related. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18378 | 1220648-2024-19449 | 20379209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379209&pc=OZD | 03/01/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/08/2023 | 10/05/2024 |  |  | 04/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219340 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the provider pulled impella back into the aorta when attempting to reposition. the provider was not able to re-cross impella into the ventricle. the patient was decompensated later, and another pump was successfully implanted. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the positioning issues was wireless re-access. the impella device is not indicated for wireless re-access per ifu 0048-9007. this report is being filed as part of a retrospective review of historical records.
18379 | 1220648-2024-19450 | 20379218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379218&pc=OZD | 06/11/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 06/11/2022 | 10/05/2024 |  |  | 07/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121538 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the pump would not prime and was alarming. the pump was replaced. there was no patient harm. | [5 empty] | The investigation for the reported priming issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the retuned pump went through the priming process with no issues or alarms. purge fluid could be seen exiting the pump. the root cause of the failure to prime was determined to be a use related issue. this report is being filed as part of a retrospective review of historical records.
18380 | 1220648-2024-19451 | 20379223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379223&pc=OZD | 12/22/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 12/22/2022 | 10/05/2024 |  |  | 09/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275690 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. shortly after insertion into the left ventricle, the pump stopped and was unable to restart. the pump was removed and replaced by another cp. there was no lasting harm or injury. | [5 empty] | The investigation for the reported issue has been completed. data logs showed that impella stopped alarm 119 triggered upon attempting to start the pump; the alarm triggered again with a subsequent restart attempt. no external damage or abnormalities were found, including no bond failures. electrical continuity testing revealed an open line on motor phase 2. red handle was then opened, and the red motor wire was found to be damaged near the kink protector. extra length of catheter found within the red handle. no other damage or abnormalities were found. in conclusion, it was determined that the cause of the pump failing to start was an open electrical line. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
18381 | 1220648-2024-19452 | 20379225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379225&pc=OZD | 06/13/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 06/13/2022 | 10/05/2024 |  |  | 03/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203892 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the impella was unable to be delivered or positioned correctly due to the patient¿s cardiac anatomy. the pump was removed and a new impella placed. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was patient condition related. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18382 | 1220648-2024-19453 | 20379241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379241&pc=OZD | 03/01/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/01/2023 | 10/05/2024 |  |  | 10/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295990 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient for pre-operation placement was attempted to be implanted with an impella cp device for mechanical circulatory support. during attempted implant, device had delivery issues and procedure was aborted. no harm to the patient was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on clinical description, case was aborted due to patient condition (hyperdynamic left ventricle) & no issue for the pump. therefore, the root cause of delivery issue was due to patient condition. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18383 | 1220648-2024-19454 | 20379245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379245&pc=OZD | 06/15/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 06/15/2022 | 10/05/2024 |  |  | 04/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207614 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella was placed in a suboptimal position, due to the pump moving, and the catheter having a loop in the aorta. the physician made the decision to keep the position of the impella. the patient was hemodynamically stable. the patient remained on impella support and was successfully weaned. there was no patient harm reported. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that fluoroscopic videos received from the field showed the catheter was curved along the aortic arch. data logs showed "impella position in aorta" alarms at the beginning of support, but the alarms soon resolved, and the placement signal metrics were within normal limits. there were no issues with the other pump performance metrics. visual inspection of the returned product revealed a significant catheter kink between the 19-cm and 20-cm marks. no other damage or abnormalities were found. in conclusion, it was determined that the root cause of the positioning issues was most likely patient condition related. the catheter kink was likely a result of an interaction between the pump and aorta, leading to looping of the catheter and the repositioning difficulties. this report is being filed as part of a retrospective review of historical records.
18384 | 1220648-2024-19455 | 20379247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379247&pc=OZD | 03/01/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/01/2023 | 10/05/2024 |  |  | 11/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298209 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis during impella support. the patient was monitored after their urine output came back blood tinged. no intervention was deemed necessary for the condition. support continued following the event and the blood tinge was resolved. there was no patient harm reported. | [5 empty] | The investigation for the reported hemolysis issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. pump was not returned. the cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
18385 | 1220648-2024-19456 | 20379271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379271&pc=OZD | 03/27/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 03/27/2022 | 10/05/2024 |  |  | 12/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178365 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. when removing the peel away sheath and advancing the repositioning sheath, the pump pulled out of the patient's left ventricle. the pump was removed and replaced with a second impella, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the positioning issues was determined to be use related as the pump was accidently pulled from the ventricle while placing the repo sheath. this report is being filed as part of a retrospective review of historical records.
18386 | 1220648-2024-19457 | 20379278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379278&pc=OZD | 06/15/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 06/15/2022 | 10/05/2024 |  |  | 12/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180974 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was successfully placed. however, when removing the peel away sheath the pump collapsed onto itself and kinked. the physician was unable to recross the valve, despite multiple attempts to pull the pump out of the body the pump was unable to unkink. the decision was made to remove this pump and replace it with a new impella pump, which was successfully placed. there was no patient harm reported. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was patient condition related. this report is being filed as part of a retrospective review of historical records.
18387 | 1220648-2024-19458 | 20379279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379279&pc=OZD | 03/28/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 03/28/2022 | 10/05/2024 |  |  | 12/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177767 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the cp came out of position and was unable to be advanced back into the left ventricle. the pump was removed and replaced with a second impella cp, and there was no patient harm reported. | [5 empty] | The device was returned, and the investigation for the reported positioning issue has been completed. the root cause of the positioning issues was patient movement. this report is being filed as part of a retrospective review of historical records.
18388 | 1220648-2024-19459 | 20379286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379286&pc=OZD | 03/28/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 03/28/2022 | 10/05/2024 |  |  | 11/30/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167158 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Damaged by Another Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. impella cp was damaged by attempting to insert over 0.035" wire vs 0.018" wire. the pump was not implanted due to the damage. a new pump was used, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported product damage has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the damage was a use issue as the physician tried to insert the pump using the 0.035 wire instead of the 0.018 wire. per cp optical ifu, the impella 0.018 wire is used for placement. this report is being filed as part of a retrospective review of historical records.
18389 | 1220648-2024-19460 | 20379292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379292&pc=OZD | 03/02/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/02/2023 | 10/05/2024 |  |  | 07/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251530 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implant, device had delivery issues due to the easy guide lumen and interaction with the guidewire. the team decided to replace the entire system. new pump was implanted. no harm or injury to the patient from delay in support initiation. | [5 empty] | The investigation for the reported issue has been completed. the product was returned, a kink was noted in the cannula. the root cause of the delivery issue is most likely a use issue due to the kink in the cannula before insertion. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
18390 | 1220648-2024-19461 | 20379297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379297&pc=OZD | 06/16/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 06/16/2022 | 10/05/2024 |  |  | 10/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023159227 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon explant at the end of impella support, the clinician observed a large mass of thrombus caught on the pigtail. there was no harm to the patient. | [5 empty] | The investigation for the reported thrombus has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. a significant mass of biomaterial was observed on the impella upon explant. the origin of the biomaterial was not determined. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18391 | 1220648-2024-19462 | 20379301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379301&pc=OZD | 09/10/2021 | 10/05/2024 |  | 10/05/2024 |  |  | 03/29/2022 | 10/05/2024 |  |  | 05/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022085964 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced bleeding at the access site, requiring a blood transfusion. in addition, the patient's anticoagulants were adjusted and the access site was redressed. it was reported that the patient recovered. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the access site bleeding was most likely use related, since replacing the gauze resolved the issue. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18392 | 1220648-2024-19463 | 20379303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379303&pc=OZD | 03/30/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 03/30/2022 | 10/05/2024 |  |  | 09/30/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022137934 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after approximately one week of support, the pump stopped. the pump was explanted and there was no reported harm to the patient. | [5 empty] | The investigation for the reported bearing wear has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. visual inspection of the returned pump revealed a significantly large purge gap. no other external damage or abnormalities are found. therefore, it was determined that the cause of the pump stop was bearing wear. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18393 | 1220648-2024-19464 | 20379310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379310&pc=OZD | 03/30/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 03/30/2022 | 10/05/2024 |  |  | 12/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023179012 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. shortly after insertion, it was found that the flows were diminished. the physician discovered a kink in the device. attempts were made to reposition, but the kink remained. the pump was then removed and replaced, and no patient harm was reported. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was a kinked cannula, most likely related to technique/use. this report is being filed as part of a retrospective review of historical records.
18394 | 1220648-2024-19465 | 20379322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379322&pc=OZD | 06/16/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 06/16/2022 | 10/05/2024 |  |  | 02/29/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202182 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. during implantation, multiple attempts to advance the pump through the iliac artery were unsuccessful. the case was aborted. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was patient condition related. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18395 | 1220648-2024-19466 | 20379325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379325&pc=OZD | 03/02/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/02/2023 | 10/05/2024 |  |  | 10/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283719 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. angiogram showed severely diseased bilateral illiacs, resulting in delivery issues during implant. decision was made to remove impella and place a new impella cp pump. staff noted successful insertion of new impella pump. no harm was done to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was patient condition related. severely diseased bilateral iliacs (pad) reported. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18396 | 1220648-2024-19467 | 20379327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379327&pc=OZD | 03/30/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 03/31/2022 | 10/05/2024 |  |  | 08/31/2023 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127656 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient receiving an aortic valvuloplasty was implanted with an impella cp device for mechanical circulatory support. when the impella was removed, a piece of tissue was discovered on its tip. the patient survived and had no further reported issues. | [5 empty] | The investigation for the reported thrombus has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the thrombus was most likely patient condition. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18397 | 1220648-2024-19468 | 20379333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379333&pc=OZD | 09/07/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 09/07/2022 | 10/05/2024 |  |  | 01/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183639 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. the us complainant had a failed delivery of the cp for support of a patient. the cp was replaced by the intra-aortic balloon pump (iabp). there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issue has been completed. no product was returned. after reviewing the clinical information, it was determined that the cause of the delivery issue was patient condition related, specifically the patient's small vasculature. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18398 | 1220648-2024-19469 | 20379334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379334&pc=OZD | 06/17/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 06/17/2022 | 10/05/2024 |  |  | 08/31/2022 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1504363 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, transesophageal echocardiography was performed which revealed an apical thrombus with possibility of clot on the pigtail catheter. the impella stopped suddenly after 16 days of support. the pump was unable to be restarted, and the patient was brought to the operating room for impella explant and surgical wound closure. | [5 empty] | The investigation for the reported pump stop and thrombus issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs showed that the pump stopped immediately upon the occurrence of the "impella stopped - motor current high" alarm. the impella was unable to be restarted. visual inspection of the returned pump revealed a large impeller purge gap. no other abnormalities were found. therefore, it was determined that the root cause of the pump stop was bearing wear after 16 days of use, which is beyond the 8-day design life. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18399 | 1220648-2024-19471 | 20379376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379376&pc=OZD | 03/30/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/30/2023 | 10/05/2024 |  |  | 10/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024296938 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump flow rate began dropping to insufficient levels. the pump was removed and replaced with new impella cp and rp pumps, and the explanted pump was found to have a clot in the inlet. there was no patient harm reported. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on clinical narrative & data, the root cause of the low flow was most likely due to biomaterial ingestion into the pump. this report is being filed as part of a retrospective review of historical records.
18400 | 1220648-2024-19470 | 20379380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379380&pc=OZD | 03/29/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/29/2023 | 10/05/2024 |  |  | 11/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303407 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited sudden suction alarms at p8. the physician lowered the pump to p2, but the alarms continued. the pump was explanted, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the low pump flow could not be determined. this report is being filed as part of a retrospective review of historical records.
18401 | 1220648-2024-19472 | 20379387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379387&pc=OZD | 06/17/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 06/17/2022 | 10/05/2024 |  |  | 01/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190411 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the physician was unable to advance the pump after multiple attempts due to anatomy of the patient. the insertion was aborted. there was no allegation of harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was patient condition related. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18402 | 1220648-2024-19473 | 20379388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379388&pc=OZD | 02/18/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 02/18/2023 | 10/05/2024 |  |  | 04/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220006 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a fluid leak from the purge sidearm during impella support. no additional information was provided related to the reported issue. there was no harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the returned product was visually inspected and no abnormalities were observed. a syringe with colored fluid was connected to the yellow luer, pressure was applied and a small filter hole leak was reproduced. the cause of the purge leak was sidearm filter damage. as noted in the impella ifu: impella cp with smart assist system section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18403 | 1220648-2024-19474 | 20379390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379390&pc=OZD | 06/17/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 06/19/2022 | 10/05/2024 |  |  | 04/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207750 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp did not start. the device was replaced. there was no reported patient harm. | [5 empty] | The investigation for the reported pump did not start issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump was visually inspected and a catheter kink next to the motor was observed along with a fractured guidewire in the outflow cage. the catheter kink was indicative of excessive force and likely the cause of the guidewire fracture. the root cause of the pump unable to start was a damaged guidewire as a result of excessive force. this report is being filed as part of a retrospective review of historical records.
18404 | 1220648-2024-19475 | 20379391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379391&pc=OZD | 02/20/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 02/20/2023 | 10/05/2024 |  |  | 01/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189033 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. the complainant reported a delivery issue involving an impella cp optical pump. during attempted implant, it was noted that the motor housing on the impella was unable to pass through an existing stent within the patient's right femoral artery (rfa). the implant was aborted, and it was discovered that the pump was also unable to pass through the sheath because of the introducer being kinked. there was no patient harm because of the delivery issue. | [5 empty] | The investigation for the reported issues have been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. root cause of the issue was patient condition related due to tortuous and stent in right femoral artery (rfa) and totally occluded left femoral artery (lfa). introducer sheath kinked while attempted to insert the pump. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
18405 | 1220648-2024-19476 | 20379393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379393&pc=OZD | 06/17/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 06/17/2022 | 10/05/2024 |  |  | 02/29/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193532 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the physician confirmed placement under fluoroscopy. the impella cp was alarming for suction and mal positioning. the physician repositioned the impella, but the impella continued to alarm. the physician did not feel comfortable sending the patient to the unit with a malfunctioning impella cp, so he decided to exchange for new impella cp device. there was no reported harm to the patient. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
18406 | 1220648-2024-19477 | 20379396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379396&pc=OZD | 03/30/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 03/30/2022 | 10/05/2024 |  |  | 10/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290392 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. the pump insertion was aborted due to the patient anatomy and an intra-aortic balloon pump (iabp) was placed for support. there was no harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was patient anatomy, most likely due to patient condition. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18407 | 1220648-2024-19478 | 20379397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379397&pc=OZD | 03/31/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/31/2023 | 10/05/2024 |  |  | 09/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281660 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the device was unable to achieve optimal flows with good pump position. the pump was removed and replaced, and the second pump achieved adequate flows. there was no reported harm to the patient. | [5 empty] | The device was returned, and the investigation for the reported low pump flow has been completed. visual inspection found no issues on the pump. the pump ran without issue under pump run test. the failure mode could not be reproduced. the fluctuated flow & ps waves happened with suction alarms on data and no issues were found on the pump, indicating that the low flow was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
18408 | 1220648-2024-19479 | 20379398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379398&pc=OZD | 06/18/2022 | 10/05/2024 |  | 10/05/2024 |  |  | 06/18/2022 | 10/05/2024 |  |  | 02/29/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195193 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a pump stop occurred on the 11th day of impella cp support. the pump was unable to be restarted and was explanted as a result. there was no report of patient harm. | [5 empty] | The investigation for the reported pump stop issue has been completed. the data logs confirm a high motor current pump stop. the root cause of the pump stop was unable to be determined as no product was returned for analysis. the pump stop occurred on day 11 of use which is beyond the 8 day design life per design trace matrix. impella cp with smart assist system instructions for use for the related event are as follows: section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18409 | 1220648-2024-19480 | 20379399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379399&pc=OZD | 03/31/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/31/2023 | 10/05/2024 |  |  | 10/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285531 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the pump was started outside the body, then inserted prior to consulting the field team. during use, the pump had less than optimal flows. however, the procedure was completed with the impella pump, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the low pump flow could not be determined. this report is being filed as part of a retrospective review of historical records.
18410 | 1220648-2024-19481 | 20379403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379403&pc=OZD | 04/01/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 04/01/2023 | 10/05/2024 |  |  | 11/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312008 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was apparent that the impella had tracked into the sub valvular structures of the mitral valve. attempts to free the cp from the sub valvular structures resulted in a straightening of the pigtail, until finally there was a sudden jump of the catheter and it was displaced into the aorta. multiple attempts were then made to advance the impella cp back across the aortic valve, but failed. the impella was then explanted and replaced, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was wireless insertion. the impella device is not indicated for wireless insertion. this report is being filed as part of a retrospective review of historical records.
18411 | 1220648-2024-19482 | 20379404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379404&pc=OZD | 02/21/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 02/21/2023 | 10/05/2024 |  |  | 10/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024294031 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a placement signal not displaying on the automated impella controller (aic) console during support. while no harm was endured, the pump was removed and replaced, and support continued. | [5 empty] | The investigation for the reported issue has been completed. the device was returned for evaluation that found the light continuity test showed light exiting at the sensor face. a damaged sensor face was observed after sensor removal and imaging. the logs show a change in 5s parameters consistent with a sensor break. the cause of the optical signal issue was a damaged sensor. this report is being filed as part of a retrospective review of historical records.
18412 | 1220648-2024-19483 | 20379405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379405&pc=OZD | 04/01/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 04/12/2023 | 10/05/2024 |  |  | 08/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023264688 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. when the patient was returned to the intensive care unit, it was found that the impella had moved into the aorta. the impella was removed and replaced, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on clinical description, the root cause of the positioning issue was a use issue. this report is being filed as part of a retrospective review of historical records.
18413 | 1220648-2024-19484 | 20379415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379415&pc=OZD | 04/08/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 04/08/2023 | 10/05/2024 |  |  | 12/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318089 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was running as expected and then started developing suction alarm that could not be resolved. went through troubleshooting steps and could not be resolved. replaced with new pump and performed as expected. patient was stable post replacement. | [5 empty] | The investigation into the reported low pump flow has been completed. product was not returned for review. data logs confirmed the failure mode. data showed pump was started on p-9 with flow was in specification (3.1 l/min) for about six minutes then decreased gradually to 1.5 l/min. p level was changed to p-6 after that but the flow was not improved. pump then was run on p-2 with flow about 0.9 -1.0 l/min to the rest of the case. the root cause of the low flow was unable to be determined as no product was returned for reviewing. this report is being filed as part of a retrospective review of historical records.
18414 | 1220648-2024-19485 | 20379417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379417&pc=OZD | 02/22/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/01/2023 | 10/05/2024 |  |  | 12/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024299791 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. pump priming was not possible, cause priming procedure on automated impella controller (aic) being not available. procedure was aborted and percutaneous coronary intervention (pci) performed without impella support. no harm done to the patient. | [5 empty] | The investigation for the reported issue has been completed. the cause of the priming issue was use related and related to the pump and purge cassette already priming prior to the user initiating case start steps on the console. this report is being filed as part of a retrospective review of historical records.
18415 | 1220648-2024-19486 | 20379418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379418&pc=OZD | 03/03/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/03/2023 | 10/05/2024 |  |  | 11/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024294496 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported a kinked cannula after delivery of the device. several attempts to reposition were unsuccessful, therefore, the pump was removed and percutaneous coronary intervention (pci) proceeded with no harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. base on the clinical data root cause of the low pump flow was cannula kink in the ascending aorta and reposition failed due to patient condition related. this report is being filed as part of a retrospective review of historical records.
18416 | 1220648-2024-19487 | 20379423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379423&pc=OZD | 03/03/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/03/2023 | 10/05/2024 |  |  | 07/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254675 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patent had an access site bleed. the bleed prompted the team to angle match and plan for explant the following day. there was no lasting harm or injury. the bleed did not prompt any further intervention or blood products. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the issue was not determined since product was not returned and sufficient information was not provided. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18417 | 1220648-2024-19488 | 20379461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379461&pc=OZD | 06/10/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/10/2023 | 10/05/2024 |  |  | 03/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024339580 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. due to high motor currents and widening separation of left ventricle and aortic signals with adequate placement it was determined that the pump is failing. the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of optical signal issue. the device was not received for evaluation. the root cause of the optical signal issue was unable to be determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
18418 | 1220648-2024-19489 | 20379462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379462&pc=OZD | 03/13/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 03/13/2023 | 10/05/2024 |  |  | 01/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2024304130 |  | 0048-0014 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, demographics unknown, with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported the placement of the impella was aborted due to the patient¿s vasculature. reportedly the patient¿s small aortic arch anatomy made advancing of the impella impossible. there was no patient harm reported, and this device was removed and replaced with another impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. per clinical review the cause of the delivery issue was patient condition related due to the patient¿s aortic anatomy. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluation" this report is being filed as part of a retrospective review of historical records.
18419 | 1220648-2024-19490 | 20379463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379463&pc=OZD | 06/29/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/29/2023 | 10/05/2024 |  |  | 03/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2024333461 |  | 0048-0014 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, race and ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the placement of the impella was aborted due to the patient¿s vasculature. reportedly the patient¿s anatomy made advancing of the impella impossible to insert a pump into the tortuous femoral artery. additionally, reportedly an attempt was made to switch to the long 14f x 25-centimeter abiomed (peel-away), but it could not be inserted because of the anatomy. there was no reported patient harm reported, and this device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the delivery issue has been completed. per clinical review the cause of the delivery issue was patient condition related due to the patient¿s tortuous femoral artery. the root cause of the issue was patient condition related. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluation" this report is being filed as part of a retrospective review of historical records.
18420 | 1220648-2024-19491 | 20379464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379464&pc=OZD | 06/29/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/29/2023 | 10/05/2024 |  |  |  |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350426 |  | 1000080 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while attempting to access the left main artery the guide wire dissected the left main artery. reportedly the patient underwent an extensive percutaneous coronary intervention with complete revascularization, totaling over five hours, during which time they never required any type of pressors. additionally, a clot was suspected as the pump flows never flowed above 1 liter and had continuous suction. the impella was removed and successfully replaced with another impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the low pump flow and vascular damage has been completed. per clinical details there were low flows seen at the start of the case with some spikes in motor current. no positioning issues noted but there were suction issues. a clot was seen upon removal of the pump. the root cause of the low pump flow is most likely due to biomaterial. while attempting to access the left main artery the guide wire dissected the left main. the root cause of the vascular damage is most likely due to improper technique by the end user. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk pci & impella 5.5 with smart assist for use during section: warnings & cautions: cautions as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: use of impella with transcatheter aortic valves as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18421 | 1220648-2024-19492 | 20379465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379465&pc=OZD | 06/29/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 06/29/2023 | 10/05/2024 |  |  | 02/28/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2024317437 |  | 0048-0002 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, demographics unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that on the 9th day of support the purge pressure was high, concerns of possible impending pump stop. this impella was removed and a new impella placed. there was no patient harm reported, and this device was replaced. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the high purge pressure has been completed. per clinical details, high purge pressure was observed with 1065 millimeters of mercury at a purge flow of 2 milliliters per hour. no purge line kinks were observed during removal of the impella. the impella cp was used for more than 8 days of support and high purge pressure issue was observed on the 9th day. the data logs were not returned for review. the impella cp optical device is not indicated for support more than 8 days of design life period according to the design. the cause of the high purge pressure issue was not determined due to insufficient clinical information and no product, or data logs were returned for review. this report is being filed as part of a retrospective review of historical records.
18422 | 1220648-2024-19493 | 20379466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379466&pc=OZD | 07/03/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 07/03/2023 | 10/05/2024 |  |  | 06/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235363 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support during electrophysiology ablation. while on support, the patient experienced access site bleeding. a hematoma was noted at the left femoral midway through the case, manual pressure was and toward the end of the case the hematoma became more noticeable and the patient hemodynamics more unstable. reportedly manual pressure was held for another 30 minutes. the decision was made to remove the impella and place an intra-aortic balloon pump. the impella was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the access site bleeding has been completed. the cause of the access site bleeding cannot be determined since the complaint device not returned for investigation and insufficient data returned. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18423 | 1220648-2024-19495 | 20379482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379482&pc=OZD | 05/26/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/26/2023 | 10/05/2024 |  |  | 09/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023280224 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient Device Interaction Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. the complainant reported that the doctor decided to abort the implantation of the impella because the patient's coronary disease was not as severe as initially suspected and hemodynamically stable. the doctor decided to continue with percutaneous coronary intervention (pci). there was no harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. root cause of the delivery issue was patient condition related due to tortuous femoral arteries and stenosis. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after this report is being filed as part of a retrospective review of historical records.
18424 | 1220648-2024-19496 | 20379483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379483&pc=OZD | 05/26/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/26/2023 | 10/05/2024 |  |  | 01/31/2025 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024332114 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient Device Interaction Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. there were noted delivery issues; delivery was complicated by the length needed for deployment. the cp had the sterile sleeve removed and delivered for single access. there was no patient harm. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was patient condition related because the difficulty with insertion was related to the patient's long vasculature. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18425 | 1220648-2024-19497 | 20379486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379486&pc=OZD | 05/19/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 05/19/2023 | 10/05/2024 |  |  | 10/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023277361 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the physician was concerned that even though the purge liquid flow rate was 20 milliliters per hour, and purge liquid was replaced (with a 500 ml bag of glucose) twice a day, the purge pressure low alarm sounded after replacing the purge fluid. there was no patient harm reported the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation into the low pump flow issue has been completed. the impella pump was returned for investigation. the data logs were not available for review. the returned pump was visually inspected which revealed the pump was covered with mold. the visual did not observe a cannula kink, broken blade or biomaterial, the only visual observation was the mold that surrounded the impella. the impella cp pump was run in water to reproduce the purge alarm in the investigation laboratory, no purge pressure alarms were reproduced, and the pump flows were within specification. a catheter kink was observed. the cause of the purge issue was improper technique by the end user, use related as the low pump flow could not be reproduced. this report is being filed as part of a retrospective review of historical records.
18426 | 1220648-2024-19498 | 20379487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379487&pc=OZD | 04/18/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 04/18/2023 | 10/05/2024 |  |  | 10/31/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285797 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. during support, the impella cp pump came out of position. the pump could not be delivered back to the position within the left ventricle (lv) so the pump was explanted. the patient was reportedly stable. | [5 empty] | The investigation for the reported positioning issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the positioning issue was use related. this report is being filed as part of a retrospective review of historical records.
18427 | 1220648-2024-19499 | 20379490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20379490&pc=OZD | 04/19/2023 | 10/05/2024 |  | 10/05/2024 |  |  | 04/19/2023 | 10/05/2024 |  |  | 11/30/2024 |  | 10/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311174 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/19/2023 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. low pump flow was reported during support. physician gave fluids but flows dropped lower to 1l/min. physician looked under fluoro and discovered a kink in the cannula. physician decided to remove and place a new impella to continue with the procedure. | [5 empty] | Data confirmed the low flow for entire case. no other issues were found on the data logs. product was not returned for review. based on clinical description & data analysis, the root cause of low flow was due to cannula kink & the root cause of cannula kink was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
18428 | 1220648-2024-19209 | 20367810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20367810&pc=OZD | 12/08/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 12/08/2022 | 10/04/2024 |  |  | 04/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220133 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that an impella cp had low flows.. low flows remained throughout support with normal mc and echocardiogram (echo) images. it is unknown how the reported issue was resolved. it was noted the patient was successfully wean and the device was removed in cath lab. | [5 empty] | The investigation for the reported access site bleeding issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the low pump flow was most likely biomaterial. this report is being filed as part of a retrospective review of historical records.
18429 | 1220648-2024-19210 | 20367818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20367818&pc=OZD | 12/18/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 12/18/2022 | 10/04/2024 |  |  | 09/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272852 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA-CORPOREAL MEMBRANE OXYGENATION (ECMO) | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that impella cp support was used in addition to extracorporeal membrane oxygenation (ecmo). the team sent the patient for ecmo decannulation in the operation room (or) and afterwards observed a cp purge leak. the physician performed a purge bypass, and the patient remained on impella cp support. the pump remained on for support until wean and explant without report of harm or injury. it was reported that they survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a cause. based on the clinical account, the purge filter was found to be leaking after the patient returned from operating room for extra-corporeal membrane oxygenation (ecmo) decannulation. based on the clinical account, the cause of the low purge pressure is due to a leak at the filter. as noted in the impella ifu: impella cp with smart assist system section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18430 | 1220648-2024-19211 | 20367849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20367849&pc=OZD | 12/27/2022 | 03/10/2025 |  | 10/04/2024 |  |  | 12/27/2022 | 10/04/2024 |  |  | 06/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240351 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was elevated motor current/saturated flows that suggested an impending device failure. the pump was removed. after removal, clot/fibrin was noted by physician in impella. there was no patient harm and no new impella was inserted. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | Data logs showed the motor current gradually increased with no p-level change, accompanied by saturated flows. recurrent placement signal spikes occurred throughout this period. additionally, the purge pressure rapidly rose towards the tail end of the motor current rise. remnants of biomaterial were found on the impeller of the retuned pump. therefore, when considering the clinical information, it was determined that the cause of the saturated flow was ingested biomaterial.
18431 | 1220648-2024-19212 | 20368445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368445&pc=OZD | 09/11/2024 | 12/03/2025 |  | 10/04/2024 |  |  | 09/12/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1208383 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that while the patient was on impella 5.5 support the automated impella controller (aic) turned off and then back on while in use. the aic was replaced and no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name has been updated. h3 has been updated to device eval completed. h6 type of investigation code 10 added. | The product was returned for analysis. after 9 hours since the pump was started, console recorded "process imcgui no heartbeat" a couple of times. then, there was a "watchdog partial reset," which made the aic restart automatically. after restart, immediately after the pump starts, the console shutdown unexpectedly. then the console was restarted in 18 seconds. this confirms the reported failure mode. this console was tested after arriving in fs. the reported failure mode could not be reproduced naturally upon running the pump. no problem was found with the flash cards. there was no log evidence to confirm the a/c power issues. the unexpected reboot was misreported as 'aic - a/c power issues.' the root cause for the unexpected reboot was 'imcgui' process crashing, but the failure was not reproduced, so the cause of the 'imcgui' crash was not determined.
18432 | 1220648-2024-19213 | 20368482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368482&pc=OZD | 08/11/2024 | 09/25/2025 |  | 10/04/2024 |  |  | 08/11/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520010 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that on day 22 of patient with impella cp support, the pump stopped. the patient is hemodynamics stable, and inotropes remain the same. the impella was removed and the procedural outcome was the patient survived surgery. | The customer also reported that a placement signal not reliable alarm was also observed on 08/11/24. |  |  |  |  | The investigation for the pump stop has been completed. impella cp pump was not returned. per log analysis, small motor current spike without p-level change and additional resistance to the motor with saturated flows seen. increased motor current was found with higher flows greater than 3 to 3.5 l/min and at same p-level(p-6). the cause of the pump stop issue was most likely due to biomaterial. | The investigation has been completed. no product was returned. per the clinical investigation, the cause of the pump stop issue was most likely due to biomaterial and no other information as pump was not returned. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no problem found with the pump. upon review of data logs it is apparent that the console is the potential cause of the optical signal issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. d10 concomitant medical products and therapy dates.
18433 | 1220648-2024-19216 | 20368569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368569&pc=OZD | 09/13/2024 | 10/29/2024 |  | 10/04/2024 |  |  | 09/13/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1615676 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the plastic around the impella plug on the automated impella controller (aic) is broken. the dust shield is not able to move properly due to broken plastic. a different aic was used and no patient harm was reported. | [5 empty] | The investigation of automated impella controller (aic) - damage has been completed. the root cause of this broken pump connector cover was most likely use related. e4 should have been left blank on manufacturer device report 1220648-2024-19216. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
18434 | 1220648-2024-19214 | 20368605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368605&pc=OZD | 09/12/2024 | 10/28/2024 |  | 10/04/2024 |  |  | 09/12/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1786102 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported that the automated impella controller (aic) has a broken purge pressure sensor. the was no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. visual investigation showed the purge disc retainer was found to be broken. the root cause of the issue was broken purge disc retainer.
18435 | 1220648-2024-19215 | 20368610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368610&pc=OZD | 09/12/2024 | 02/14/2025 |  | 10/04/2024 |  |  | 09/12/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025503608 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant had a failed delivery of the impella cp due to extensive calcification of left femoral artery. the team removed and accessed successfully via the contralateral leg. the patient remained stable throughout the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to be delivered due to calcification in lfa. | The pump was returned for inspection and no abnormalities were noted upon inspection. the root cause of delivery issue remained the same as previously reported on manufacturer device report 1220648-2024-19215. d9 date device returned to manufacturer added. e4 should have been left blank on manufacturer device report 1220648-2024-19215. g1 reporting contact email revised on manufacturer device report 1220648-2024-19215 in accordance with updated procedures.
18436 | 1220648-2024-19217 | 20368616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368616&pc=PYX | 09/13/2024 | 11/07/2024 |  | 10/04/2024 |  |  | 09/13/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025529666 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support and the patient developed hemolysis. the patient had continuous renal replacement therapy started on support and eventually the pump needed to be explanted due to the hemolysis. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the hemolysis has been completed. the rp flex pump was returned and was visually inspected. very small biomaterial was observed around the impeller. no sharp edge or broken blade found. data logs were reviewed which showed no motor current (mc) spike observed outside of p-level changes. some negative cvp indicating likely suction at the beginning of the case. pump mostly run at lower p-levels. at higher p-levels (p-7) flows. the cause of the hemolysis issue was most likely patient condition related.
18437 | 1220648-2024-19219 | 20368635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368635&pc=OZD | 09/14/2024 | 12/03/2024 |  | 10/04/2024 |  |  | 09/14/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025539794 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp placed to escalate care from the intra-aortic balloon pump placed for the patient admitted in cardiogenic shock. the pump was placed, but there was an access site bleed that prompted a suture being placed to stop the bleeding. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be use issue because the bleeding was resolved by applying a suture at the site. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-19219.
18438 | 1220648-2024-19220 | 20368637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368637&pc=OZD | 05/17/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/17/2022 | 10/04/2024 |  |  | 12/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163879 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was oozing from the repo sheath. repositioning the sheath with side port did not resolve the oozing. continuous venovenous hemodiafiltration (cvvhd) ongoing, low dose. no other information provided. the patient survived the procedure. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the pump was returned and the repo unit was pressure tested and no leaks were found. the repo unit was examined, and no damage was found. the root cause of the repo leak was most likely use related due to patient movement. if patient movement shifted the insertion angle or positioning of the repo unit it can create a gap between the sheath and the blue butterfly. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18439 | 1220648-2024-19218 | 20368646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368646&pc=OZD | 09/12/2024 | 11/21/2024 |  | 10/04/2024 |  |  | 09/12/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481208 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 34 YR |  | Female |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and while on support, the purge pressure rose to >1000mmhg. no kinks were observed and the purge cassette was replaced with no success. tissue plasma activator protocol was initiated and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. no product or data logs were returned for evaluation. clinical details noted tissue plasminogen activator (tpa) protocol was administered and resolved high purge pressure alarms. the root cause of the high purge pressure was most likely due to biomaterial as it was able to be resolved with tpa. added-h6 impact code 4644.
18440 | 1220648-2024-19221 | 20368653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368653&pc=OZD | 09/15/2024 | 11/07/2025 |  | 10/04/2024 |  |  | 09/15/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1423376 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that while changing the purge tubing, the automated impella controller (aic) displayed purge disc not detected alarm. new purge tubing was used, and the alarm persisted. new aic was used. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | Additional information is listed in g3, g6, h3, h6, h10 the device was returned for evaluation. the cause is component failure. the purge disc flag was found to be broken (missing at blue retainer).
18441 | 1220648-2024-19222 | 20368693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368693&pc=OZD | 09/15/2024 | 12/03/2024 |  | 10/04/2024 |  |  | 09/15/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533455 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during support for the male patient, the impella cp was repositioned and after pulling back of the catheter the impella was pulled out of the left ventricle. re-crossing the valve was unsuccessful. the impella was therefore explanted. patient is stable and left ventricle (lv) function is noted to have normalized. | [5 empty] | The investigation into positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue most likely was use related during repositioning causing the device to pull out of the left ventricle. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-19222. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
18442 | 1220648-2024-19223 | 20368775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368775&pc=OZD | 09/17/2024 | 11/11/2024 |  | 10/04/2024 |  |  | 09/17/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545752 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support during high risk percutaneous coronary intervention. while closing the groin, bleeding was noted around the left femoral artery ¿ access site. physician attempted to close groin by perclose x2. a peripheral balloon was inserted to balloon tamponade the vessel. post intervention angio showed no bleeding. patient is noted to be stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18443 | 1220648-2024-19224 | 20368779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368779&pc=PYX | 09/12/2024 | 11/04/2024 |  | 10/04/2024 |  |  | 09/12/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025529634 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support. while on support, the sheath dilator kinked and split. another sheath was used and there was no patient harm reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the introducer damage has been completed. only short dilator was returned for investigation. dilator tip was found damaged. as per the given clinical, sheath dilator was kinked and split. doctor was successfully placed the new one. the cause of the introducer damage issue was not determined since damaged introducer was not returned and sufficient clinical information was not provided. the instructions for use for the impella rp smart assist system with automated impella controller and rp flex with smart assist system with automated impella controller states the following: section: warnings and cautions: cautions: ¿handle with care. the impella rp with smartassist system catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
18444 | 1220648-2024-19225 | 20368783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368783&pc=OZD | 09/18/2024 | 12/06/2024 |  | 10/04/2024 |  |  | 09/18/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535081 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support continued oozing from the impella access site. femstop applied and heparin further reduced. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed has been completed. the returned product was visually inspected and no abnormalities were found. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. added d9 device return date.
18445 | 1220648-2024-19227 | 20368821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368821&pc=OZD | 09/18/2024 | 11/01/2024 |  | 10/04/2024 |  |  | 09/18/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025453984 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhagic Stroke | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp had a computed tomography scan that showed cerebellar hemorrhage. the patient not a neurosurgical candidate. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
18446 | 1220648-2024-19226 | 20368826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368826&pc=OZD | 09/18/2024 | 11/13/2024 |  | 10/04/2024 |  |  | 09/18/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025462093 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Defective Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the 57-year-old male patient was on impella cp support for percutaneous coronary intervention (pci). following procedure, the decision was made to remove peel away sheath and place repositioning sheath. after the peel away sheath was removed, it was discovered that the clear collar lock of the dilator was still attached. it had broken off free from the dilator. when the staff attempted to remove the collar, the pump got displaced into the aorta. the decision was made to remove the pump. the patient is noted to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 & cp impella cp with smartassist system & impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the positioning issue and the introducer damage has been completed. product was not returned for review. the root cause of introducer damage was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of the positioning issues was most likely use related since the attempt to remove the clear collar lock of the dilator caused pump to displace into aorta. added-h6 impact code 4627.
18447 | 1220648-2024-19228 | 20368868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368868&pc=OZD | 05/28/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 05/28/2023 | 10/04/2024 |  |  | 03/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213506 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 42 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella cp for the support of a patient admitted in with multiple cardiac and systemic comorbidities. the patient was enrolled in the abiomed's long-term outcome and quality indicator registry database. the patient had known ischemic cardiomyopathy, mitral regurgitation, tricuspid regurgitation, and acute on chronic heart failure. the impella cp was placed via the axillary insertion site and was supporting for just under 19 days when the pump stopped and was unable to be restarted. the team attempted to troubleshoot by automated impella controller exchange, but found the pump was in need of explant and medical therapy. the patient survived the event. | [5 empty] | A.5 is unknown. the investigation has been completed. the impella device was discarded by the customer. data logs at the time of the pump stop were not available for review and the product was not returned. therefore, the root cause of the pump stop was unable to be determined. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
18448 | 1220648-2024-19229 | 20368892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368892&pc=OZD | 05/18/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/18/2022 | 10/04/2024 |  |  | 02/29/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199230 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during insertion, there were multiple attempts to pass the impella through bifurcation but would not cross. multiple attempts were made but unsuccessful. fluid bolus given and hemodynamics improved. insertion was aborted due to anatomic limitations. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the issue was patient condition related. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18449 | 1220648-2024-19230 | 20368946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20368946&pc=OZD | 09/14/2024 | 11/15/2024 |  | 10/04/2024 |  |  | 09/14/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1664272 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, the automated impella controller (aic) displayed a controller error (yellow alarm). a different aic was used and no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the console (aic) controller alarm yellow has been completed. data logs were reviewed which showed right as pump was starting up, the user got a controller 1045 opm communication error. after the pump is removed to switch to a different console, there was a 1035 error that causes an optical bench reset followed by another controller error but this time for opc communication stopping. the console was returned and the failure mode was reproduced as a 1035 error was received while logs were being downloaded. after opening the device, the ribbon cable between the 4-in-1 and optical bench was found to be slightly unplugged on the 4-in-1. after plugging it in fully, the controller error resolved and did not return. the cause of the controller errors due to opm communication was an unplugged ribbon cable between the optical bench and the 4-in-1.
18450 | 1220648-2024-19231 | 20369144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20369144&pc=OZD | 06/10/2024 | 03/18/2025 |  | 10/04/2024 |  |  | 06/10/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1477755 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support, the automated impella controller (aic) was showing a controller failure alarm. it was recommended to exchange the aic, but the physician made a decision to explant the impella. the procedural outcome was the patient survived the surgery. | [5 empty] | A.5 is unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of the automated impella controller (aic) - pump stop has been completed. ten days into running the impella cp pump, there was a motor current high alarm that triggered a pump stop. this persisted through numerous attempts to restart the pump over the following three hours that were unsuccessful. there was no problem found with the console and it was able to run a purge and pump as expected. there was no issue found during the case. the device functioned/operated to specification. d.2 revised common device name since the initial manufacturer device report was submitted in accordance with updated procedures. g.1 revised reporting contact email since the initial manufacturer device report was submitted in accordance with updated procedures. h.6 code 4627 was reported incorrectly on the in the initial manufacturer device report 1220648-2024-19231 and a new code was added.
18451 | 1220648-2024-19232 | 20369185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20369185&pc=OZD | 05/19/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/20/2022 | 10/04/2024 |  |  | 01/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174062 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there was bleeding noted at the access site. the pump was removed, and a new pump was reinserted using an introducer with a dry seal. the procedure was completed without any problems. | [5 empty] | D6 explant date unknown. the investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18452 | 1220648-2024-19234 | 20369376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20369376&pc=PYX | 09/10/2024 | 11/19/2024 |  | 10/04/2024 |  |  | 09/10/2024 | 10/04/2024 |  |  | 05/31/2025 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024346906 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella rp was pulled back into the right ventricle, and it was unable to be advanced. the impella rp was almost pulled out for removal when it got stuck in the right femoral vein because there was a wound vac above the site where a cutdown was done for bypass cannulation. there were extra staples, and gauze in the site. that is what stopped the pump from being fully removed. it was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported positioning issue has been completed. the product was not returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review.
18453 | 1220648-2024-19233 | 20369399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20369399&pc=PYX | 09/15/2024 | 11/18/2024 |  | 10/04/2024 |  |  | 09/16/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025529692 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a surgical 52-year-old female patient admitted for coronary artery bypass graft (cabg) and valve. the impella rp pump positioning was causing oxygen desaturation. the pump was repositioned but the issue did not resolve. the pump remained for 13 days of support and was successfully explanted. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ | The investigation for the positioning issue has been completed. the rp flex was returned and visually inspected. no cannula kink or catheter kink observed. pigtail deformed during support or removal. no any other abnormality observed. the cause of the positioning issue was patient condition related due to patient small size and small aortic arch/ventricle.
18454 | 1220648-2024-19235 | 20369434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20369434&pc=OZD | 05/19/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/19/2022 | 10/04/2024 |  |  | 01/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174061 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Female |  |  | Asian | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, bleeding was noted from the surgical anastomotic site of the graft, so the impella was removed. it was reported that there was not enough room in the bellows of the artificial blood vessel when fixing the repositioning sheath and the artificial blood vessel. because the artificial blood vessel was short, there was a possibility of bleeding due to bed movement or body movement. additionally, it was noted that the floppy part of the 0.018 inch wire was stretched when j-shaped with a dilator. | [5 empty] | The investigation for the reported access site bleeding and guidewire damage issues have been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the bleeding was most likely patient movement. the root cause of the guidewire damage was most likely due to use as the damage was reported as it was being inserted. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18455 | 1220648-2024-19237 | 20369589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20369589&pc=OZD | 09/10/2024 | 11/12/2024 |  | 10/04/2024 |  |  | 09/10/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522236 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while a 62-year-old male patient was on support, the impella cp showed signs of failure. the staff was aware of the length of duration an impella cp should be in place and does not see escalation as a viable solution as the patient has a poor prognosis. purge solutions were reviewed and the staff was aware that the pump could fail. troubleshooting for pump failure was reviewed. the pump eventually stopped but was restarted. care was withdrawn and the patient was expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation has been completed. no product was returned for analysis. the data logs showed "impella stopped - motor current high' alarms. the motor current spike was confirmed in the logs before the pump stop. there were no purge related issues mentioned. it was mentioned that there was no systemic heparin in the days prior and on 08-sep-2024 heparin was removed from the purge. there were restart attempts and the pump was able to run after initial pump stop. there was a console shutdown request at the end of the procedure. the root cause of the pump stop was not determined as no product was returned and insufficient information related to failure was provided.
18456 | 1220648-2024-19238 | 20369591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20369591&pc=OZD | 02/14/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 02/14/2022 | 10/04/2024 |  |  | 10/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156953 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported a pump side arm leak after approximately ten days of support. the clinician discussed with the team reducing the amount of interaction connecting and disconnecting the yellow hub. there was no patient harm reported, and this device was removed and replaced with a new impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the product purge leak has been completed. the root cause of the purge system leak was unable to be determined as the exact location of the purge leak could not be determined due to the product not being returned. the purge system leak occurred on day ten of support which is beyond the design specification of eight days. the cause of the purge leak was not determined. as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi section: cleaning as noted in the impella ifu ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18457 | 1220648-2024-19236 | 20369759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20369759&pc=OZD | 02/14/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 02/14/2022 | 10/04/2024 |  |  | 09/30/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022130302 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 26 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with myocarditis was implanted with an impella device for mechanical circulatory support. a kink was observed under fluoro. the pump was replaced. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18458 | 1220648-2024-19239 | 20369790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20369790&pc=OZD | 05/20/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/20/2022 | 10/04/2024 |  |  | 02/29/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194949 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implantation, delivery past the patient's stenotic aortic valve failed. the impella use was aborted and an intra aortic balloon pump (iabp) was inserted. there was no harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was a stenotic valve. impella cp with smartassist for use during cardiogenic shock and high risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
18459 | 1220648-2024-19240 | 20369900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20369900&pc=OZD | 02/14/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 02/15/2022 | 10/04/2024 |  |  | 06/30/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100474 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was taken to the cardiac catheterization laboratory (ccl) for impella pre-operative placement. while on support, the patient developed limb ischemia. no further information regarding treatment was noted. the patient remained on impella support and was successfully weaned. patient survived the procedure. | [5 empty] | The investigation into limb ischemia has been completed. the impella device was not received from the customer, therefore, investigation of the device was not possible. the root cause of the leg ischemia was most likely patient condition. no product or logs were returned for this investigation. as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
18460 | 1220648-2024-19241 | 20369975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20369975&pc=OZD | 02/15/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 02/15/2022 | 10/04/2024 |  |  | 11/30/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166790 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the physician attempted to reposition the device several times due to the continuous suction issues, by pulling the pigtail to the valve. during repositioning attempts, the impella popped across the valve. an echocardiogram revealed an old effusion was the root cause. this device was removed and replaced with another impella. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the positioning issue has been completed. the root cause of the positioning issues was determined to be improper use technique related as the pump was accidentally pulled back across the valve during repositioning. this report is being filed as part of a retrospective review of historical records.
18461 | 1220648-2024-19242 | 20370113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370113&pc=OZD | 05/20/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/20/2022 | 10/04/2024 |  |  | 02/29/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200308 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support for 15 hours, the pump stopped. the pump stop caused the team to give extra pressors and cardiopulmonary resuscitation (cpr) from pulseless electrical activity (pea). the impella was explanted and a new impella cp was placed. the patient was noted to be stable. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs were returned and show an abrupt high motor current pump stop. the pump was restarted but experienced high purge pressure and purge system blocked alarms. the pump was returned and biomaterial was found in the purge gap. the root cause of the high motor current pump stop was determined to be ingested biomaterial. impella cp with smart assist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18462 | 1220648-2024-19244 | 20370118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370118&pc=OZD | 02/22/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 02/22/2022 | 10/04/2024 |  |  | 10/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022145448 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the device stated to generate aspiration alarms and low flow. the physician made the decision to replace the device with another impella cp pump. no patient harm was reported. | [5 empty] | The investigation into the low pump flow has been completed. the impella pump was returned for investigation. the suction event was reported to have occurred on 15feb2022, incomplete logs were retuned that do not show this date. the product was returned, and biomaterial was found in the cannula. no kinks could be seen on the device. the pump was run in a basin of water and the biomaterial was ejected and the pump ran within specification. the root cause of the low flow and suction was concluded to be biomaterial ingestion from the patient¿s anatomy. this report is being filed as part of a retrospective review of historical records.
18463 | 1220648-2024-19245 | 20370152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370152&pc=OZD | 02/16/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 02/16/2022 | 10/04/2024 |  |  | 10/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156172 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump had a low purge pressure issue and leaking at the sidearm. the clinician was unable to visualize the leak due to sterile drapes and the impella being removed before locating the leak. there was no patient harm reported. | [5 empty] | The investigation into the purge leak has been completed. the impella pump was returned for investigation. a visual inspection of the returned pump revealed a large crack in the filter and the leak was reproduced. the cause of the purge leak was sidearm filter damage. the complaint was confirmed. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi section: cleaning as noted in the impella ifu ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18464 | 1220648-2024-19246 | 20370156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370156&pc=OZD | 09/10/2024 | 11/19/2024 |  | 10/04/2024 |  |  | 09/10/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532965 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaintant reported once the foley catheter was placed in the cardiovascular intensive care unit, the patient¿s urine noted to be bright red. the patient¿s lab work reportedly hemolyzed. mean arterial pressure was >120. milrinone and nicardipene was being ordered for afterload reduction. the uurse stated the device was repositioned under point of care united states guidance, but no further details were provided. the patient survived the explant. | [5 empty] | The investigation of hemolysis has been completed. the cause of the hemolysis issue could not be determined since the complaint device not returned for analysis and insufficient clinical data provided. data logs were reviewed and no motor current spike was observed. positioning issue alarms occurred that resolved with quick trouble shooting. a few suction alarms occurred. no indication observed to confirm the cause of the reported hemolysis issue. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
18465 | 1220648-2024-19249 | 20370164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370164&pc=OZD | 02/20/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 02/20/2022 | 10/04/2024 |  |  | 11/30/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161814 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, blood was leaking from the repositioning sheath hub. heparin was being withheld out of the purge solution. it was also noted that the patient was bleeding at access site from what appeared to be directly through the impella. multiple attempts to mitigate bleeding had not been successful. dressing was being changed every 2 hours. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation into the access site bleeding has been completed. the impella pump was returned for investigation. the investigator reported the gwru failed leak reproduction testing as leaking was observed at the interface between the catheter and the hemostatic valve after applying approximately 19 millimeters of pressure. after opening the gwru, the o-ring valve was found to be significantly deformed. the root cause of the access site bleeding was a damaged o-ring valve in the gwru. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18466 | 1220648-2024-19247 | 20370184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370184&pc=OZD | 01/14/2023 | 03/07/2025 |  | 10/04/2024 |  |  | 01/14/2023 | 10/04/2024 |  |  | 08/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270617 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis, urine output was reported to be dark. the resolution for this issue is unknown. additionally, the device exhibited continuous suction, and the motor current increased with flows close to 4.3. the physician decided to explant the device and replace it with another impella cp. it was reported that the patient was stable post replacement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | The long term log confirms suction alarms and a drop in flow. at the time of the first suction alarm there is a spike in motor current. a visual inspection of the pump revealed biomaterial on the impeller. the pump was run at p9 and the low pump flow was reproduced. a second visual inspection revealed that biomaterial remained on the impeller after the pump was run in the lab. the cause of the low pump flow was biomaterial. this report is being filed as part of a retrospective review of historical records.
18467 | 1220648-2024-19248 | 20370234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370234&pc=OZD | 02/20/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 02/20/2022 | 10/04/2024 |  |  | 03/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022065028 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding and low pump flow. the physician could not control the bleeding from the introducer sheath and cracked the peel away sheath while still in the groin. the physician then removed the impella, and a clot was observed. a new short abiomed introducer sheath was then used with the new impella which resolved the bleeding. there was no patient harm reported. | [5 empty] | The investigation into the access site bleed and low pump flow has been completed. the impella device was not received from the customer, therefore, investigation of the device was not possible. the cause of the access site bleeding was improper technique/use related as the introducer was peeled away in the patient. the cause of the low pump flow was not determined, no clinical details provided were sufficient to determine a root cause. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18468 | 1220648-2024-19243 | 20370324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370324&pc=OZD | 09/16/2024 | 11/18/2024 |  | 10/04/2024 |  |  | 09/16/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526923 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Malfunction | No | Yes | Material Puncture/Hole | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that in an attempt to insert the impella cp via axillary approach, it was noted that the white catheter was damaged before pump entered the female patient¿s body. a new impella was therefore used. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 & cp impella cp with smartassist system & impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the product damage has been completed. no product was returned for evaluation. clinical details noted that white catheter on pump was damaged before pump was inserted in patient's body. it was not noted what kind of damage occurred. clinical details states that the pump was never inserted into the patient and a puncture was not noted. the root cause of the product damage (cannula kink) cannot be determined with limited clinical details were provided.
18469 | 1220648-2024-19252 | 20370325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370325&pc=OZD | 02/21/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 02/21/2022 | 10/04/2024 |  |  | 11/30/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165921 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that post weaning and removal of the impella pump, a thrombus was found adhered to the impella. patient survived the procedure. | [5 empty] | The investigation into the thrombus issue has been completed. the impella device was not received from the customer, therefore, investigation of the device was not possible. the thrombus was sent out for analysis by the hospital and found to be a fibrous clot. a visual inspection of an image returned from the field confirmed the thrombus. the root cause of the thrombus noted upon explant was most likely patient condition as no issues were reported with the pump performance. no product or logs were returned for this investigation. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smart assist system. section: warnings & cautions: warnings: section: pre-support evaluation . section: impella cp with smart assist catheter insertion . section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. as noted in the impella ifu "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". this report is being filed as part of a retrospective review of historical records.
18470 | 1220648-2024-19251 | 20370332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370332&pc=OZD | 01/15/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/15/2023 | 10/04/2024 |  |  | 09/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278937 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported suction issues during impella support. the cause for the issue was found to be a fibrin clot within the motor. repositioning and fluid manipulation were attempted to troubleshoot the issue, but suction remained. as a result, the decision was made to remove the pump. there was no reported harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. root cause of the low pump flow most likely biomaterial as the clinical data stated clot observed close to the outlet area. data log reviewed and showed many suction alarms and motor current spikes occurred. this report is being filed as part of a retrospective review of historical records.
18471 | 1220648-2024-19255 | 20370406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370406&pc=OZD | 02/22/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 02/22/2022 | 10/04/2024 |  |  | 11/30/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162678 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the pump stopped during support. the device was removed and replaced with no harm to the patient. | [5 empty] | The investigation for the reported pump stop issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause of the pump stop was bearing wear. this report is being filed as part of a retrospective review of historical records. as noted in impella ifu: impella cp smartassist system: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
18472 | 1220648-2024-19256 | 20370553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370553&pc=OZD | 09/10/2024 | 11/14/2024 |  | 10/04/2024 |  |  | 09/10/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545537 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a femoral impella cp for the support of a patient admitted in acute myocardial infarction/cardiogenic shock. the patient had an ooze at the groin site that lead to heparin removal and sodium bicarbonate added to the purge. a day later, the patient went and had a computed tomography and the imaging found an aortic thrombus. the team gave the patient red blood cells and a thrombectomy. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of access site bleeding and thrombus has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was determined to be use issue because the bleeding was controlled by removing and redoing the sutures. as the necessary information was not provided, the root cause of the thrombus was not determined.
18473 | 1220648-2024-19254 | 20370562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370562&pc=OZD | 09/16/2024 | 11/18/2024 |  | 10/04/2024 |  |  | 09/16/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025511344 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support for 55-year-old male patient, an air in purge system alarm occurred followed by purge pressure high, and ultimately impella stopped motor current high. the purge cassette was replaced but the issue did not resolve. the impella 5.5 was explanted and another impella 5.5 was placed. there was no reported patient harm. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the high purge pressure and the saturated flow has been completed. there were no data logs returned but the product was returned. the pump was returned for evaluation. upon functional testing, the high purge pressure and pump stop were reproduced and there was biomaterial found in the purge gap as well as in the purge line indicating blood ingress into the motor. this blockage that caused the high purge pressure led to the high motor current pump stop. there were no "air in purge" alarms when attempting to purge the pump. the sidearm was cut and pressurized and no leaks were reproduced. there was no problem found regarding the purge leak - pump. the root cause of the pump stop was due to biomaterial.
18474 | 1220648-2024-19257 | 20370620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370620&pc=OZD | 05/20/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/20/2022 | 10/04/2024 |  |  | 11/30/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165942 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella cp device for mechanical circulatory support. during implantation, the impella device kinked upon removal of the supplied guidewire. the physician believes that the guidewire was sheared into two pieces. one piece still remained in the cannula at removal, and the other was discarded and couldn¿t be retrieved. the pump was removed and a new impella was used without further issue. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the pump was returned, and a small segment of guidewire was found in the cannula. the guidewire damage was confirmed. imaging from the insertion was provided that confirms the cannula kink. the root cause of the delivery issue was most likely use related, this report is being filed as part of a retrospective review of historical records.
18475 | 1220648-2024-19250 | 20370651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370651&pc=OZD | 05/20/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/20/2022 | 10/04/2024 |  |  | 10/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023159084 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a pump stop was noted upon attempts to initiate the device. multiple troubleshooting factors performed. the pump was removed and replaced with no harm to the patient. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the data logs show high motor current pump stops when the pump attempts to start up. the cause of the issue was not determined since product was not returned. impella cp with smart assist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
18476 | 1220648-2024-19253 | 20370748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20370748&pc=OZD | 09/10/2024 | 10/29/2024 |  | 10/04/2024 |  |  | 09/10/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025502041 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and experienced limb ischemia after the implant. pressor support was lowered. however, the limb ischemia continued. it was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
18477 | 1220648-2024-19261 | 20371093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371093&pc=OZD | 09/17/2024 | 11/25/2024 |  | 10/04/2024 |  |  | 09/17/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523713 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported when the 58-year-old male patient was on impella cp support, the pump had high motor current and the pump stopped. several attempts were made to restart the pump to no avail. the pump was successfully explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation into saturated flow and pump stop has been completed. the 'saturated flow' failure mode was deleted because there was no mention of saturated flow in the customer complaint intake, regulatory evaluation, or patient updates. no product was returned for analysis. the data logs returned are insufficient and do not go up till complaint occurrence date. the returned logs are from (b)(6) 2024 till (b)(6) 2024 while reported pump stop occurred on (b)(6) 2024. the returned logs show no motor current spikes and placement signal was trending down. the root cause of the pump stop was not determined as no product was returned and insufficient data logs were provided
18478 | 1220648-2024-19264 | 20371126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371126&pc=OZD | 01/19/2023 | 10/20/2025 |  | 10/04/2024 |  |  | 01/19/2023 | 10/04/2024 |  |  | 08/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268398 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that following impella insertion, the user was unable to obtain a good position for the pump. multiple attempts were made, however the positioning remained poor. due to the noted issue, the pump was removed and replaced with another pump. no further issues were noted following the replacement. there was no reported harm to the patient. | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. root cause cannot be determined since the device not returned for investigation and lack of clinical details.
18479 | 1220648-2024-19263 | 20371175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371175&pc=OZD | 01/16/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/16/2023 | 10/04/2024 |  |  | 10/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023280018 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced ventricular fibrillation (v-fib) requiring defibrillation and flow saturation requiring second device to be placed. no further information is available. | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
18480 | 1220648-2024-19265 | 20371187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371187&pc=OZD | 01/16/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/16/2023 | 10/04/2024 |  |  | 07/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023258946 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. the complainant reported a purge leak from the pressure disc. the pump was removed and replaced with no reported harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. leaking was observed from the sidearm reservoir during reproduction testing. the leaking originated from a puncture hole in the blue diaphragm of the reservoir. no other damage or abnormalities were found. therefore, it was determined that the cause of the purge leak was use-related sidearm reservoir damage. as noted in the impella ifu: impella cp with smart assist system section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18481 | 1220648-2024-19262 | 20371192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371192&pc=OZD | 01/15/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/15/2023 | 10/04/2024 |  |  | 09/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276805 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported viewing possible air bubbles in the patient's ventricle during support. there was no other indication of an issue with the device. the pump was removed proactively, and a new pump was placed. there was no reported adverse impact to the patient. | [5 empty] | The investigation for the reported issue has been completed. the data logs were reviewed, the pump and purge cassette appeared to be primed for approximately 10 minutes before the pump was inserted. the pump was returned and primed with dyed purge fluid. no leaks in the system were observed. the pump was then run in a bucket of water and no air bubbles were produced. the root cause of the suspected air bubbles was most likely use related as no functional issues with the impella were found and the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
18482 | 1220648-2024-19258 | 20371193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371193&pc=OZD | 10/01/2024 | 11/15/2024 |  | 10/04/2024 |  |  | 09/16/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474117 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Material Twisted/Bent | [5 empty] | Arrhythmia | Atrial Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, the patient had episode of unsustained ventricular tachycardias (vt) and atrial fibrillation (af). the patient was shocked out by implantable cardioverter defibrillators (icd). additionally, impella stopped alarm occurred. attempts to restart were unsuccessful. pump was removed due to unexpected impella failure. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of vf/vt/af/at and pump stop has been completed. the pump was returned for investigation. the data log was missing for the last two hours of the pump run. a catheter kink was noted at the kink protector. the pump catheter was found stretched near the red handle, and the motor current line was found open inside the red handle. no other physical abnormalities were observed on the returned product. additionally, the pump stopped and could not be restarted after patient movement from chair to bed. the cause of the pump stop issue was an open electrical line inside red handle. as the necessary information was not provided, the root cause of the vt/vf was not determined. d9 date device returned to manufacturer added.
18483 | 1220648-2024-19266 | 20371410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371410&pc=OZD | 09/17/2024 | 11/19/2024 |  | 10/04/2024 |  |  | 09/17/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523106 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the physician attempting to reposition impella cp when it was inadvertently pulled across the valve. the physician made decision to remove the cp. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was a use issue since the pump was inadvertently pulled across the valve during repositioning.
18484 | 1220648-2024-19259 | 20371491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371491&pc=OZD | 05/19/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/20/2022 | 10/04/2024 |  |  | 10/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144089 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella cp device for mechanical circulatory support. during priming, there was a purge liquid leak alarm. it was restarted several times but was not resolved. the purge line was checked and no issue could be found. the pump was replaced after the cassette replacement. there was no patient harm. | [5 empty] | The investigation for the reported priming issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that returned compliant pump and purge cassette visually inspected. pump and purge cassette connected to run no purge observed in the pump at the beginning. purge cassette cut open leak observed in thf bonding. the pump connected to another purge cassette. returned pump and new purge cassette working normal and priming normal. no priming error alarm detected. the root cause of the priming issue was determined to be bond failure. this report is being filed as part of a retrospective review of historical records.
18485 | 1220648-2024-19268 | 20371500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371500&pc=OZD | 05/22/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/22/2022 | 10/04/2024 |  |  | 12/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178112 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the impella cp could not cross the aortic valve. after multiple attempts to pass the pump through the sheath with no success, the case was aborted. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was determined to be patient condition as the patient had aov stenosis. impella cp with smartassist for use during cardiogenic shock and high risk cpi instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18486 | 1220648-2024-19271 | 20371516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371516&pc=OZD | 01/19/2023 | 10/18/2024 |  | 10/04/2024 |  |  | 01/19/2023 | 10/04/2024 |  |  | 09/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274436 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was presented for impella cp device for mechanical circulatory support. it was reported that the procedure was aborted due to patient status. no reported harm done to the patient. | [5 empty] | Initial mfr 1220648-2024-19271 was incorrectly submitted as a reportable product malfunction due to delivery issues from patient incompatibility. upon further review, it was determined that no product malfunction occurred that was likely to cause or contribute to death or serious injury. the patient was in acute myocardial infarction with coronary microvascular disease with a critical left main coronary artery and given the patient's condition, the medical team made the change in how to treat the patient. the medical team decided to support the patient with bypass instead of using the impella device. the impella device was never inserted nor attempted to be inserted into the patient due to the medical team's choice of treatment. | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was determined to be patient condition as the patient's condition was noted to be poor so the physician decide to abort the insertion of the device. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18487 | 1220648-2024-19270 | 20371532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371532&pc=OZD | 01/17/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/17/2023 | 10/04/2024 |  |  | 07/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250448 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was presented for impella cp device for mechanical circulatory support (mcs). the complainant reported an inability to implant due to vasculature of patient anatomy. no reported harm done to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was patient condition related, specifically the patient's vascular condition. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18488 | 1220648-2024-19260 | 20371545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371545&pc=OZD | 01/15/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/15/2023 | 10/04/2024 |  |  | 08/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023271103 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the device purge system was not working so pump placement was aborted and new pump used without issue. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the priming issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
18489 | 1220648-2024-19275 | 20371698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371698&pc=OZD | 09/13/2024 | 11/22/2024 |  | 10/04/2024 |  |  | 09/18/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025497121 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient was successfully weaned off the pump and aic support without incident which was one day earlier than planned. there were no complications related to the aic. the patient did expire from a ventricular tachycardia episode post support and assessed as related to the impella pump. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | Complainant in japan reported the patient was on impella cp smart assist support and there were low purge pressure alarms. the low purge pressure occurred around the fifth day of support. the purge cassette was replaced to no avail. the pump was explanted a day earlier than planned. additionally, a few days after weaning from the impella the patient decompensated and expired from ventricular tachycardia (vt). neither the vt nor cause of patient's death was found to be impella related. |  |  |  |  | The investigation for the purge leak-pump has been completed. the pump and 2 purge cassettes were returned. all devices functioned to specification and no leak was reproduced. the pump was run with both purge cassettes and there were no leaks; in addition, the purge pressures during the lab run were within specification. there were no physical abnormalities with any of the devices. the data logs show low purge pressure alarm and high purge flow towards the end of the case. the root cause of the purge leak - pump was most likely a use-related issue. corrections to the initial mfr report: b1-selected adverse event. b2-selected death and added date of death. b5-mso review. h1-type of reportable event changed to death. h6 impact code added 1802. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
18490 | 1220648-2024-19276 | 20371732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371732&pc=OZD | 01/25/2023 | 10/13/2025 |  | 10/04/2024 |  |  | 01/25/2023 | 10/04/2024 |  |  | 09/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023282457 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Defective Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, a patient had low pump flow. after verifying the device position, a kink in the cannula was observed on fluoroscope near outlet cage. the pump was removed and replaced with a new impella cp. no patient harm was reported. | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that after impella support was initiated, the pump flows would not go above 1.2l, motor current flat, and there were suction alarms. after verifying position kink in cannula noted on fluoroscopy near outlet cage. the decision was made to replace the impella cp pump with a new device. it was also noted that the short peel away sheath appeared to be leaking with no obvious cause. the sheath was replaced when the new pump was opened. the second impella cp serial number was (b)(6). it was reported that the patient experienced an expanding hematoma at the left femoral impella access site and the patient's hemoglobin dropped from 8 to 6.4. the patient was transferred to the cicu due to concern of hemorrhagic shock and the patient was transfused with two units of packed red blood cells. the patient was late successfully weaned and explanted. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The returned pump was visually inspected and a cannula kink was observed. introducer was not returned for review. the cause of the issue was a kinked cannula. no additional information was provided. the cause of the introducer bleeding was not determined since product was not returned.
18491 | 1220648-2024-19269 | 20371760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371760&pc=OZD | 09/17/2024 | 11/20/2024 |  | 10/04/2024 |  |  | 09/17/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025506863 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that a patient was on impella 5.5 support. while on support, the impella was dislodged from optimal position and unable to be repositioned. impella was explanted. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined.
18492 | 1220648-2024-19267 | 20371771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371771&pc=OZD | 01/16/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/16/2023 | 10/04/2024 |  |  | 08/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260746 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  |  | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the pump experienced suction issues during impella support. support continued throughout the event; however, it was noted that the catheter appeared to be kinked under fluoroscopy. as the catheter was not kinked prior to the percutaneous coronary intervention procedure, it was ascertained that the catheter had been damaged during attempts to achieve a return of spontaneous circulation. the pump was explanted as a result of the failure. no further information is available. | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
18493 | 1220648-2024-19272 | 20371824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371824&pc=OZD | 05/22/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/22/2022 | 10/04/2024 |  |  | 12/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174692 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the pump would not pass through the aortic valve and the catheter crimped about 1 inch from the motor housing. the pump was then removed. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smartassist for use during cardiogenic shock and high risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18494 | 1220648-2024-19279 | 20371909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371909&pc=OZD | 09/18/2024 | 11/19/2024 |  | 10/04/2024 |  |  | 09/18/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025496202 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Migration or Expulsion of Device | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump delivered to the left ventricle (lv) to support a patient admitted in with acute myocardial infarction / cardiogenic shock. the pump came out of position with patient movement. the pump was explanted and then replaced back to the lv. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of positioning issue was most likely patient movement.
18495 | 1220648-2024-19280 | 20371965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20371965&pc=OZD | 05/14/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/24/2022 | 10/04/2024 |  |  | 02/29/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191539 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had bleeding and limb ischemia. the impella device was surgically removed for transition to intra-aortic balloon pump (iabp). the surgeon confirmed the condition of bleeding, but he said that he could not find a clear bleeding point because hemostasis had already been achieved by manual compression. it was thought that ischemia was the trigger, so percutaneous coronary intervention (pci) was performed on the treatable blood vessels, and mitral valve surgery was considered for future. | [5 empty] | The investigation for the reported access site bleeding and limb ischemia issues have been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the issue was most likely patient condition related since the bleeding was resolved with manual pressure and the patient had disseminated intravascular coagulation (dic) in which the proteins that control blood clotting become overactive. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18496 | 1220648-2024-19277 | 20372000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372000&pc=OZD | 01/25/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/25/2023 | 10/04/2024 |  |  | 12/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176285 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while inserting an impella, a delivery issue occurred. during attempted placement, it was noted that the pump would not cross the iliac due to the right iliac being completely occluded and the left iliac being 90 percent occluded. the pump implant was aborted due to the restrictions caused by the patient's existing anatomy. no further information is available. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
18497 | 1220648-2024-19273 | 20372094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372094&pc=OZD | 01/17/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/17/2023 | 10/04/2024 |  |  | 10/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022147668 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported a failure to deliver the device due to extreme calcium in the aorta along with the tortuous anatomy. insertion was aborted with no reported harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issues was patient condition related, specifically the significant calcification and tortuosity in the patient's vessels. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18498 | 1220648-2024-19278 | 20372101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372101&pc=OZD | 01/25/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/25/2023 | 10/04/2024 |  |  | 10/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283482 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | Migration | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp came out of position, this coincided with placement of the extra-corporeal membrane oxygenation cannula. the team was unable to redeliver the impella cp to the left ventricle and the decision was made to explant the impella cp. no lasting harm or injury reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the positioning issue was use related and related to pulling back of the device during ecmo cannula placement. this report is being filed as part of a retrospective review of historical records.
18499 | 1220648-2024-19274 | 20372187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372187&pc=OZD | 01/25/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/25/2023 | 10/04/2024 |  |  | 11/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023280651 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after a successful implant the device was dislocated during computer tomography exam, repositioning was not possible, the device was explanted and new pump used. no reported harm done to the patient. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the positioning issue was use relate (tuohy borst valve not correctly secured). this report is being filed as part of a retrospective review of historical records.
18500 | 1220648-2024-19283 | 20372246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372246&pc=OZD | 09/18/2024 | 11/19/2024 |  | 10/04/2024 |  |  | 09/18/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506722 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that the procedure to insert the impella cp was aborted. dilator was reinserted and attempted to advance further without any luck. the process was repeated with 6mm and 7mm balloon with no success in advancing further. the patient¿s outcome is unknown. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
18501 | 1220648-2024-19281 | 20372339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372339&pc=OZD | 01/26/2023 | 03/27/2025 |  | 10/04/2024 |  |  | 01/26/2023 | 10/04/2024 |  |  | 09/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274580 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was bleeding from the access site. systematic heparin was withheld. additionally, the motor was not running properly and flows never exceeded 0.9 l/min. upon taking the proper actions of turning the pump down, restarting multiple times, the pump then started running. flows achieved normal levels, but the motor current exceeded 1200+. the issue was unable to be resolved. decision was made to replace the pump with a new impella cp. the patient was reported stable post replacement. no further informtion is available. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported saturated flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the device passed all post-sterile inspection checks. the evaluation of the device noted that the field real time (rt) log was reviewed and spikes in motor current, suction, pump position in ventricle alarm, and trending purge pressure and motor current were observed. the returned pump was visually inspected, and biomaterial was observed in the purge gap. the cause of the issue was biomaterial in the motor. this report is being filed as part of a retrospective review of historical records.
18502 | 1220648-2024-19284 | 20372487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372487&pc=OZD | 01/27/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/27/2023 | 10/04/2024 |  |  | 10/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024287651 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing high-risk surgery was attempted to be implanted with an impella cp device for mechanical circulatory support. the pump was inserted via the axillary artery. a calcified lesion in the subclavian artery prevented the pump from being advanced into the aorta. after multiple failed attempts to pass the calcified lesion, the case was aborted. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause was patient condition because a calcified lesion in the subclavian artery prevented the pump from being advanced. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18503 | 1220648-2024-19288 | 20372507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372507&pc=OZD | 03/07/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 03/07/2023 | 10/04/2024 |  |  | 04/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210698 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, it was noted that there was thrombus on the catheter in the descending aorta. due to bleeding complications, coagulation optimization was kept very short and below recommendations. no intervention was required. impella was successfully weaned and patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was most likely improper anticoagulation management, specifically the patient's act was below the recommended therapeutic range. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18504 | 1220648-2024-19285 | 20372520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372520&pc=OZD | 01/27/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/30/2023 | 10/04/2024 |  |  | 07/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251679 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing aortic valvuloplasty was implanted with an impella cp device for mechanical circulatory support. during support, the pump exhibited low flows. the team attempted to reposition, but the low flows remained so the team explanted and replaced the pump. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the low pump flow was most likely the calcified valve patient condition resulting in a cannula kink and guidewire damage. this report is being filed as part of a retrospective review of historical records.
18505 | 1220648-2024-19286 | 20372532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372532&pc=OZD | 01/27/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/27/2023 | 10/04/2024 |  |  | 10/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024286831 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient of unknown race/gender noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. during implantation, the team was unable to gain femoral access with the 14fr introducer due to the patient's peripheral artery disease. the team aborted this attempt, and it was not specified how they continued. there was no reported patient harm. | [5 empty] | The investigation for the reported difficult insertion has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the difficult insertion was patient anatomy, most likely due to patient condition. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18506 | 1220648-2024-19287 | 20372613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372613&pc=OZD | 01/29/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/29/2023 | 10/04/2024 |  |  | 08/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269805 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump flow was saturated post insertion/initiation. the pump was removed and replaced with no reported harm to the patient. | [5 empty] | The investigation into the impella cp pump saturated flow has been completed. the impella pump was returned for investigation. the ltlog showed a spike in motor current and placement signal. after conducting a product evaluation on the device, it was determined that there was biomaterial in the purge gap. the root cause of the saturated flow is most likely biomaterial. this report is being filed as part of a retrospective review of historical records.
18507 | 1220648-2024-19291 | 20372617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372617&pc=OZD | 03/09/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 03/09/2023 | 10/04/2024 |  |  | 03/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203605 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, impella cp pump stopped. per the physician, the pump would not restart after being taken out of surgical mode. the automated impella controller (aic) was swapped in an attempt to troubleshoot the issue, however the pump remained inactive. the pump was subsequently removed, and an intra-aortic balloon pump was inserted to maintain patient support. there was no reported patient harm. | [5 empty] | The investigation for the reported pump stop issue has been completed. the data logs confirm a high motor current pump stop occurred after the pump had been in surgical mode and that the pump was unable to restart. the product was returned and the outflow cage was found to be deformed and contacting the blades of the impeller. the deformation likely occurred when clamping down on the pump. the root cause of the pump stop was determined to be use related as clamping down on the pump deformed the outflow cage resulting in a high motor current pump stop. as noted in the impella ifu: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18508 | 1220648-2024-19290 | 20372663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372663&pc=OZD | 03/07/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 03/07/2023 | 10/04/2024 |  |  | 10/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269518 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. during attempted implant, the device due to "dissociation "was only able to be inserted in the false lumen of the vessel. the insertion was aborted, and the patient was supported by the prior placed extracorporeal membrane oxygenation (ecmo). no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the issue was not determined since product was not returned. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18509 | 1220648-2024-19282 | 20372691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372691&pc=OZD | 01/26/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 01/26/2023 | 10/04/2024 |  |  | 03/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206318 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. during implantation, the device was unable to reach the heart due to the calcified and tortuous anatomy of the patient's aorta. the pump insertion was aborted, and an intra-aortic balloon pump was used instead. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause in the impella delivery issue was patient condition due to calcified and tortuous anatomy from the ascending aorta to the descending aorta. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18510 | 1220648-2024-19289 | 20372710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372710&pc=OZD | 05/24/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/24/2022 | 10/04/2024 |  |  | 11/30/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169213 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp device had low flows of only 0.5 l/min. and suction events. the activated clotting time (act) was 230 and an additional 1,000 units of heparin was given. it was decided to explant the device and replace with a new impella pump. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the pump was visually inspected and connected. suction alarm reproduced and low flow observed .5 l/min. no kink observed. biomaterial around impella blades observed which sucked from patient when pump started. the root cause of the low pump flow was determined to be biomaterial. this report is being filed as part of a retrospective review of historical records.
18511 | 1220648-2024-19296 | 20372717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372717&pc=OZD | 03/14/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 03/14/2023 | 10/04/2024 |  |  | 10/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283248 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced access site was bleeding and there was non-pulsatility in both feet. no further information was provided regarding intervention or patient outcome. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia could not be determined. the investigation for the reported access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on the clinical details, the root cause of the access site bleeding is most likely due to use. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ limb ischemia: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
18512 | 1220648-2024-19292 | 20372726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372726&pc=OZD | 02/27/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 02/27/2023 | 10/04/2024 |  |  | 07/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254361 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a loud noise was heard from the device after priming. per the physician, the catheter was placed on the guidewire while priming. upon completion, and prior to patient insertion, a loud noise was heard. the physician opted to use another device as a result of the noted issue. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
18513 | 1220648-2024-19293 | 20372777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372777&pc=OZD | 09/13/2021 | 10/04/2024 |  | 10/04/2024 |  |  | 09/14/2021 | 10/04/2024 |  |  | 06/30/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022106999 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. during implantation, the medical staff was unable to pass the guide wire through the easy lumen and the guidewire could not be removed with manual effort. the decision was made to use a new impella pump. the device was never entered into the body and never turned on and was discarded due to covid precautions. | [5 empty] | A.5 is unknown. the investigation for the reported difficulty inserting/removing introducer issue has been completed. the impella device was discarded by the customer due to covid precautions and clinical information provided was sufficient to determine the root cause. the root cause of the guidewire issues was issues backloading the guidewire through the red lumen. the exact cause of this was not determined as product was not returned. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18514 | 1220648-2024-19295 | 20372783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372783&pc=OZD | 03/13/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 03/14/2023 | 10/04/2024 |  |  | 11/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024301433 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient of unknown race/ethnicity noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the flows were found to be inadequate, unable to run at p2 without suction alarms. the device was repositioned several times, but the problem persisted. the physician felt that the line was kinked and decided to replace the device for ongoing support after the procedure was finished. there was no reported patient harm. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the low pump flow could not be determined. this report is being filed as part of a retrospective review of historical records.
18515 | 1220648-2024-19294 | 20372977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372977&pc=OZD | 03/11/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 03/11/2023 | 10/04/2024 |  |  | 11/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024301340 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a low flow issue involving an impella pump was reported. during patient support, an impella in ventricle alarm was triggered prompting repositioning of the pump. however, once the impella was pulled back into position, the pump flow was notably diminished and a spike in motor current was observed. as a result, the pump was explanted and replaced with another impella cp pump. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. base on the clinical data root cause of the low pump flow issue was biomaterial (thrombus) ingestion due to patient condition. complaint device not returned for investigation. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
18516 | 1220648-2024-19298 | 20372994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20372994&pc=OZD | 03/15/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 03/15/2023 | 10/04/2024 |  |  | 11/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024301822 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was reported that there were suction alarms at every performance level. the team also attempted to reposition the pump multiple times, but the issue persisted. the pump was removed and replaced with another impella, there was no reported harm to the patient. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the low pump flow could not be determined. this report is being filed as part of a retrospective review of historical records.
18517 | 1220648-2024-19297 | 20373040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20373040&pc=OZD | 02/23/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 02/23/2022 | 10/04/2024 |  |  | 11/30/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169052 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump stopped after approximately 8 days of use. a new impella cp device was used with no harm done to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. the cause of the pump stop was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
18518 | 1220648-2024-19299 | 20373582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20373582&pc=OZD | 03/15/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 03/15/2023 | 10/04/2024 |  |  | 10/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272120 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was inserted via a reaccess sheath, but there were severe bleeding issues at the access site. all attempts to stop the bleeding were unsuccessful, so the pump was explanted. the patient survived this procedure. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the access site bleeding could not be determined. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18519 | 1220648-2024-19303 | 20373604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20373604&pc=OZD | 09/19/2024 | 12/06/2024 |  | 10/04/2024 |  |  | 09/19/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529279 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, bleeding was noted at the groin. the patient was given one unit of red blood cells and dressing was applied. the patient survived the event. | [5 empty] | The device has not been received for evaluation. upon completion of the investigation, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, dressing change minimized the bleeding as per the clinical description provided. corrections to the initial mfr report: d4-udi number.
18520 | 1220648-2024-19308 | 20373923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20373923&pc=OZD | 05/26/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/26/2022 | 10/04/2024 |  |  | 02/29/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195415 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were low flows. the team made the choice to remove and replace the pump. the second new pump successfully supported the patient. there was no patient harm reported. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the data logs were analyzed and confirm the suction alarms. the root cause of the low flow was unable to be determined as no product or imaging was returned for this investigation. this report is being filed as part of a retrospective review of historical records.
18521 | 1220648-2024-19311 | 20374394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20374394&pc=OZD | 05/26/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/26/2022 | 10/04/2024 |  |  | 01/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182813 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump exhibited lower flows would not increase more than 0.5 and continued to get suction alarms. the device was removed and replaced without any harm. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18522 | 1220648-2024-19304 | 20374441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20374441&pc=OZD | 12/14/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 12/15/2022 | 10/04/2024 |  |  | 08/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263742 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. during placement, the physician was unable to fully deliver the impella cp due to the patient's vascular anatomy. the pump was removed and a second impella cp was delivered with contralateral leg access. there was no reported harm to the patient. the patient survived the explant. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was discarded by the customer. based on the clinical details, the patient had many health condition problems. the root cause was patient condition due to her anatomy. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18523 | 1220648-2024-19302 | 20374636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20374636&pc=OZD | 12/13/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 12/13/2022 | 10/04/2024 |  |  | 04/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217194 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute decompensated chronic cardiomyopathy was implanted with an impella cp for mechanical circulatory support. shortly after beginning support, the impella cp exhibited suction and low flows. the pump was explanted and replaced. there was no reported harm to the patient. the patient survived the explant. | [5 empty] | The investigation for the reported low pump flow has been completed. clinical details and data logs were reviewed. the issue cause could not be determined definitively without the product. the root cause could not be determined since the complaint device was not returned for investigation. this report is being filed as part of a retrospective review of historical records.
18524 | 1220648-2024-19300 | 20374925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20374925&pc=OZD | 03/16/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 03/16/2023 | 10/04/2024 |  |  | 10/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272785 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. after insertion of the sheath, the pump experienced high friction and became impassable in the final third, causing damage to the sheath. the impella was then pulled back, which also required excessive force. the physician did not attempt entry on the left side due to patient's anatomy, and the procedure was aborted. there was no reported harm to the patient. | [5 empty] | The device was returned, and an investigation for the reported delivery issue has been completed. a crack in the mid portion of the sheath on the score line was observed. there was also kinking on the sheath. no abnormalities were observed on the pump. the dilator was curved at the end and there were scratch marks on the dilator. scratch marks indicate the likely presence of hard calcium in the patient's vessels. the cause of the delivery issue was patient condition related. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18525 | 1220648-2024-19306 | 20374976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20374976&pc=PYX | 09/19/2024 | 11/20/2024 |  | 10/04/2024 |  |  | 09/19/2024 | 10/04/2024 |  |  |  |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025487266 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Failure to Sense | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp flex support. while on support, placement signal and flow issues were noted. patient became symptomatic with lower flows. impella rp flex was removed due to low pump flows. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation placement signal issue has been completed. the pump was returned for investigation and passed all electrical testing. no biomaterial was found on the impeller or in the purge gap. additionally, no issues were noted with the optical signal parameters or external damage to the dp sensor. data logs show an upward placement signal with psnr alarms after 5 hours of case run, along with a loss of pump flow accompanied by an upward trend, with no noted trend in motor current. the cause of the placement signal issue was dp sensor drift. h6 medical device problem code 1248 was reported in error on manufacturer device report 1220648-2024-19306. pump flow is calculated based on the dp sensor reading. field use does not indicate true low flows but rather results from the dp sensor drift due to malfunction. the low pump flow failure mode was removed, and only keep the placement signal issue. the cause of the placement signal issue was dp sensor drift.
18526 | 1220648-2024-19310 | 20375229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20375229&pc=OZD | 02/25/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 02/25/2022 | 10/04/2024 |  |  | 08/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022135787 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 20 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the complainant reported pump stopped during use of the device. the pump was removed, and the patient was managed medically without the need for an additional impella. no patient harm was reported. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the pump stop could not be determined. this report is being filed as part of a retrospective review of historical records.
18527 | 1220648-2024-19334 | 20375330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20375330&pc=OZD | 02/11/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 02/11/2023 | 10/04/2024 |  |  | 09/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023277111 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the impella cp was placed with a flow rate of approximately 3.5 liters per minute (l/m). reportedly, the flows suddenly dropped to 0.61 l/m and flows were unable to be increased. the placement of this impella was aborted due to concerns that a clot may have been ingested. however upon removal, a clot was not visible on the cannula. the pump was replaced with another impella. | [5 empty] | A.5 race is unknown. the investigation into the low pump flow issue has been completed. the impella pump was returned for investigation. a visual inspection of the pump revealed biomaterial on the impeller. the pump was run at p-9 in water and the low pump flow was reproduced. the pump was re-examined after the run and it was found that the biomaterial remained on the impeller. the cause of the low pump flow was biomaterial. this report is being filed as part of a retrospective review of historical records.
18528 | 1220648-2024-19301 | 20375398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20375398&pc=OZD | 05/26/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/26/2022 | 10/04/2024 |  |  | 12/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178116 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, there were lower than expected pump flow during initialization of the device. attempts were made to reposition the pump, but nothing would resolve the issue. the pump continued to support the patient during the procedure without any harm. no additional information provided. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a large amount of biomaterial was found around the impeller. the initial log shows low flow, and the second log shows flow saturation. it is likely at the beginning of the case this was stuck in the cannula, causing low flows. at some point, the biomaterial began to interact with the impeller, causing high motor current and saturated flows. after reviewing the returned pump and the data logs, the root cause of low flow as well as flow saturation was likely due to biomaterial. this report is being filed as part of a retrospective review of historical records.
18529 | 1220648-2024-19307 | 20375525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20375525&pc=OZD | 12/14/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 12/14/2022 | 10/04/2024 |  |  | 09/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273332 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there was bleeding at the access site which was resolved by a purse string stitch. the patient survived the procedure with no further reported complications. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device has not been returned for evaluation. the root cause of the access site bleeding was unable to be determined as the product was not return and insufficient clinical details were provided. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18530 | 1220648-2024-19320 | 20375564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20375564&pc=OZD | 06/06/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 06/06/2022 | 10/04/2024 |  |  | 05/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022103586 |  | 0048-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, throughout the case, there were continuous suction alarms, even after volume and positioning corrections. flows decreased from 3.5 to 1.2 liters per minute in auto mode and guided repositioning under echocardiogram did not resolve the low flows. the decision was made to remove the pump with a flow of 2.5 liters per minute. there was no patient harm reported. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18531 | 1220648-2024-19317 | 20375588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20375588&pc=OZD | 10/21/2021 | 10/04/2024 |  | 10/04/2024 |  |  | 10/21/2021 | 10/04/2024 |  |  | 07/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121387 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported low pump flows followed by a pump stop. the pump was removed and replaced without any adverse impact to the patient. | [5 empty] | The cause of the pump stop was most likely ingested biomaterial per the clinical report. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
18532 | 1220648-2024-19313 | 20375693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20375693&pc=OZD | 05/27/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/31/2022 | 10/04/2024 |  |  | 02/29/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199957 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the outflow cage was crimped while passing up through the aorta. a new impella was placed. there was no patient harm reported. | [5 empty] | The investigation for the reported product damage issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was a most likely damaged outflow cage from excessive force. this report is being filed as part of a retrospective review of historical records.
18533 | 1220648-2024-19343 | 20375719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20375719&pc=OZD | 11/03/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/04/2022 | 10/04/2024 |  |  | 06/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023233914 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, the patient was diagnosed with a cerebral hemorrhage. the hemorrhage is not attributed to the use of impella. patient's condition was stable post procedure. | [5 empty] | The investigation has been completed for the reported issue of stroke. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the cerebral hemorrhage was most likely patient condition related. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18534 | 1220648-2024-19309 | 20375769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20375769&pc=OZD | 02/24/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 02/24/2022 | 10/04/2024 |  |  | 10/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022148779 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with high-risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. during use, a pump stop was reported. upon removal, biomaterial was found in the cannula. there was no patient harm reported. | [5 empty] | The investigation for the reported pump stop issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause of the pump stop issue was ingested biomaterial. this report is being filed as part of a retrospective review of historical records. as noted in impella ifu: impella cp with smartassist for use during cardiogenic shock section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
18535 | 1220648-2024-19315 | 20375836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20375836&pc=OZD | 05/28/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/28/2022 | 10/04/2024 |  |  | 02/29/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193615 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump was started, and low flows were experienced. the impella was removed. there was no patient harm reported. | [5 empty] | D6 implant and explant date unknown. the investigation for the reported low pump flow issue has been completed. data logs showed there were recurrent "suction" alarms upon pump activation; the alarms were accompanied by low pump flow and low motor current pulsatility. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18536 | 1220648-2024-19312 | 20375857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20375857&pc=OZD | 05/27/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/27/2022 | 10/04/2024 |  |  | 11/30/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168469 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, lower than expected flows occurred due to a suspected thrombus within the pump. the device was removed and replaced without any harm to the patient. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs were returned and confirm low flow and suction. the product was returned, and a thrombus was found in the inflow and the outflow of the pump. the root cause of the low flow was determined to be biomaterial. this report is being filed as part of a retrospective review of historical records.
18537 | 1220648-2024-19314 | 20375862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20375862&pc=OZD | 10/09/2021 | 10/04/2024 |  | 10/04/2024 |  |  | 10/11/2021 | 10/04/2024 |  |  | 08/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2022120482 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The pump stopped, another device was used, and there was no harm to the patient. | [5 empty] | The investigation into the pump stop is complete. the device was returned for evaluation and the field rt data log was reviewed and multiple impella stopped ¿ motor current high alarms were confirmed on 09oct2021 at case start associated with motor current spiking to 1500ma. visual inspection of the returned pump revealed that the red lumen was still present within the outflow cage, obstructing the impeller, and extending out of the pigtail. the red lumen flag/label was also present. the root cause of the pump unable to start was failing to remove the red lumen. refer to es2019-234 for further information. this report is being filed as part of a retrospective review of historical records.
18538 | 1220648-2024-19355 | 20375905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20375905&pc=OZD | 11/28/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/28/2022 | 10/04/2024 |  |  | 07/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023255780 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an acute patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, a purge pressure rise alarm occurred on the 6th day of pump operation. it did not improve even after replacing the cassette. the impella was removed and switched to intra-aortic balloon pump. no further information was provided. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was completed. a review of the device log confirms high purge pressure alarms and shows that the purge system blocked alarms. a suction event is also noted at the time that the purge flow begins to drop. a short time after the purge pressure begins to steadily rise. a visual inspection of the pump did not reveal any abnormalities. there was no biomaterial observed on initial inspection and no kinks. the pump was attached with the returned purge cassette and run in the lab. the high purge pressure alarm was not reproduced. a pump tear down was then completed, and biomaterial was found within the motor. the cause of the high purge pressure was biomaterial. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18539 | 1220648-2024-19329 | 20375915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20375915&pc=OZD | 04/04/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 04/04/2023 | 10/04/2024 |  |  | 12/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303394 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella access site. the patient received one unit of packed red blood cells to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of access site bleeding - major. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the access site bleeding - major was unable to be determined as limited clinical details were provided and no product was returned for review. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18540 | 1220648-2024-19305 | 20375939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20375939&pc=OZD | 05/26/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 05/26/2022 | 10/04/2024 |  |  | 11/30/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170069 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was accidentally pulled back to where only the pigtail was across the valve. after many attempts in several minutes, the physician was able to get device back into left ventricle. but after that, suction was present all the way down to p2 and low flows of 1-1.5 l present. 1l of volume given to help with suction, but did not change. the pump was explanted and another impella cp was inserted. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that upon inspection of the pump, a cannula kink was found. the kink most likely occurred during the repositioning of the pump. the low pump flow was most likely due to improper positioning and the repositioning of the pump after it was pulled back with only the pigtail across the aortic valve. the root cause of the issue was wireless re-access. the impella device is not indicated for wireless re-access. this report is being filed as part of a retrospective review of historical records.
18541 | 1220648-2024-19323 | 20376102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376102&pc=OZD | 03/27/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 03/27/2023 | 10/04/2024 |  |  | 10/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290037 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the patient lost pulsatility several times. the impella was explanted, and the patient survived the procedure. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the limb ischemia was most likely patient condition. as noted in the impella ifu: limb ischemia: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
18542 | 1220648-2024-19318 | 20376229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376229&pc=OZD | 06/04/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 06/04/2022 | 10/04/2024 |  |  | 02/29/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198911 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Fluid/Blood Leak | Increase in Pressure |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the system alarmed high purge pressure purge system blocked, and purge pressure was not present on the console. to resolve the pressure issue, the cassette was changed resulting in the console starting to purge the system open. during changing of the cassette, a hole, purge leak, the size of a pin was noted in the cassette tubing, the end user discarded the tubing. additionally, it was reported a kink was noted and the device collapsed on itself, the physician was not comfortable repositioning by pulling the device back because it was tugging at the aorta, the physician was unable to successfully recross the valve with the rewire access port. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation for the reported high purge pressure, purge leak, and positioning issues have been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the high purge pressure and purge leak was unable to be determined as no product was returned. however, clinical details provided were sufficient to determine a root cause of the positioning issue. the root cause of the positioning issue was determined to be use related as the pump was pulled back across the valve. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi section: cleaning as noted in the impella ifu ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18543 | 1220648-2024-19356 | 20376236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376236&pc=OZD | 11/28/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/28/2022 | 10/04/2024 |  |  | 08/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269955 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the device exhibited suction alarms indicating low flow. the physician manipulated the device by advancing and retracting, however this did not resolve the issue. the physician used a fluoroscope and observed that the cannula distal to the motor was kinked. the physician decided to explant the device and replace with another impella cp. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was possible. clinical details provided were sufficient to determine a root cause. the pump was returned and a kink in the cannula was confirmed. the clinical details confirm that the kink was present after the insertion of the device. the root cause of the low flow was most likely patient condition related as the patient's anatomy likely caused the cannula to kink during insertion. this report is being filed as part of a retrospective review of historical records.
18544 | 1220648-2024-19325 | 20376289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376289&pc=OZD | 04/03/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 04/03/2023 | 10/04/2024 |  |  | 06/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248017 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, the patient experienced significant bleeding from the right groin impella cp site, requiring the application of a femstop at the referring facility, which resulted in mottling of the leg. several attempts at angle matching were unsuccessful in achieving hemostasis. manual pressure was maintained for over an hour; however, any release of pressure resulted in renewed bleeding from the site. the physician subsequently applied a hemostatic suture and adjusted the skin tension to mitigate further bleeding. the patient was escalated to a 5.5. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
18545 | 1220648-2024-19358 | 20376313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376313&pc=OZD | 11/29/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/29/2022 | 10/04/2024 |  |  | 07/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022123391 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  |  | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the physician attempted to reposition the device and caused a tear in the sterile sleeve. transferred to another facility for escalation of care aa the torn sterile sleeve was thought to be an infection risk. support continued with the pump following the event. there was no reported harm to the patient, and it was noted that they survived the procedure. | [5 empty] | The investigation for the reported package damage issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was a damaged sterile sleeve during repositioning due to improper technique. this report is being filed as part of a retrospective review of historical records.
18546 | 1220648-2024-19362 | 20376397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376397&pc=OZD | 12/02/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 12/02/2022 | 10/04/2024 |  |  | 05/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228812 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the pump stopped, the patient was stable so the physcian decided to explant the device. it was noted that the impella cp pump was utilized well beyond indication when it stopped. the cp ran for 28 days. there was no harm or injury reported from the pump stop issue. no additional information required. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. device used beyond indication for use. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
18547 | 1220648-2024-19327 | 20376484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376484&pc=OZD | 04/03/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 04/03/2023 | 10/04/2024 |  |  | 09/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023282769 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. despite multiple attempts, the physician was unable to advance the impella motor through the distal aortic bifurcation. a long impella sheath was introduced, however, it did not pass through the transition. the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of difficulty inserting/removing introducer. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the difficulty inserting/removing introducer was due to patient condition with patient having calcified iliac artery. this report is being filed as part of a retrospective review of historical records.
18548 | 1220648-2024-19335 | 20376529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376529&pc=OZD | 10/28/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 10/28/2022 | 10/04/2024 |  |  | 05/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231054 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on day 12 of support, the physician observed a pump stop. the pump was able to be restarted. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of pump stop. it was noted that the pump was used beyond the 8-day design life. the device was not received for evaluation. the root cause of the pump stop was unable to be determined as no product was returned for analysis. this report is being filed as part of a retrospective review of historical records.
18549 | 1220648-2024-19331 | 20376537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376537&pc=OZD | 11/11/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/11/2022 | 10/04/2024 |  |  | 04/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216249 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. prior to support, an extra-corporeal membrane oxygenation (ecmo) was inserted into the right leg and an intra-aortic balloon pump (iabp) was inserted into the left leg. after iabp removal, the impella cp was inserted using a 14fr sheath. the patient was returned to the intensive care unit and bleeding was observed from the impella cp access site. swelling of the lower extremities and hematoma occurred. the impella was removed and a 16fr sheath was inserted, but the bleeding did not stop. the physician decided to explant the impella cp. it was reported that cardiac function returned to normal, and the bleeding issue was treated with compression. no further information is available. | [5 empty] | The investigation has been completed for the reported issue of access site bleeding. the device was not received for evaluation. however, clinical details provided were sufficient to determine a root cause. per the results of the clinical review and investigation: the physician said the following about the case, "impella should have been used from the beginning instead of ecmo + iabp. since the iabp insertion angle is different from the impella insertion angle, a softer sheath should have been used when switching." no product was returned for investigation. the cause of the access site bleeding was use related and related to the change of angle from inserting the iabp to inserting the impella. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18550 | 1220648-2024-19352 | 20376612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376612&pc=OZD | 11/23/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/23/2022 | 10/04/2024 |  |  | 05/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022084104 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was going to be implanted with an impella cp device for mechanical circulatory support. it was reported the staff was unable to deliver the impella cp due to native anatomy and previously placed iliac stents. the medical staff attempted all troubleshooting and sheath exchanges but no success. the patient's pressure had stabilized on the levophed and the decision was made to abort impella insertion and continue to monitor patient. no harm or injury reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause of the delivery issue was patient condition related. stents in the iliac caused resistance for inserting the impella cp and caused damage to the long sheath and long dilator. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18551 | 1220648-2024-19341 | 20376620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376620&pc=OZD | 11/03/2022 | 03/27/2025 |  | 10/04/2024 |  |  | 11/03/2022 | 10/04/2024 |  |  | 08/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262206 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during the removal of the impella insertion sheath, the medical staff noted significant tension between the sheath and the impella pump. the pump was inadvertently pulled back and was unable to advance across the valve. the impella was explanted and replaced with a new impella cp. the patient survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning issues and difficulty inserting/removing pump through introducer issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that only the pump was returned for investigation. no abnormalities were observed on the returned product. the clinical information does not indicate that a guidewire was used for re-access. the cause of the issue was most likely wireless re-access. the impella device is not indicated for wireless re-access. the cause of the introducer issue was not determined since the introducer was not returned. this report is being filed as part of a retrospective review of historical records.
18552 | 1220648-2024-19322 | 20376655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376655&pc=OZD | 03/03/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 03/03/2023 | 10/04/2024 |  |  | 11/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278864 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited a rise of pressure. after disconnecting the yellow luer connection on the purge line, the pressure went down. the physician decided to explant the device. no further information is available. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the high purge pressure was not determined since neither product nor data logs were returned. this report is being filed as part of a retrospective review of historical records.
18553 | 1220648-2024-19339 | 20376717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376717&pc=OZD | 10/31/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 10/31/2022 | 10/04/2024 |  |  | 08/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259576 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to being implanted with an impella cp device for mechanical circulatory support. the red lumen was removed prior to the guidewire placement. the impella was explanted and replaced with a new impella cp to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the product was not returned. from the clinical review: physician removed the red lumen prior using 0.18" guide wire. the impella was unable to cross the valve in the patient. root cause of the failure is improper technique use. this report is being filed as part of a retrospective review of historical records.
18554 | 1220648-2024-19342 | 20376840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376840&pc=OZD | 11/03/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/03/2022 | 10/04/2024 |  |  | 05/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023222388 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. no purge fluid visible coming from the exhaust basket during priming. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of priming. the device was received for evaluation. there was no priming issue with the pump. the device functioned/operated to specification. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
18555 | 1220648-2024-19345 | 20376844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376844&pc=OZD | 02/18/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 02/18/2023 | 10/04/2024 |  |  | 10/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024294432 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis. it was reported that the patient¿s laboratory results were hemolyzed and the urine output was dark and coke colored. the impella was repositioned under echocardiogram and the following day the urine reportedly cleared, and the laboratory results were not hemolyzed. the patient was a heart transplant patient and was upgraded to the impella 5.0 pending possible transplant. | [5 empty] | A.5 race is unknown. the investigation has been completed . the impella device was not received from the customer. the cause of the issue was not determined since thr product was not returned and insufficient clinical details were provided. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
18556 | 1220648-2024-19328 | 20376874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376874&pc=OZD | 04/04/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 04/04/2023 | 10/04/2024 |  |  | 11/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303090 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the team observed a kink while positioning, consequently, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. no patient harm was reported due to the issue. | [5 empty] | The investigation has been completed for the reported issue of product damage (cannula kink). the device was not received for evaluation. the root cause of product damage (cannula kink) was most likely a kinked cannula. this report is being filed as part of a retrospective review of historical records.
18557 | 1220648-2024-19353 | 20376911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376911&pc=OZD | 11/25/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/25/2022 | 10/04/2024 |  |  | 06/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236870 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported upon insertion and pump start, flows were low. on auto, "impella flows reduced". multiple attempts were made to reposition in central venous line (cvl), but still getting same alarms. flows 1.0-1.8 despite adequate volume and moderate rv function. no reported harm to the patient. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause of the low pump flow issue was most likely improper positioning of the device deep in the left ventricle (lv) resulting in a cannula kink. this report is being filed as part of a retrospective review of historical records.
18558 | 1220648-2024-19350 | 20376991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20376991&pc=OZD | 11/20/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/20/2022 | 10/04/2024 |  |  | 06/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237687 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the pump failed advance over the 0.018" guidewire during delivery. a kink was noted in the pigtail. a new pump was used with no harm to the patient reported. | [5 empty] | The investigation for the reported product damage (cannula kink) issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause of the product damage (cannula kink) issue was most likely use related since the pump accidentally pulled out of the left ventricle (lv) during delivery and resulted in pigtail kink. this report is being filed as part of a retrospective review of historical records.
18559 | 1220648-2024-19364 | 20377023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377023&pc=OZD | 12/05/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 12/05/2022 | 10/04/2024 |  |  | 07/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257143 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, continuous bleeding was observed from the impella site despite manual pressure, fem stop, new suture and dressing. the resolution for this issue is unknown. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the access site bleeding was most likely a lack of vessel recoil onto the impella repositioning sheath attributed to patient condition. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18560 | 1220648-2024-19333 | 20377041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377041&pc=OZD | 10/27/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 10/27/2022 | 10/04/2024 |  |  | 07/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251678 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, saturated flows were observed. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of saturated flow. the device was returned for evaluation. the root cause of the saturated flow was determined to be biomaterial. this report is being filed as part of a retrospective review of historical records.
18561 | 1220648-2024-19316 | 20377235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377235&pc=OZD | 06/01/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 06/01/2022 | 10/04/2024 |  |  | 12/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176021 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient n acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was successfully placed, however upon removal of the peel away sheath the pump was accidentally advanced into the left ventricle, which resulted in the outlet cage advancing completely in the ventricle with kinking of the cannula below the motor housing, causing continuous suction and low pump flows. the kink was unable to be removed by pulling back on the catheter and repositioning. the physician made the decision to remove this impella and place a new impella. there was no patient harm reported, and this device was replaced. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the low flow was determined to be a cannula kink caused by improper positioning. the pump came out of positioning when placing the repo unit which is a use related issue. this report is being filed as part of a retrospective review of historical records.
18562 | 1220648-2024-19357 | 20377244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377244&pc=OZD | 11/28/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/28/2022 | 10/04/2024 |  |  | 05/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230901 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a positioning issue. it was documented that the white cable got caught during the patient's transfer to their bed and the impella was subsequently pulled back out of position. multiple attempts were made to reposition the pump, however repositioning ultimately failed as the pigtail had gotten caught in the chordae. the patient was noted to be stable with minimal support and the decision was made to remove the pump. there was no harm to the patient as a result of the noted issue and it was reported that they survived the procedure. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. after reviewing the clinical details: the positioning issue occurred after initial placement of impella. the clinical staff accidentally pulled the impella back while moving the patient to bed. several attempts were made to reposition the device without success and the physician decided to explant the device. the root cause is patient movement. this report is being filed as part of a retrospective review of historical records.
18563 | 1220648-2024-19319 | 20377257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377257&pc=OZD | 10/09/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 10/09/2022 | 10/04/2024 |  |  | 04/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224003 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with unknow demographics was implanted with an impella cp device for mechanical circulatory support. it was reported the impella cp pump flow suddenly dropped sharply. the physician decided to remove the impeller. the impella device was successfully removed, and a clot was noticed in impeller, after removal. no patient harm or injury was noted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation, but clinical details were sufficient to determine the root cause of the issue was biomaterial. this report is being filed as part of a retrospective review of historical records.
18564 | 1220648-2024-19337 | 20377271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377271&pc=OZD | 10/28/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 10/28/2022 | 10/04/2024 |  |  | 07/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257276 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was going to be implanted with an impella cp device for mechanical circulatory support. the complainant reported that the physician was unable to implant the device. it was reported that this issue occurred due to the patient's anatomy, therefore, the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
18565 | 1220648-2024-19324 | 20377302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377302&pc=OZD | 04/02/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 04/02/2023 | 10/04/2024 |  |  | 11/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303983 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with unknown race was implanted with an impella device for mechanical circulatory support. a clot was visualized on echocardiography, and the placement signal alarms were determined to be inaccurate due to low flow volume, with suction alarms also noted at p2. the patient began to decompensate. consequently, the decision was made to explant the device and replace it with a new impella cp device. no patient harm was reported. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was returned for evaluation. the root cause of the low pump flow was determined to be was most likely ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
18566 | 1220648-2024-19347 | 20377410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377410&pc=OZD | 11/21/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/21/2022 | 10/04/2024 |  |  | 09/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259811 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported days after impella cp removal, the patient experienced retroperitoneal bleeding. two blood products were administered. the physician noted no causal relationship between the impella procedure and bleeding, referred to multimorbid intensive care unit (icu) patient as contributing factor. no further information is available. | [5 empty] | The investigation for the reported access site bleeding issues has been completed. the impella device has not been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the patient injury was most likely patient condition related and related to the patient's multiple chronic conditions. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18567 | 1220648-2024-19330 | 20377639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377639&pc=OZD | 11/06/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/06/2022 | 10/04/2024 |  |  | 07/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252385 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient grabbed/pulled catheter. flows decreased and echo was performed at bedside. it was observed that the catheter was looped/kinked. multiple attempts to reposition were unsuccessful. the impella was explanted and replaced with a new impella cp. the patient survived the procedure. | [5 empty] | The investigation for the reported low flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. a review of the data logs was performed and show a decrease in motor current and motor speeds. additionally, many suction alarms were observed. based on the clinical details, the position of the pump had moved due to patient movement. the root cause is most likely improper position of the pump, this is based on the changes observed in flow and motor current in the data logs. this report is being filed as part of a retrospective review of historical records.
18568 | 1220648-2024-19321 | 20377692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377692&pc=OZD | 11/05/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/05/2022 | 10/04/2024 |  |  | 04/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221234 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to being implanted with an impella cp device for mechanical circulatory support. during delivery, the device got kinked in the aortic arch and the wire was out, the physician was unable to overcome the kink issue. the device was explanted, and the kink was visible. the physician decided to implant a replacement impella cp. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely use related as the wire was removed before the pump had crossed the valve. wireless re-access is not permitted per ifu 0048-9007. impella cp with smartassist system. instructions for use for the related event are as follows: section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18569 | 1220648-2024-19326 | 20377698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377698&pc=OZD | 04/03/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 04/03/2023 | 10/04/2024 |  |  | 04/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231852 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on day 15 of support, a sudden pump stop occurred. normal troubleshooting attempts to restart the device were unsuccessful. the patient continued on support until the device was successfully weaned. no patient harm was reported. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the device was not received for evaluation. the root cause of the pump stop was unable to be determined as no product or logs were returned for analysis. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
18570 | 1220648-2024-19340 | 20377711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377711&pc=OZD | 10/31/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/04/2022 | 10/04/2024 |  |  | 04/30/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211949 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced hemolysis. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of hemolysis. the device was not received for evaluation. the root cause of the hemolysis was unable to be determined since product was not returned. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. this report is being filed as part of a retrospective review of historical records.
18571 | 1220648-2024-19349 | 20377767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377767&pc=OZD | 11/21/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/21/2022 | 10/04/2024 |  |  | 05/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223065 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was going to be implanted with an impella cp device for mechanical circulatory support. it was reported that the physician made attempts to insert the impella cp device. delivery failed and the impella cp kinked with attempted insertion. the impella cp device was removed, and the team adjusted the graft and replaced with an impella 5.5 device. no patient harm or injury noted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the issue was patient condition related. this report is being filed as part of a retrospective review of historical records.
18572 | 1220648-2024-19348 | 20377773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377773&pc=OZD | 11/19/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/19/2022 | 10/04/2024 |  |  | 05/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229473 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the impella cp exhibited low pump flows. the device was explanted, and the patient remained on extracorporeal membrane oxygenation (ecmo) support. no harm or injury reported, no further information is available. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has not been returned for evaluation. the root cause of the low pump flow was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
18573 | 1220648-2024-19360 | 20377775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377775&pc=OZD | 12/02/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 12/02/2022 | 10/04/2024 |  |  | 05/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227894 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the device exhibited low flow. the physician decided to explant the device and replace it with a new impella. it was reported that the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. this report is being filed as part of a retrospective review of historical records.
18574 | 1220648-2024-19361 | 20377790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377790&pc=OZD | 12/02/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 12/02/2022 | 10/04/2024 |  |  | 07/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257116 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. while on support, the device exhibited low purge flow and subsequent high purge pressure alarm. it was discovered that the patient was not administered heparin, and the customer was aware that heparin should be left in the purge while on impella support, the medical staff administered the patient with argatroben to help resolve the issue. the physician decided to explant the device early as a precaution. there was no reported harm to the patient and the patient survived the procedure. | [5 empty] | The investigation for the reported high purge pressure issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, a review of the data logs found the following: after the removal of heparin from the purge system the purge pressure began to gradually rise. the root cause of the high purge pressure was most likely biomaterial buildup due to the lack of heparin. clinical details provided were sufficient to determine a root cause. the cause of the issue was biomaterial. this report is being filed as part of a retrospective review of historical records.
18575 | 1220648-2024-19332 | 20377793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377793&pc=OZD | 12/15/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 12/15/2022 | 10/04/2024 |  |  | 08/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272894 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, there was oozing from the impella access site. the team attempted angle matching, forward tension, securing the blue pad, and manual pressure but the issue persisted. the pump was then removed and there was no further intervention or patient harm reported. the patient survived the explant. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device was discarded by the customer. however, clinical details provided were sufficient to determine a root cause. the cause of the access site bleeding was most likely a lack of vessel recoil onto the sheath since there was oozing despite following best practices for access site management. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18576 | 1220648-2024-19359 | 20377861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377861&pc=OZD | 12/01/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 12/01/2022 | 10/04/2024 |  |  | 08/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269852 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp stopped after just 18 minutes of support. the pump was explanted and replaced with another impella cp. no lasting harm or injury from interruption in the support and pump replacement. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the pump stop and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18577 | 1220648-2024-19336 | 20377953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377953&pc=OZD | 10/28/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 10/28/2022 | 10/04/2024 |  |  | 07/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257492 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump started outside of patient body by accident. consequently, the decision was made to replace it with another imeplla cp. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of packaging issues- sterility. the device was not received for evaluation. the root cause of the product damage was use related per the clinical account. this report is being filed as part of a retrospective review of historical records.
18578 | 1220648-2024-19338 | 20377974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377974&pc=OZD | 02/18/2023 | 10/04/2024 |  | 10/04/2024 |  |  | 02/18/2023 | 10/04/2024 |  |  | 01/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176204 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported that the patient started having hematuria signs with down streaming platelets. the clinician recommended changing the position of the impella and volume status as soon as possible. no further information regarding repositioning was noted. the patient's urine cleared. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
18579 | 1220648-2024-19344 | 20377999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20377999&pc=OZD | 11/04/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/04/2022 | 10/04/2024 |  |  | 10/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154237 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on day 11 of support, the physician observed a pump stop. the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the device was not received for evaluation. it was noted that the clinicians were made aware that the impella is not designed for this length of support. the root cause of the pump stop was unable to be determined as no product was returned for analysis. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
18580 | 1220648-2024-19346 | 20378014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378014&pc=OZD | 11/18/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/18/2022 | 10/04/2024 |  |  | 07/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259534 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported an impella cp was kinked and exhibited low flows. the physician was unable to gain appropriate flows with multiple attempts of repositioning under echocardiogram and fluoroscope. the pump was replaced with another impella cp and there was no report of harm or injury to the patient. | [5 empty] | The investigation for the reported low pump flow and positioning/kinked cannula issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the cause of the reported issues. no logs are available on constellation and no logs were returned from the field. a visual inspection of the pump revealed a kink in the cannula. no other abnormalities were observed. based on the provided clinical information, the cannula kink was caused by the patient's anatomy. the cause of the low pump flow was therefore patient condition related. this report is being filed as part of a retrospective review of historical records.
18581 | 1220648-2024-19351 | 20378076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378076&pc=OZD | 11/23/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/23/2022 | 10/04/2024 |  |  | 06/30/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022103369 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was going to be implanted with an impella cp device for mechanical circulatory support. it was reported the device was unable to be delivered due to the patient's anatomy. despite multiple attempts to insert, the device would not advance past the iliac stents and insertion was aborted. the impella device was removed. there was no reported harm to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the cause of the issue was patient condition related. stents in the iliac caused resistance for inserting the impella cp. instructions for use for the related event are as follows: impella cp with smartassist system section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
18582 | 1220648-2024-19354 | 20378101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378101&pc=OZD | 11/25/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 11/25/2022 | 10/04/2024 |  |  | 07/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252180 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported an impella cp device had low flow, due to a suspected clot entrapment. the pump was successfully exchanged for a new impella cp, with no reported harm to the patient. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the root cause of the low pump flow issue was due to improper positioning of the device deep in the ventricle, resulting in a kinked cannula and low pump flow. this report is being filed as part of a retrospective review of historical records.
18583 | 1220648-2024-19363 | 20378528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378528&pc=OZD | 12/05/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 12/05/2022 | 10/04/2024 |  |  | 08/31/2024 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023251179 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was discovered that a detached easyguide lumen remained in the cannula and prevented the impella from starting. the detached component and pump were removed, and a new pump was placed for patient support. there was no reported harm to the patient. | [5 empty] | The investigation for the reported pump did not start issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was completed. the product was evaluated, and an intact red lumen was found inside the pump. thus, the root cause of the pump being unable to start is the red lumen not being removed properly. this report is being filed as part of a retrospective review of historical records.
18584 | 1220648-2024-19365 | 20378747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378747&pc=OZD | 02/24/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 02/24/2022 | 10/04/2024 |  |  | 12/31/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170206 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the pump was unable to restart after it stopped. the device was removed and replaced with no harm to the patient. | [5 empty] | The investigation for the reported pump stop issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the cause of the pump stop issue was use related. the pump stop issue was unable to be reproduced and the pump operated to specification. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpiated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18585 | 1220648-2024-19366 | 20378749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20378749&pc=OZD | 02/24/2022 | 10/04/2024 |  | 10/04/2024 |  |  | 02/24/2022 | 10/04/2024 |  |  | 11/30/2023 |  | 10/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169510 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, access site bleed was noted. pump was explanted and replaced with another impella cp pump. hemostasis was achieved and no further intervention was needed. patient survived the procedure. | [5 empty] | The investigation for the reported access site bleeding issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely due to a lack of vessel recoil onto the repositioning sheath. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18586 | 1220648-2024-18993 | 20356627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356627&pc=OZD | 01/19/2023 | 10/02/2024 |  | 10/03/2024 |  |  | 01/19/2023 | 10/03/2024 |  |  | 10/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269150 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the automated impella controller (aic) exhibited the error "impella catheter defective". the medical staff attempted to disconnect and reconnect the cable to the impella catheter and also to aic device, however the error still persisted. the resolution for this issue is unknown, no additional information is available. | [5 empty] | The investigation for the reported issue has been completed. the returned pump was connected to an aic and the pump serial number was recognized and case start was prompted. there was no issue found, the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
18587 | 1220648-2024-18994 | 20356638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356638&pc=OZD | 01/19/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 01/19/2023 | 10/03/2024 |  |  | 09/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274929 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was presented for mechanical circulatory support. the complainant had an impella cp that failed during delivery despite exchanging the sheath. the sheath was caught on iliac plaque and the physician was unable to advance the sheath past the lesion. the impella was inserted, and the physician tried to advance through the patient's iliac/aorta area with no success and support was aborted. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was determined to be patient condition as the patient had peripheral artery disease (pad)/peripheral vascular disease (pvd). as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18588 | 1220648-2024-18995 | 20356642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356642&pc=OZD | 01/19/2023 | 03/06/2025 |  | 10/03/2024 |  |  | 01/19/2023 | 10/03/2024 |  |  | 10/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285211 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp stopped during use of the device while on support. the pump was removed and replaced with another impella cp. it was reported that during removal a clot was discovered in the pump. no further information is available. | [5 empty] | The impella device evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | Device was returned for investigation. biomaterial was found wrapped around the impeller. based on the visual inspection and the data logs showing saturated flows, the root cause of the pump stop was biomaterial. this report is being filed as part of a retrospective review of historical records.
18589 | 1220648-2024-18996 | 20356644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356644&pc=OZD | 01/23/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 01/23/2023 | 10/03/2024 |  |  | 09/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023264397 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support as pre-operation placement. during support, the patient experienced limb ischemia. this was resolved when the venoarterial extracorporeal membrane oxygenation (va ecmo) was removed and blood was sent from the side tube of the 16fr sheath to the anterograde sheath. no further information is available. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the limb ischemia was determined to be due to the use of a larger sheath than what is provided by abiomed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
18590 | 1220648-2024-18997 | 20356665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356665&pc=OZD | 01/19/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 01/19/2023 | 10/03/2024 |  |  | 07/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259496 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was presented for mechanical circulatory support. the complainant reported a delivery issue involving an introducer and impella pump. upon initial insertion, the long 14 fr introducer kept getting caught on an existing stent. a short sheath was subsequently placed, however upon attempted pump implant, the device became stuck on a strut and would not advance. as a result of the noted issue, the implant was aborted. there was no reported patient harm, no additional information is available. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the patient had several health conditions. there was a calcium shelf with a tight stenosis in the aorta not supported to pass the sheathes. root cause was patient condition because of the patient anatomy issue. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18591 | 1220648-2024-18998 | 20356678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356678&pc=OZD | 01/20/2023 | 03/10/2025 |  | 10/03/2024 |  |  | 01/20/2023 | 10/03/2024 |  |  | 09/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273020 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported discovery of a retroperitoneal bleed post explant of the impella cp. one unit of blood product was delivered with no further intervention, and it was reported that the patient survived the procedure. | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18592 | 1220648-2024-18999 | 20356680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356680&pc=OZD | 01/20/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 01/20/2023 | 10/03/2024 |  |  | 09/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281428 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | During delivery of the device, best practices used to obtain right femoral artery access. this includes injection of the bilateral iliac and common femoral artery anatomy. it was confirmed that the vessel size and pathway were clear for impella insertion. upon diagnostic catheter insertion to cross the left ventricle, the catheter would not advance due extreme ¿s¿ shaped tortuosity above the bifurcation and leading into the aortic arch. multiple catheters and wires were used to attempt access. ultimately, the physician decided to abort impella insertion. no further information is available. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. no product was returned. the cause of the delivery issue was patient condition related and related to the vessel tortuosity. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18593 | 1220648-2024-19000 | 20356685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356685&pc=OZD | 01/20/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 01/20/2023 | 10/03/2024 |  |  | 09/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275491 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that before impella cp implantation, the locking mechanism on the sterile sleeve was fully loosened but was unable to be pulled proximal on the catheter. the scrub technician tried multiple times, until the outer sleeve of catheter looked compromised from shear force. it was reported that impella support was aborted. no reported harm to the patient, no further reformation is available. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the sterile sleeve tuohy not loosening was not determined since product was not returned and there is insufficient clinical evidence. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18594 | 1220648-2024-19001 | 20356687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356687&pc=OZD | 01/25/2023 | 04/16/2025 |  | 10/03/2024 |  |  | 01/25/2023 | 10/03/2024 |  |  | 10/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276707 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  |  | Asian | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella was removed because there was a risk of pump stoppage due to a purge pressure rise alarm. the patient's condition was stable after the impella was removed. the device was explanted successfully, no further information is available. | [5 empty] | The impella device investigation evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. a visual inspection of the pump did not show any biomaterial in the purge gap or any kinks. the pump was run in a bucket of water and the high purge pressure was not reproduced. a pump tear down was performed and biomaterial was found in the motor. the cause of the high purge pressure was biomaterial. section b1:: adverse event selected for due to investigation findings section h1: reportability type corrected to serious injury due to investigation findings section d9: device available for evaluation changed to yes, device returned to manufacturer changed to yes, date device returned to manufacturer added.
18595 | 1220648-2024-19002 | 20356690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356690&pc=OZD | 01/21/2023 | 03/10/2025 |  | 10/03/2024 |  |  | 01/21/2023 | 10/03/2024 |  |  | 09/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023284564 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was implanted in a very dilated left ventricle. the physician placed the pump without issue; however, he accidentally moved the impella cp back into the patient's aorta. multiple attempts were made to place the impella cp back across the aortic valve without success. the decision was made to replace the device with another impella cp. it was reported that the patient survived the procedure. | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. impella cp with smartassist for use during cardiogenic shock. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, after reviewing the clinical information, it was determined that the cause of the positioning issues was wireless re-access. the impella device is not indicated for wireless re-access per ifu 0048-9007. this report is being filed as part of a retrospective review of historical records.
18596 | 1220648-2024-19003 | 20356691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356691&pc=OZD | 01/20/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 01/20/2023 | 10/03/2024 |  |  | 08/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023266164 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced bleeding from the impella cp and the nose. it was noted that the patient was on patient was on anticoagulants and antiplatelets. heparin was withheld, however bleeding persisted, and the physician decided to explant the device. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was most likely high patient activated clotting time (act) values since the bleeding suddenly occurred during support from around the sheath and from the patient nose at the same time. patient was on anticoagulants and antiplatelets. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18597 | 1220648-2024-19004 | 20358192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358192&pc=OZD | 08/15/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 08/15/2022 | 10/03/2024 |  |  | 05/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217418 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to being implanted with an impella cp device for mechanical circulatory support. during insertion the pump was pushed in too deep into the ventricle, which caused a kink in the cannula. the impella was explanted and replaced with a new impella cp to address the issue. the patient survived the procedure. | [5 empty] | The investigation for the reported delivery issue has been completed. the product was returned, and the delivery issue was confirmed. the returned complaint device was visually inspected. deforming was observed in the cannula. the device did not return to its original shape even after retraction. in this case the root cause is use issue. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
18598 | 1220648-2024-19005 | 20358213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358213&pc=OZD | 08/17/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 08/18/2022 | 10/03/2024 |  |  | 04/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216061 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, flows were low due to the pump kinking after a complicated insertion. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the product was not returned, per the results of the clinical evaluation: the root cause of the low pump flow was most likely a kinked cannula due to the interaction with the patient's tortuous vasculature during delivery. this report is being filed as part of a retrospective review of historical records.
18599 | 1220648-2024-19006 | 20358233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358233&pc=OZD | 08/19/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 08/19/2022 | 10/03/2024 |  |  | 03/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191833 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Female |  |  | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, saturated flows were observed as there was an unfixable kink in the catheter shaft. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of saturated flow. the device was returned for evaluation. the returned pump was visually inspected and foreign material was observed under the impeller. the root cause of the saturated flow was biomaterial. this report is being filed as part of a retrospective review of historical records.
18600 | 1220648-2024-19007 | 20358277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358277&pc=OZD | 08/19/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 08/19/2022 | 10/03/2024 |  |  | 05/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023222663 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 27 YR |  | Female |  |  | White | Injury | Yes | Yes | Crack | Fluid/Blood Leak | [5 empty] | Embolism/Embolus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed an embolism, and the purged sidearm broke. this led to the presence of artifacts and bubbles in the aortic root observed on echocardiography. the sidearm bypass did not resolve the issue. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of purge leak and embolism. the product was returned, and the purge leak was confirmed. the root cause of the purge leak was sidearm filter damage. the root cause of the embolism was not determined since the air in purge was unable to be reproduced. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18601 | 1220648-2024-19010 | 20358314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358314&pc=OZD | 08/22/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 08/22/2022 | 10/03/2024 |  |  | 10/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151728 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, low flows and suction alarms occurred. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was not received for evaluation. the root cause of the low pump flow was unable to be determined as no product or logs were returned for analysis. this report is being filed as part of a retrospective review of historical records.
18602 | 1220648-2024-19012 | 20358385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358385&pc=OZD | 08/23/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 08/23/2022 | 10/03/2024 |  |  | 05/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228981 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. following the successful placement of the impella cp, the catheter was displaced during the ramp-up phase. suction alarms were triggered and could not be resolved without reducing the p level to p-2. the physician attempted to reposition the device but was unsuccessful. a kink in the cannula was observed, leading to the decision to explant the device and replace it with a new impella cp. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was returned for evaluation. the product was returned, and a kink was confirmed in the cannula. the clinical account confirms that the pump was out of position at the time of the suction alarms and kink. the doctor was unable to reposition the device. the root cause of the low flows was due to improper positioning. this report is being filed as part of a retrospective review of historical records.
18603 | 1220648-2024-19013 | 20358391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358391&pc=OZD | 08/22/2024 | 10/25/2024 |  | 10/03/2024 |  |  | 09/03/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1369445 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant had placed a impella 5.5 pump for support of a patient admitted in for surgery and suffering from postcardiotomy cardiogenic shock. the team noted the impella 5.5 had issues with the automated impella controller, which displayed a controller failure (red alarm) after it was exchanged for a second time. the patient remained in stable condition on venoarterial extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the console (aic) controller failure (red alarm) has been completed. data logs were reviewed which showed that the console was running 12 days when the pump stopped due to phase current differences. the pump was moved over to a different console but again there were phase current differentials causing pump stops. the console was returned for investigation but the failure mode was not reproduced. the console was able to successfully run pumps without a pump stop. because there was an electrical open found in the pump, and the failure modes occurred on both consoles, neither console malfunctioned causing the pump stops. the electrical issue with the pump is captured under mfr report #1220648-2024-19014.
18604 | 1220648-2024-19016 | 20358402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358402&pc=OZD | 09/06/2024 | 10/29/2024 |  | 10/03/2024 |  |  | 09/06/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507292 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support due to st elevated myocardial infarction (stemi). while on support, the patient experienced an episode of ventricular tachycardias shortly after arrival to the cardiovascular intensive care unit. the patient was successfully cardioverted and the impella was repositioned, under echo guidance. it was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. added- h6 impact code 4641.
18605 | 1220648-2024-19015 | 20358408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358408&pc=OZD | 09/05/2024 | 11/08/2024 |  | 10/03/2024 |  |  | 09/09/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505842 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella 5.5 for the support of a patient admitted in acute myocardial infarction/cardiogenic shock. while on the impella 5.5 pump, the patient experienced an ischemic stroke. the patient had known left ventricle thrombus. there was no evidence of pump issues or clot ingestion. the treatment, explant date and patient outcome are unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of stroke and thrombus could not be determined as the device was not returned nor sufficient clinical details.
18606 | 1220648-2024-19017 | 20358414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358414&pc=OZD | 08/23/2022 | 03/10/2025 |  | 10/03/2024 |  |  | 08/23/2022 | 10/03/2024 |  |  | 03/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203926 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. upon explantation, a significant amount of old thrombus was found in the outflow and within the cannula. fresh clot was also present from pulling through the sheath. the patient was escalated to an impella 5.5. no further information is available. | [5 empty] | The impella device investigation evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombus issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18607 | 1220648-2024-19008 | 20358422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358422&pc=OZD | 08/20/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 08/20/2022 | 10/03/2024 |  |  | 03/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212426 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, low flows and suction alarms occurred. repositioning of the catheter did not resolve the pump flow. as a result, the decision was made to explant the device and replace it with a new impella cp device. upon explantation, biomaterial was observed in the inlet cage, cannula, and the sheath. the procedural outcome was that the patient survived. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
18608 | 1220648-2024-19009 | 20358424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358424&pc=OZD | 08/20/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 08/20/2022 | 10/03/2024 |  |  | 02/29/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203529 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella stopped on the 12th day of support. normal troubleshooting attempts to restart the device were unsuccessful. the impella cp was explanted and replaced to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the device was not received for evaluation. the root cause of the pump stop was unable to be determined as no product was returned for analysis. it was noted that the pump stop occurred on day 12 of support which is beyond the 8-day design life impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18609 | 1220648-2024-19011 | 20358440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358440&pc=OZD | 08/22/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 09/07/2022 | 10/03/2024 |  |  | 03/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207578 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. it was reported that a thrombus was observed on the impella device upon removal of the device. action taken to resolve the issue are unknown. the patient survived the procedure. | [5 empty] | The investigation for the reported thrombus has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the thrombus was due to biomaterial was observed on the pump at the time of pump removal. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18610 | 1220648-2024-19018 | 20358448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358448&pc=OZD | 09/08/2024 | 10/23/2024 |  | 10/03/2024 |  |  | 09/08/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025515974 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, purge pressures issues were reported. tissue plasminogen activator started, and purge blocked was resolved. there were multiple suction alarms. an echocardiogram showed pigtail slightly diverted posterior and concerns regarding being under pulmonary artery pressure or at apex. the pump was repositioned as it was noted to be deep, and device settled at p2 flow with no further suction alarms. the patient was declared brain dead and then expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the high purge pressure has been completed. the pump was not returned for inspection. the data logs confirm purge pressure high and purge system blocked alarms. there were no motor current spikes and there was a gradual rise in purge pressure before alarm was triggered. tissue plasminogen activator (tpa) was added to the purge which resolved the issue after about 4 hours. the root cause of high purge pressure was most likely biomaterial. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi states the following: section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
18611 | 1220648-2024-19019 | 20358458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358458&pc=OZD | 09/09/2024 | 11/16/2024 |  | 10/03/2024 |  |  | 09/09/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025492039 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella 5.5 support. while on support, the patient developed hemolysis, but hemolysis was resolved with repositioning. additionally, septic shock was noted. the patient eventually expired. use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined.
18612 | 1220648-2024-19020 | 20358460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358460&pc=OZD | 09/09/2024 | 11/12/2024 |  | 10/03/2024 |  |  | 09/09/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492930 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support due to st elevated myocardial infarction. while on support, the patient developed a suction alarm. volume was given and echocardiogram was done at bedside with attempt to reposition. however, an attempt to reposition was unsuccessful. the decision was made to replace with a new pump and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the repositioning issue most likely was use related due to improper technique.
18613 | 1220648-2024-19021 | 20358472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358472&pc=OZD | 09/09/2024 | 11/04/2024 |  | 10/03/2024 |  |  | 09/09/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493937 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that on the same day of impella cp placement, the cp was getting suction alarms and the impella was unable to be repositioned at bedside. a decision was made to place a new cp through the left femoral artery. the patient was pulseless electrical activity arresting at the time of insertion. it was unable to achieve return of spontaneous circulation and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. data logs were reviewed which showed pump ran without issue during support. there were 2 suction alarms triggered but resolved quickly. product was not returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. added-h6 impact code 4628. h6 impact code 4627 changed to 4629.
18614 | 1220648-2024-19014 | 20358492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358492&pc=OZD | 08/22/2024 | 11/01/2024 |  | 10/03/2024 |  |  | 09/03/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1308463 |  | 0042-0010-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Abiomed received notification of of the impact study reporting the complainant had placed a impella 5.5 pump for support of a patient admitted in for surgery and suffering from post-cardiotomy cardiogenic shock. the team noted the impella 5.5 had issues with the automated impella controller, which displayed a controller failure (red alarm). the console was exchanged and patient remained in stable condition on venoarterial extracorporeal membrane oxygenation. | [5 empty] | The investigation for the console (aic) controller failure (red alarm) has been completed. data logs were reviewed which showed the console was in use 12 days when the pump stopped due to phase current differences. this happened multiple times over the next 10 minutes as the pump was attempted to be restarted and the pmd was power cycled but the phase current differential persisted and the pump continued to stop. the pump was moved over to a different console but again there were phase current differentials causing pump stops. the console was returned for investigation but the failure mode was not reproduced and the console was able to successfully run pumps without a pump stop. because there was an electrical open found in the pump, neither console malfunctioned causing the pump stops when consoles were used with known good pumps. the pump associated with the pump stop has been reported in mfr report #1220648-2024-19013. | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
18615 | 1220648-2024-19022 | 20358516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358516&pc=OZD | 09/09/2024 | 11/14/2024 |  | 10/03/2024 |  |  | 09/09/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534751 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Intraoperative Cardiac Valve Injury | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support following cardiac arrest to attempt to wean the patient off veno-arterial extracorporeal membrane oxygenation. the impella 5.5 was placed without difficulty. however, the left ventricle was perforated. the impella was pulled back to ascending and repair was initiated. trans-esophogeal echocariogram subsequently showed a torn aortic valve leaflet and aortic (valve) insufficiency so the decision was made to replace the aortic valve. the patient was unable to sustain blood pressure and the patient ultimately expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation has been completed. no product was returned. the log review showed positioning alarms throughout the case (impella position in aorta, impella position unknown). the ventricle perforation and heart valve damage occurred during initial placement of the pump. the failure mode "vascular damage - great vessel" was updated to "heart valve damage" because the aortic valve leaflet was torn. the root cause of the heart valve damage was not determined since insufficient clinical details were provided and the product was not returned. the failure mode "ventricle perforation" was added because there was a perforation of the left ventricle, which was repaired. the root cause of the ventricle perforation was not determined since insufficient clinical details were provided and the product was not returned. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ h.6 added a health effect - clinical code due to investigation result findings since the initial manufacturer device report number 1220648-2024-19022 was submitted. h.10 add instructions for use due to investigation result findings since the initial manufacturer device report number 1220648-2024-19022 was submitted.
18616 | 1220648-2024-19023 | 20358557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358557&pc=OZD | 09/04/2024 | 11/21/2024 |  | 10/03/2024 |  |  | 09/04/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530086 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella cp implant and the doctors were prepared and intended to implant the impella cp in a patient who required resuscitation, but the patient passed away before the impella could be implanted. therefore, only the introducer kit was used. there was no allegation towards the impella for cause of death. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. a.1, a.2, a.3, a.4 and a.5 are unknown. d.6a and d.6b implant and explant dates are unknown. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, the patient passed away before the impella could be implanted as per the clinical description provided.
18617 | 1220648-2024-19024 | 20358569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358569&pc=OZD | 04/29/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/31/2022 | 10/03/2024 |  |  | 01/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190455 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implantation, the impella was having delivery issues due to the device getting stuck in the short and long introducers. a gore sheath was used with a new impella device with no harm done to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump and the sheath were visually inspected. several pigtail, cannula and catheter kinks were observed from the pump. no kink or tip broken observed from the sheath. the pump could be inserted inside the introducer. the root cause of the delivery issue was determined to be patient condition related. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18618 | 1220648-2024-19025 | 20358602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358602&pc=OZD | 09/07/2024 | 10/29/2024 |  | 10/03/2024 |  |  | 09/07/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433384 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported loss of pulses to the lower extremity after the patient was taken to the intensive care unit from impella cp placement. the patient was unable to be escalated to impella 5.5 due to severe vascular disease. the patient eventually expired. medical safety review of the event noted that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
18619 | 1220648-2024-19026 | 20358629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358629&pc=OZD | 09/25/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 09/25/2021 | 10/03/2024 |  |  | 01/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022039026 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was needing an impella cp device for mechanical circulatory support. while trying to prime the impella cp, a 'defective impella catheter' alarm was triggered. the aic was replaced and the same alarm was triggered on the replacement aic. it was reported that there were no adverse events to the patient. the pump was replaced prior to insertion and the new impella cp was used without issue. | [5 empty] | The investigation for the reported priming issue has been completed. the product was not returned and the priming issue could not be confirmed. the root cause of the priming issue was unable to be determined as no product or logs were provided for evaluation. this report is being filed as part of a retrospective review of historical records.
18620 | 1220648-2024-19028 | 20358638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358638&pc=OZD | 09/05/2024 | 11/04/2024 |  | 10/03/2024 |  |  | 09/05/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025464481 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Hemolysis | Renal Failure | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support for 43-year-old male patient, there were impella cp position in ventricle alarms. urine was noted in foley to be blood tinged. transesophageal echocardiogram (tee) performed showed impella position pointed towards mitral valve with possible entanglement of pigtail in papillary muscle. the doctor attempted to optimize position but did not want to free pigtail as they were afraid the device would pop into aorta and would not be able to get it back across. additionally, there were suction events above p-1. six units of blood products were given to the patient. the patient also had organ damage due to the hemolysis and the patient was put on dialysis. furthermore, the patient was said to have ventricular tachycardia during this time. support continued following the events. the patient was on impella support for 48.13 hours before the patient expired. the relationship between the impella and cause of death is unknown. medical safety review of the event noted there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the hemolysis, renal failure, positioning issues, low pump flow, and ventricular arrhythmias has been completed. data logs were reviewed and only the first 3 hours of data was returned for review. data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle. the alarm then was resolved and pump ran without issue to the rest of the case. per clinical description, ventricular tachycardia and renal failure were determined to be unrelated to impella before placing impella. product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. the patient has multiple organ failure, therefore, the root cause of low flow was most likely patient condition related.
18621 | 1220648-2024-19027 | 20358650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358650&pc=OZD | 05/01/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/01/2022 | 10/03/2024 |  |  | 01/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191126 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after over 8 days of impella support, the pump stopped and troubleshooting measures of controller exchange failed. the pump did not restart. the physician explanted the pump and chose not to replace it. no patient harm was reported. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18622 | 1220648-2024-19029 | 20358654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358654&pc=OZD | 09/27/2021 | 05/22/2025 |  | 10/03/2024 |  |  | 02/01/2022 | 10/03/2024 |  |  | 05/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089415 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  | Asian | Injury | Yes | No | Positioning Failure | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, hemolysis was evident by red urine. no plasma-free hemoglobin testing was conducted to confirm the hemolysis. the condition improved by adjusting the device's performance level and repositioning. additionally, it was reported that the patient developed thrombocytopenia. no additional information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the hemolysis and positioning issues was not determined since the product was not returned and insufficient clinical details were not provided. this report is being filed as part of a retrospective review of historical records.
18623 | 1220648-2024-19030 | 20358658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358658&pc=OZD | 09/06/2024 | 11/07/2024 |  | 10/03/2024 |  |  | 09/06/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025515807 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, hemolysis was noted and thrombus was confirmed. the impella cp was removed due to thrombus. it was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis and thrombus has been completed. product and data logs were not returned for review. the root cause of hemolysis and thrombus was unable to be determined as limited clinical details were provided and no product was returned for review.
18624 | 1220648-2024-18721 | 20358661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358661&pc=OZD | 08/20/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 08/20/2021 | 10/03/2024 |  |  | 05/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022082338 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella pump was found to have a broken luer lock on its sidearm. there was no reported patient involvement. | [5 empty] | The device was returned, and the investigation has been completed. the sidearm was returned for review and found to have a broken middle joint. the cause of the sidearm damage was excessive force to the middle joint. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18625 | 1220648-2024-19031 | 20358666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358666&pc=OZD | 09/29/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 09/29/2021 | 10/03/2024 |  |  | 02/28/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022067832 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella passed the aortic valve diagonally and faced the posterior wall, confirming aortic regurgitation. as the patient¿s hemodynamics were stable, the patient was successfully weaned and the impella was removed. | [5 empty] | The investigation for the reported heart valve damage issue has been completed. the impella device was not received from the customer. analysis of the provided data logs showed the pump operating as expected. the clinical information provided was not sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. as noted in the impella ifu: impella cp with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ this report is being filed as part of a retrospective review of historical records.
18626 | 1220648-2024-19032 | 20358671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358671&pc=OZD | 09/30/2021 | 05/28/2025 |  | 10/03/2024 |  |  | 09/30/2021 | 10/03/2024 |  |  | 07/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022116518 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during implantation, the impella catheter would not pass the iliac into the aorta. the medical staff considered that the cannula was kinked affecting the push ability of the impella. the decision was made to abort the process for the initial device and a new impella was utilized. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | Section h6: medical device problem code added. the site reported that impella cp would not pass from the iliac into the aorta within the aaa (abdominal aortic aneurysm) stent graft. visual inspection of the returned pump revealed catheter kinks at approximately the 24cm, 31cm, 32cm, 43cm, and 45cm marks. therefore, the cause of the reported delivery issues was determined to be multiple catheter kinks. furthermore, the catheter likely kinked due to a pump interaction with the existing aaa stent graft. this report is being filed as part of a retrospective review of historical records.
18627 | 1220648-2024-19033 | 20358680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358680&pc=OZD | 09/30/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 09/30/2021 | 10/03/2024 |  |  | 07/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110957 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. during implantation, the device could not pass the aortic valve and the impella process was aborted. the medical staff made the decision to proceed with veno-arterial extracorporeal membrane oxygenation (va ecmo) without impella support. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was determined to be patient condition as the physician stated that the valve would barely open as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18628 | 1220648-2024-19034 | 20358691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358691&pc=OZD | 10/02/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 10/02/2021 | 10/03/2024 |  |  | 07/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121198 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, as the patient was being transferred, a pump stop alarm was received. the medical staff attempted to restart the device but were unsuccessful. it was reported that the initial impella pump was explanted, and an intra-aortic balloon pump (iabp) was placed for support. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs showed that the pump stopped immediately upon the sudden occurrence of the "impella stopped - rpm deviations" alarm (alarm 115) with the ps and mc simultaneously railing to 3000 and 30000, respectively, indicating that the cables in the catheter had been damaged at this time. the pump was returned with the catheter completely severed. the optical sensor lumen had been dislodged from the catheter. a significant catheter kink was observed near the red handle. within the red handle, there was damage to the motor cable lines, with necking and discontinuity observed between the pcb and the catheter. additionally, the optical fiber ferrule and lumen were broken. electrical resistance testing confirmed that there were open lines on all 3 motor phase cables between the pcb and catheter. additionally, the ol was recreated by stretching the motor cables. the connections between patient cable and the red handle were intact. in conclusion, it was determined that the root cause of the pump stop was an open line due to excessive tensile force on the catheter. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18629 | 1220648-2024-19035 | 20358694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358694&pc=OZD | 09/09/2024 | 11/05/2024 |  | 10/03/2024 |  |  | 09/09/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445642 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support experienced bleeding from the impella access site. pressure was held and sandbag was placed. blood products were given. additionally, the patient has hemolysis. the impella was repositioned, but hemolysis did not resolve with repositioning. patient's prognosis was poor and patient decided to withdraw care. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the access site bleed and the hemolysis has been completed. pump was not returned for investigation. data logs were not returned as well. clinical mentions that the patient had high activated clotting time (acts) of 296. aggratstat was discontinued to control bleeding along with pressure and sandbag. therefore, the root cause of the access site bleeding issue was improper anticoagulant management which wad high patient act values. without the product and data logs, the hemolysis issue cannot be investigated. therefore, the root cause of the hemolysis issue was not determined since product and data logs were not returned.
18630 | 1220648-2024-19036 | 20358701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358701&pc=OZD | 05/01/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/01/2022 | 10/03/2024 |  |  | 01/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189850 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump stopped after a power spike. the impella cp was replaced with an impella 5.0 device. there was no lasting harm or injury. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs were returned and confirmed a sudden high motor current pump. the product was returned, and biomaterial was found in the outflow and in the purge gap. the high motor current pump stop was reproduced in the lab. the root cause of the pump stop was determined to be due to biomaterial. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18631 | 1220648-2024-19037 | 20358710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358710&pc=OZD | 09/05/2024 | 11/07/2024 |  | 10/03/2024 |  |  | 09/05/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520567 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Migration or Expulsion of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support due to st-elevated myocardial infarction (stemi). while on support, the patient arrested several times. cardio-pulmonary resuscitation (cpr) was performed, and the impella pulled back to the aorta. the interventional cardiologist couldn¿t recross the aortic valve. impella was removed and new device was implanted successfully. patient arrested again and patient expired. use of the device cannot conclusively be associated with the outcome. | The patient expired during the procedure. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related. b5 missing the mso review statement on manufacturer device report 1220648-2024-19037.
18632 | 1220648-2024-19038 | 20358732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358732&pc=OZD | 10/03/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 10/03/2021 | 10/03/2024 |  |  | 08/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022119335 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella pump did not start after implant. attempts to restart the device were unsuccessful. the impella device was explanted and rinsed with saline solution. no thrombotic material was identified on the pump. a new impella device was implanted to manage this issue. | [5 empty] | The investigation for the reported pump did not start issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that data logs showed that the pump stopped immediately upon activation with the occurrence of the "impella stopped - motor current high" alarm and the corresponding mc spike to 1500. multiple unsuccessful restart attempts were made, and the alarm recurred with each attempt. no other abnormal metrics were noted. visual inspection of the returned pump confirmed the presence of the full red lumen in the cannula. the outflow end of the red lumen was twisted around the impeller. no other abnormalities were found. in conclusion, it was determined that the root cause of the pump being unable to start was a fractured red lumen. this report is being filed as part of a retrospective review of historical records.
18633 | 1220648-2024-19039 | 20358738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358738&pc=OZD | 09/10/2024 | 11/16/2024 |  | 10/03/2024 |  |  | 09/10/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504274 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was explanted due to the patient experiencing limb ischemia. the patient eventually expired after care was withdrawn. use of the device cannot conclusively be associated with the outcome. | The patient eventually expired after care was withdrawn. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report: 1220648-2024-19039. b5: mso statement missing. e4 should have been left blank.
18634 | 1220648-2024-19040 | 20358744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358744&pc=OZD | 09/11/2024 | 10/31/2024 |  | 10/03/2024 |  |  | 09/11/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025502071 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support experienced groin oozing at the impella access site. repositioning repo sheath in different angles was done with little success. surgical was added after last dressing. the left groin dressing was fixed, angled of insertion matched by medical doctor at bedside, no further oozing or bleeding noted at site. the family decided to withdraw care due to a massive un-recoverable stroke. all product was removed and discarded. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, the repo sheath was re-angled and redressed to resolve oozing as per the clinical description provided.
18635 | 1220648-2024-19041 | 20358748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358748&pc=OZD | 12/05/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 12/05/2021 | 10/03/2024 |  |  | 07/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | IMPELLA CP |  |  | US |  | 2022121414 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support the pump experienced constant suction alarms and had lower than expected flows. the pump was repositioned with no improvement to suction alarms or flow. fluoroscopy was performed and revealed the pump was kinked. as a result, the impella was explanted and the patient was placed on extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18636 | 1220648-2024-19044 | 20358768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358768&pc=OZD | 12/04/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 12/04/2021 | 10/03/2024 |  |  | 09/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140414 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient n acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient had a case start delay due to alarms on the impella console for purge fluid not detected. the impella was explanted and successfully replaced with a new impella cp as a result of the purge fluid alarm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18637 | 1220648-2024-19042 | 20358782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358782&pc=PYX | 09/12/2024 | 11/06/2024 |  | 10/03/2024 |  |  | 09/12/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025487369 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Sepsis | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp flex support due to mitral valve disease. while on support, flows started to drop, the pulmonary artery pressure started to increase and motor current was slightly higher. the patient did experience pneumonia during hospitalization and may be septic. due to probable sepsis and poor prognosis, the family and team decided to pursue comfort care measures. patient expired. use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the low pump flow and infection has been completed. the pump was returned for investigation. data log was not provided for this case. the pump passed the laser continuity test and resistance test on the dp sensor. all the 5s optical parameters were within specification. some biomaterial was found on the inlet cage and was flushed out during testing in a bucket of water. purge pressure was within specification, and pump flow was found to be saturated without any observed biomaterial on the impeller. biomaterial injection could lead to saturated pump flow. according to the clinical notes, low pump flow of 1.5 l/min was reported on p7 with slightly elevated motor current. pa placement signal was reported 85/57 (70). the pump was in good position during low pump flow event. also, patient did experience pneumonia during hospitalization and may be septic. the cause of the low pump flow issue was most likely biomaterial, based on given clinical details and returned product investigation. the root cause of the sepsis could not be determined. added: d8 device returned to manufacture changed to yes corrections to the initial mfr report: b5: the statement "use of the device cannot conclusively be associated with the outcome." should have said "the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome."
18638 | 1220648-2024-19046 | 20358800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358800&pc=OZD | 09/05/2024 | 11/05/2024 |  | 10/03/2024 |  |  | 09/05/2024 | 10/03/2024 |  |  | 10/31/2025 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424110 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 63-year-old female patient was on impella cp support due to st elevated myocardial infarction. the surgeon attempted to remove the 0.18 wire but felt a lot of resistance. it was noted that the wire and impella were both kinked. the impella and wire were both removed. the patient ultimately expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella pump and guidewire was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the product damage/kinked cannula has been completed. the device was returned for evaluation. visual inspection confirmed a kink on the cannula and catheter. the root cause of cannula kink was most likely use related. h.6 code 4627 was inadvertently omitted from the initial manufacturer device report number 1220648-2024-19046 and was added.
18639 | 1220648-2024-19045 | 20358803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358803&pc=OZD | 09/30/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 10/02/2021 | 10/03/2024 |  |  | 07/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022103827 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. priming was initialized in unit to implant the impella pump. however, the outflow of the purge system could not be confirmed. the purge cassette was replaced, and priming was again attempted without success. therefore, the impella pump placement process was aborted. the procedural outcome is unknown. | [5 empty] | The investigation for the reported priming issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a purge cassette was run through de-air and the returned pump was connected to the purge cassette and purge fluid was confirmed to exit out the purge gap. the purge flow and purge pressure were within specification. the root cause of the reported issue was determined to be use related. this report is being filed as part of a retrospective review of historical records.
18640 | 1220648-2024-19043 | 20358807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358807&pc=OZD | 09/13/2024 | 10/01/2025 |  | 10/03/2024 |  |  | 09/13/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528969 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | Eu complainant reported the patient was on impella cp support and the staff noticed a discrepancy in the automated impella controller (aic) aorta-placement signal and blood pressure. also, there were several unknown things on the left ventricle placement signal and motor current. staff suspected that this might be biomaterial. the patient eventually expired after care was withdrawn. use of the device cannot conclusively be associated with the outcome. patient's peri-procedural condition was critical. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned for analysis. the data logs from the case show continuous suction alarms and several position related alarms. the placement signal was trending downwards consistently over the duration of the case. the 5s parameters were within specification. the root cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b2 updated to death. h1 updated to death. h6 updated to include death and placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-19043. e4 should have been left blank in the initial report. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
18641 | 1220648-2024-19047 | 20358837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358837&pc=OZD | 05/02/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/03/2022 | 10/03/2024 |  |  | 09/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142113 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. while inserting the cp, the physician was unable to deliver into the left ventricle. the team replaced the pump and support proceeded with a second impella cp without harm or injury to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system the investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the issue was patient condition related. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18642 | 1220648-2024-19048 | 20358863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358863&pc=OZD | 09/14/2024 | 10/14/2024 |  | 10/03/2024 |  |  | 09/14/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1401317 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support for 63-year-old male patient, the wires of the automated impella controller (aic) power cord had become exposed. the aic was therefore successfully switched. there was no reported patient harm. | [5 empty] | The investigation of automated impella controller (aic) - damage during therapy has been completed. while investigating the logs as a precaution, there were numerous connections and disconnections of ac power observed from the reported complaint date. the reported exposed wires on the power cord could not be confirmed. there is no evidence of any damage to the power cord and the aic charges as normal. the root cause of this issue was likely use related. e4 should have been left blank on manufacturer device report 1220648-2024-19048. | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
18643 | 1220648-2024-19050 | 20358894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358894&pc=OZD | 09/17/2024 | 11/13/2024 |  | 10/03/2024 |  |  | 09/17/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539107 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that impella 5.5 was attempted to be placed but the physician was unable to pass the 5.5 beyond the subclavian into the innominate. the physician attempted multiple strategies unsuccessfully, ultimately decided to abort the impella 5.5 insertion. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-19050.
18644 | 1220648-2024-19049 | 20358913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358913&pc=OZD | 05/21/2022 | 10/21/2024 |  | 10/03/2024 |  |  | 06/01/2022 | 10/03/2024 |  |  | 02/29/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188972 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while advancing the impella cp wire the patient experienced ventricular fibrillation two times, which was promptly terminated with defibrillation each time. there was no residual problems noted and the patient remained stable afterwards. | Upon further evaluation it was determined that the patient experienced access site bleeding. the bleeding could not be controlled with manual pressure and the impella was removed. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.5 revised narrative with adverse event that was inadvertently not included on the initial manufacturer device report number 1220648-2024-19049. g.2 study selection was inadvertently omitted from the initial manufacturer device report number 1220648-2024-19049. h.6 added health effect - clinical code and health effect - impact code to reflect adverse event that was not included on the initial manufacturer device report number 1220648-2024-19049.
18645 | 1220648-2024-19051 | 20358975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358975&pc=OZD | 09/03/2024 | 10/24/2024 |  | 10/03/2024 |  |  | 09/04/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523022 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient come in with a non-st elevated myocardial infarction. two days later, the patient was taken to the cardiac catheterization lab and arrested with cardiopulmonary resuscitation for 40 minutes. the impella cp was then placed for the percutaneous coronary intervention. the first impella .018 wire kinked and was replaced. the impella cp was placed and supported for two days until the family arrived for care withdrawal from anoxic brain injury. the patient expired. medical safety review of the event noted that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system arnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to the guidewire kink. added- h6 impact code 4629.
18646 | 1220648-2024-19052 | 20358997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20358997&pc=OZD | 05/03/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/03/2022 | 10/03/2024 |  |  | 01/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183516 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 12 MO |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. while inserting the impella, the catheter had a kink in the cannula after the peel away was removed. the kink occurred when catheter was advanced too far into the left ventricle (lv) and the pig tail was caught on the shaft of the cannula. device was explanted and the case was aborted. no patient harm was reported. | [5 empty] | The investigation for the reported delivery and cannula kink issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the cannula kink was determined to be improper positioning of the pump when peeling away the introducer. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18647 | 1220648-2024-19053 | 20359013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20359013&pc=OZD | 09/06/2024 | 11/11/2024 |  | 10/03/2024 |  |  | 09/06/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531229 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and they had access site bleeding with a hematoma. manual pressure was held, femstop was placed and anticoagulation was withheld. it was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
18648 | 1220648-2024-19054 | 20359099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20359099&pc=OZD | 09/03/2024 | 10/21/2024 |  | 10/03/2024 |  |  | 09/03/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025483328 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 73-year-old male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during impella 5.5 support, the patient developed a hematoma at the right axillary insertion site that was noted in the intentive care unit the day after placement. hematoma is stable. heparin was held. the patient received one unit of packed red blood cells due to a gastrointestinal bleed after it was discovered that the blood loss was separate from the hematoma. care was ultimately withdrawn. medical safety review of the event noted that the use of the device cannot conclusively be associated with the outcome. patient's peri-procedural condition was critical. | [5 empty] | The investigation of vascular damage - peripheral vessel and access site bleeding has been completed. the impella device was not returned for evaluation. the cause of access site bleeding was low patient activated clotting time (act) values. due to the lack of clinical details provided regarding the vascular damage - peripheral vessels, the root cause cannot be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
18649 | 1220648-2024-19055 | 20359110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20359110&pc=OZD | 05/05/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/05/2022 | 10/03/2024 |  |  | 11/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168236 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. during implantation, the impella was unable to be delivered and kinked. the 14fr sheath had kinked and when the new sheath was placed the anatomy still precluded placement. impella implant procedure was aborted. no patient harm was reported. | [5 empty] | The investigation for the reported delivery and detached inflow/outflow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the pump was visually inspected. cather and cannula kink were observed. the introduction was not returned for more analysis. the root cause of the delivery issue was determined to be patient condition because of vessels diseases. the root cause of introducer issue was kink in the sheath because of vessel diseases. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18650 | 1220648-2024-19057 | 20359267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20359267&pc=OZD | 05/06/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/06/2022 | 10/03/2024 |  |  | 01/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189369 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, blood was oozing from either the impella site and or the device's repositioning sheath. the impella cp device was weaned and explanted. patient survived the procedure. no additional information was provided. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the product was inspected. upon examination, there was no visual damage on the sheath. the valve leak test was performed on the sheath and no leak was found. the root cause of the access site bleeding was determined to be use related. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18651 | 1220648-2024-19056 | 20359306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20359306&pc=OZD | 09/03/2024 | 10/25/2024 |  | 10/03/2024 |  |  | 09/03/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532074 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after an impella cp was placed. stat echocardiogram in the catheterization lab revealed the impella was aiming toward the mitral structures. multiple attempts to reposition were unsuccessful. it was further reported that the 63-year-old male patient developed oozing from the access site. the 14 french peel away sheath was left in place due to the patient's large body habitus and deep tissue tract. hematoma was noted after the procedure. manual pressure held for 15 minutes and the site was soft. blood products were administered. however, care was ultimately withdrawn. medical safety review of the event noted the use of the device cannot conclusively be associated with the outcome. patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and positioning issues has been completed. based on data logs and clinical description, the root cause of positioning issue was most likely patient condition. the root cause of access site bleeding was most likely due to patient condition.
18652 | 1220648-2024-19058 | 20359553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20359553&pc=OZD | 05/06/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/06/2022 | 10/03/2024 |  |  | 12/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173769 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the device had low flows and lost waveforms. the impella was pulled back slightly, but the pigtail remained across the aorta valve. the impella device was repositioned. troubleshooting was performed but unable to resolve waveforms and flows. the impella was removed and a second pump was successfully placed without harm or injury from interruption in the support. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the field rt log was reviewed and 7 minutes into the run, suction and a loss of mc pulsatility were observed. placement signal waveform was steady. the returned pump was visually inspected, and no abnormalities were observed. the pump was run in a basin of water and the pump flow was 3.7 l/min. the root cause of the issue was most likely improper positioning of the device since the issue occurred during the repositioning after the repo sheath. this report is being filed as part of a retrospective review of historical records.
18653 | 1220648-2024-19059 | 20359684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20359684&pc=OZD | 05/06/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/06/2022 | 10/03/2024 |  |  | 01/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191015 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a kink in the cannula in the aortic arch postdelivery. the pump had to be removed with the sheath and 14fr cook sheath inserted in its place. a second impella cp was successfully through the cook sheath. no injury or patient harm was reported. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the pump was returned and inspected. there was a visible cannula kink on the pump. the patient's anatomy most likely contributed to the cannula kink during delivery. the root cause of the issue was a kinked cannula. this report is being filed as part of a retrospective review of historical records.
18654 | 1220648-2024-19062 | 20359892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20359892&pc=OZD | 10/26/2022 | 04/03/2025 |  | 10/03/2024 |  |  | 10/31/2022 | 10/03/2024 |  |  | 02/29/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188969 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was reported that after the impella cp implantation and a switch to repositioning the sheath caused bleeding in the groin site in the intensive care unit. the issue was resolved by manual compression and superficial surgical stitch. the patient remained stable and survived the surgery. | It was further reported while advancing the impella wire, ventricular fibrillation (vf) was induced (x2), which was promptly terminated with defibrillation (x2). no residual problems. patient remained stable afterwards. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation of access site bleeding and vf has been completed. the product was not returned. the clinical details suggest that a suture was sewn additionally to resolve/control bleeding. the root cause of the access site bleeding-major issue was most likely use related since additional suture resolved bleeding issue per clinical information. as the necessary information was not provided, the root cause of the vt/vf was not determined. b5 additional information was was inadvertently omitted on the initial manufacturer device report number 1220648-2024-19062. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-19062 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-19062 in accordance with updated procedures h6 health effect clinical code and health effect impact code revised to reflect the omitted adverse events on the initial manufacturer device report number 1220648-2024-19062. h6 code 3221 was reported incorrectly as it was duplicated on the initial manufacturer device report number 1220648-2024-19062.
18655 | 1220648-2024-19060 | 20359929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20359929&pc=OZD | 05/06/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/06/2022 | 10/03/2024 |  |  | 02/29/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194905 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. access was obtained and the 14fr peel away sheath was inserted into the right femoral artery (rfa) without complications. the impella cp was turned on and flow did not get above 0.9 liters. the impella device was removed and the new impella cp was successfully implanted. no patient harm was reported | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs showed "suction" alarms occurred shortly after the start of support, accompanied by low motor current pulsatility and low flows. visual inspection of the returned pump revealed the presence of biomaterial on and around the impeller. no other abnormalities were observed. therefore, the root cause of the low pump flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
18656 | 1220648-2024-19061 | 20359963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20359963&pc=OZD | 09/10/2024 | 11/21/2024 |  | 10/03/2024 |  |  | 09/10/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519210 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was explanted due to high hemolysis values. the patient eventually expired after care was withdrawn. use of the device cannot conclusively be associated with the outcome. | The patient eventually expired after care was withdrawn. per medical safety officer review the use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b5 mso statement missing from manufacturer device report.
18657 | 1220648-2024-19063 | 20359998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20359998&pc=OZD | 08/24/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 08/24/2022 | 10/03/2024 |  |  | 05/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227373 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, low flows and suction alarms occurred. the team attempted to troubleshoot the issue by administering additional volume to the patient. the patient had a small left ventricle with trabeculations noted. the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was returned for evaluation. the root cause of the low flows was due to patient condition related. low volume / small left ventricle. returned device operated to specification. this report is being filed as part of a retrospective review of historical records.
18658 | 1220648-2024-19065 | 20360033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360033&pc=OZD | 12/07/2021 | 02/13/2025 |  | 10/03/2024 |  |  | 12/07/2021 | 10/03/2024 |  |  | 08/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022130418 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the pump abruptly stopped operating. the impella cp pump was removed and successfully replaced. there were no allegations of patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The impella device was returned, and an evaluation is completed. sufficient data logs were not returned. visual inspection of the returned pump revealed the presence of the red lumen in the pump. a portion of the red lumen was twisted around the impeller and fractured off at the end. no other damage or abnormalities were found. therefore, it was determined that the cause of the pump unable to start was a fractured red lumen.
18659 | 1220648-2024-19066 | 20360064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360064&pc=OZD | 08/25/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 08/25/2022 | 10/03/2024 |  |  | 03/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214332 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the impella stopped on the 8th day of support. normal troubleshooting attempts to restart the device were unsuccessful. the impella cp was explanted and replaced to address the issue. the patient survived the procedure. | [5 empty] | The investigation for the reported pump stop has been completed. the pump was not returned; however, the data stick was returned for evaluation. the logs were reviewed and showed that an impella pump stop alarm occurred after 8 days which is beyond the device's indications for use. as noted in the impella ifu: impella cp with smart assist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
18660 | 1220648-2024-19070 | 20360074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360074&pc=OZD | 12/18/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 12/18/2021 | 10/03/2024 |  |  | 10/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022147640 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. it was reported the impella was out of position and was not locked down prior to leaving the cardiac catheterization laboratory (ccl). reportedly, the physician was unable to achieve optimal flows despite several repositioning attempts under echocardiogram. the physician met with intermittent resistance due to the patient¿s vasculature being extremely calcified, the flows remained low at pressure level two despite troubleshooting. the physician made the decision to replace the pump and achieved adequate flows with the new pump. there was no patient harm reported. | [5 empty] | The investigation into the low pump flow has been completed. the impella pump was returned for investigation. the visual inspection of the returned pump indicated a kink on the catheter by the kink protector. the investigator noted this kink would not affect the flow output of the impella. no biomaterial was seen in the cannula, outflow, or around the impeller. the pump was run in a basin of water and performed within specifications. the root cause of the low flow condition was most likely ingested biomaterial that dislodged during explant. the pump was able to perform within specification showing that no product malfunction was present. the cause of the low pump flow was most likely ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
18661 | 1220648-2024-19067 | 20360080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360080&pc=OZD | 09/07/2024 | 10/01/2025 |  | 10/03/2024 |  |  | 09/11/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025465407 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | Ventricular Fibrillation | Gastrointestinal Hemorrhage | [10 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The us complainant had a cp support ongoing for a patient admitted in for surgery and post-cardiotomy cardiogenic shock was diagnosed. the patient was also supported by extra corporeal membrane oxygenation (ecmo). the pump lost placement signal but, the signal loss did not cause any harm or injury and the pump was not replaced. the pump remained on for support. days later the team noted there were harms during the support. there was ventricular fibrillation that prompted defibrillation and the patient experienced a gastrointestinal bleed. the patient eventually expired after care was withdrawn. use of the device cannot conclusively be associated with the outcome. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ventricular fibrillation and gastrointestinal bleed has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the ventricular fibrillation and gastrointestinal bleed and its relation to the impella automated controller device was not determined since the product, data logs, and sufficient clinical information were not provided. the additional information is reflected in b2, b3, g6, h2. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. data logs were reviewed, and the optical signal issues were confirmed. based on the clinical information provided and data logs, it was determined that the root cause of the optical signal issues was most likely to be a patient condition. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b2 updated to death. h1 updated to death. h6 updated to include death and placement signal coding corrections that were made from the initial manufacturer device report number 1220648-2024-19067. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
18662 | 1220648-2024-19064 | 20360090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360090&pc=OZD | 09/10/2024 | 11/05/2024 |  | 10/03/2024 |  |  | 09/10/2024 | 10/03/2024 |  |  | 10/31/2025 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024420756 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during impella 5.5 support the 67-year-old male patient had fasciotomy and thrombectomy. the impella was explanted after two days of support, and patient was transitioned to comfort care. the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details.
18663 | 1220648-2024-19071 | 20360117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360117&pc=OZD | 08/25/2022 | 03/27/2025 |  | 10/03/2024 |  |  | 08/25/2022 | 10/03/2024 |  |  | 04/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218515 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the side arm of the device was damaged out of the packaging. the impella was explanted and replaced with a new impella cp to address the issue. the patient survived the procedure. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported product damage (bond failure) issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a visual inspection of the returned pump revealed the purge sidearm was broken into two pieces at the bond between the air filter and the pressure reservoir. no other damage was found. therefore, the cause of the product damage was sidearm bond damage due to excessive force. this report is being filed as part of a retrospective review of historical records.
18664 | 1220648-2024-19068 | 20360141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360141&pc=OZD | 12/07/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 12/07/2021 | 10/03/2024 |  |  | 08/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127360 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported the pump position screen was displayed when priming was not complete. the reporter indicated the reason for restarting the pump was purge liquid did not flow out from the pump even in the final stage of priming. there was no patient harm reported, and this device was replaced. | [5 empty] | The investigation into the positioning issue has been completed. the impella pump was returned for investigation. the field log was reviewed, and purge flow and pressure were within specification prior to the purge disc being removed. the returned purge cassette was run through de-air, connected to the returned pump and purge fluid was confirmed to exit the purge gap. the cause of the failure mode was use related due to improper technique since the issue was unable to be reproduced and the purge was within specification during field use. this report is being filed as part of a retrospective review of historical records.
18665 | 1220648-2024-19069 | 20360151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360151&pc=OZD | 05/09/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/09/2022 | 10/03/2024 |  |  | 12/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161737 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was an observed purge leak. the patient remained on support and no harm was reported. | [5 empty] | The investigation for the reported purge leak issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump was visually inspected, and no abnormalities were observed. the red handle was then opened, and a syringe with colored fluid was connected to the sidearm and a leak was reproduced from the sidearm kink protector. the kink protector was then torn down, and the location of the leak was confirmed to be at the adhesive joint step from the large tubing to the small tubing. the root cause of the purge leak was tubing bond failure at the sidearm kink protector. impella cp with smartassist for use during cardiogenic shock and high risk cpi instructions for use for the related event are as follows: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18666 | 1220648-2024-19075 | 20360160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360160&pc=OZD | 09/27/2022 | 03/06/2025 |  | 10/03/2024 |  |  | 09/27/2022 | 10/03/2024 |  |  | 06/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249499 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the complainant reported that multiple efforts to prime the device were unsuccessful. the consoles were switched, and purge cassettes changed without change. the decision was made to remove the pump and replace it with a new impella device. the procedural outcome was the patient survived. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The real time event logs did not indicate any priming issues and no high purge pressure or purge system failure alarms were found in the logs. a visual inspection of the pump revealed dried dextrose on the impeller around the purge gap which indicates that purge fluid had been exiting the purge gap in the field. the pump was attached to a purge cassette and a console and an attempt to prime the pump was made. the first attempt was unsuccessful and high purge pressure alarms and purge system failure alarms occurred. the pump was detached from the purge cassette and a few minutes later purge fluid that was remaining in the purge line began to exit the purge gap. the pump was re-attached to the purge cassette and the console and priming of the pump was successful. due to the dried dextrose around the purge gap and successful priming of the pump during the investigation, the cause of the priming issue was most likely use related. this report is being filed as part of a retrospective review of historical records.
18667 | 1220648-2024-19073 | 20360172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360172&pc=OZD | 12/19/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 02/08/2022 | 10/03/2024 |  |  | 07/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108716 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patien in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis and experienced access site bleeding. no further information regarding treatment was noted. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the hemolysis and access site bleeding has been completed. the investigator was unable to confirm the failures as the product was not returned and no clinical details provided were sufficient to determine a root cause. the complaint failures were not confirmed, additionally there was no evidence of a relationship between the alleged hemolysis and impella device. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: access site bleeding. impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis: impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system. section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction : as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
18668 | 1220648-2024-19072 | 20360203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360203&pc=OZD | 08/26/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 08/26/2022 | 10/03/2024 |  |  | 12/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178973 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella access site. the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18669 | 1220648-2024-19074 | 20360222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360222&pc=OZD | 09/13/2024 | 04/21/2025 |  | 10/03/2024 |  |  | 09/13/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1764240 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Eu complainant reported the patient was on impella support and during support there was a controller failure on the automated impella controller (aic). the aic was replaced. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. after review of the data logs the reported controller failure alarm was not confirmed. the console was internally inspected, and all cables were fully connected. the kbd-related controller error could not be reproduced after power cycling 200 times. there were also no issues (specifically no persistent rfid errors) with detection of a known-good purge cassette during final testing. conclusion: the root cause for the kbd-related controller error could not be determined since it was not reproducible. the root cause for the intermittent purge cassette detection (rfid errors) could not be determined since it was not reproducible. b3 date of event was erroneously entered on manufacturer device report 1220648-2024-19074. g1 reporting contact email revised on manufacturer device report 1220648-2024-19074 in accordance with updated procedures.
18670 | 1220648-2024-19076 | 20360278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360278&pc=OZD | 10/15/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/16/2022 | 10/03/2024 |  |  | 07/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253263 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, low flows and suction alarms occurred. a thrombus was observed on the impella device upon removal of the device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was returned for evaluation and thrombus in the cannula near the impeller. the root cause of the low pump flow was determined to be biomaterial. this report is being filed as part of a retrospective review of historical records.
18671 | 1220648-2024-19077 | 20360321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360321&pc=OZD | 08/08/2022 | 03/10/2025 |  | 10/03/2024 |  |  | 03/02/2023 | 10/03/2024 |  |  | 05/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224097 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced hemolysis. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
18672 | 1220648-2024-19079 | 20360329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360329&pc=OZD | 10/15/2022 | 03/10/2025 |  | 10/03/2024 |  |  | 10/18/2022 | 10/03/2024 |  |  | 01/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182344 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced hemolysis. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle: ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis: ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hemolysis issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18673 | 1220648-2024-19078 | 20360346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360346&pc=OZD | 09/15/2024 | 11/18/2024 |  | 10/03/2024 |  |  | 09/15/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534118 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while on support for the 65-year-old male patient, there was some concern of the impella cp bouncing in the left ventricle. staff reviewed the position with echo to make sure the impella was not in the mitral. patient did have some ectopy as well. support continued. additionally, the patient had hemolysis on support. the pump was therefore explanted and patient escalated to an impella 5.5. the patient was noted to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis and arrhythmia has been completed. product and data logs were not returned for review. the root cause of hemolysis and the ectopy was unable to be determined as limited clinical details were provided and no product was returned for review. corrections to the initial mfr report: h6 impact code 4627 changed to 4629.
18674 | 1220648-2024-19080 | 20360414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360414&pc=OZD | 10/19/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/19/2022 | 10/03/2024 |  |  | 03/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200396 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced limb ischemia resulting in the decision to explant the pump. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of limb ischemia. the device was not received for evaluation. the root cause of the limb ischemia was the use of a sheath that was not indicated for use with impella. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
18675 | 1220648-2024-19082 | 20360423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360423&pc=OZD | 10/20/2022 | 10/09/2025 |  | 10/03/2024 |  |  | 10/20/2022 | 10/03/2024 |  |  | 04/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224120 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. during the removal of the peel-away sheath, the pump migrated into the mitral apparatus, causing issues with the inflow volume from the left artery. since the device could not be repositioned without a guidewire, the decision was made to explant the initial impella device and implant a new one. the patient survived the procedure. | The complainant reported that the impella pump was inserted properly. during removal of the peel away sheath, the pump migrated into the mitral apparatus causing issues with the inflow volume from the left artery. as the device could not be moved back into position without a wire, the decision was made to explant the initial impella device and implant a new pump. the procedural outcome is the process for the original impella pump was aborted for placement of the new device. |  |  |  |  | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. no data logs were sent. after reviewing the clinical information, it was determined that the cause of the positioning issues was use related. g1: updated mfg email address.
18676 | 1220648-2024-19081 | 20360437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360437&pc=OZD | 05/10/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/10/2022 | 10/03/2024 |  |  | 02/29/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193810 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp device failed to redeliver after it came out of position. the device was pulled back, but the device came back through the aortic valve. impella cp was explanted and a new impella inserted without cause, successful case weaned explanted in procedural area. no patient harm or injury noted. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned product was visually inspected, and a cannula kink was observed. the root cause of the pump pulling out of the ventricle issue was use related since the pump was positioned under a papillary muscle causing the cannula to kink and be stuck. this report is being filed as part of a retrospective review of historical records.
18677 | 1220648-2024-19084 | 20360550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360550&pc=OZD | 07/15/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 07/15/2021 | 10/03/2024 |  |  | 05/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022085962 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, pump stopped after placement in the left ventricle. no additional information was provided. no patient harm was reported. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the reported pump stop could not be determined through this investigation as the product was discarded and no log files were provided. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
18678 | 1220648-2024-19083 | 20360615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360615&pc=OZD | 03/23/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 03/23/2022 | 10/03/2024 |  |  | 01/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2022057150 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, there was hemolysis reported, with no specified resolution. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the issue was improper positioning of the device. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: hemolysis impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
18679 | 1220648-2024-19086 | 20360644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360644&pc=OZD | 05/10/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/10/2022 | 10/03/2024 |  |  | 02/29/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195613 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the device had low flow. the device was explanted and replaced with a new pump. no patient injury or harm noted. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that returned pump visually inspected and no kink observed. inspected under an accurate microscope. no broken impeller blades observed. no biomaterial or foreign materials found in inflow, outflow or cannula. no data logs were returned or uploaded to find positioning issue when kit implemented. the pump purged and ran for around 3 hours. no low flow pump observed. pump runs normal. the root cause of the low pump flow could not be determined. this report is being filed as part of a retrospective review of historical records.
18680 | 1220648-2024-19087 | 20360661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360661&pc=OZD | 07/31/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 07/31/2021 | 10/03/2024 |  |  | 08/31/2022 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1502456 |  | 0048-0034-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that after inserting the impella, the placement signal suddenly disappeared. exchanging the automated implantable cardioverter defibrillator (aic) did not improve the matter. the possibility of a blockage of the placement signal lumen was suspected, so the position of the tube was confirmed as per the manual. the pump worked without issue other than ps display problem. it is unknown how the reported issued was resolved. no patient injury or harm was noted. | [5 empty] | The returned data logs showed that the placement signal railed to 3000 mmhg with the sudden occurrence of the "placement signal not reliable" alarm. notably, this alarm was immediately preceded by the "epm_sensor_defekt" alarm (alarm 701). no damage or abnormalities were found on the returned product. when the pump was connected to the aic, placement signal values displayed on the console and were within specification, and no alarms related to the placement signal occurred. additionally, straining and moving the patient cable and pump did not recreate the failure mode. in conclusion, the cause of the placement signal issue was not determined since the failure mode could not be reproduced. this report is being filed as part of a retrospective review of historical records.
18681 | 1220648-2024-19088 | 20360666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360666&pc=OZD | 11/13/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 11/13/2022 | 10/03/2024 |  |  | 08/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023266174 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device experienced high purge pressure and was explanted and replaced as a result. there was no reported patient harm. | [5 empty] | The investigation for the reported high purge pressure has been completed. the impella device has not been returned for evaluation. in addition, as the necessary clinical information was not provided and the device was not returned, the root cause of the high purge pressure could not be determined. this report is being filed as part of a retrospective review of historical records.
18682 | 1220648-2024-19085 | 20360738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360738&pc=OZD | 11/12/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 11/12/2022 | 10/03/2024 |  |  | 06/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238718 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the device began showing suction alarms. the p level was lowered, and the pump produced a "pump in aorta" alarm. the p level was then increased, but the flows and motor current were much lower than expected. the patient remained on impella support, and there was no reported patient harm. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. clinical details suggest improper positioning was the root cause of the low flow. this report is being filed as part of a retrospective review of historical records.
18683 | 1220648-2024-19089 | 20360840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360840&pc=OZD | 08/23/2021 | 03/07/2025 |  | 10/03/2024 |  |  | 08/23/2021 | 10/03/2024 |  |  | 12/31/2022 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022027788 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 23 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that during support, the pump suddenly stopped. multiple unsuccessful attempts were made to restart the pump. the pump was explanted and replaced with another impella cp, and no patient harm was reported. | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. the field long term log was reviewed, and a pump stop was observed after approximately 29 minutes of support associated with an impella stopped ¿ motor current high alarm. this alarm was immediately followed by an impella stopped alarm and impella stopped ¿ reverse flow alarms. the returned pump was visually inspected and damage to the sterile sleeve at approximately the 77cm mark was observed. the pump catheter was sliced open in this location, exposing the underlying wiring. electrical resistance testing revealed multiple motor phases below specification. therefore, the root cause of the sudden pump stop was multiple shorted lines which most likely occurred due to improper use resulting in damage to the pump catheter. d8 / d9: date of return of device added this report is being filed as part of a retrospective review of historical records.
18684 | 1220648-2024-19090 | 20360934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20360934&pc=OZD | 08/24/2021 | 02/13/2025 |  | 10/03/2024 |  |  | 08/24/2021 | 10/03/2024 |  |  | 06/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | ABIOMED, INC. |  |  | ABIOMED, INC. |  |  | US |  | 2022094929 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, it was reported that the pump stopped after ingesting a clot. the physician decided to explant the device and replace it with another impella cp. it was reported that the patient survived the procedure. | [5 empty] | The impella device investigation evaluation is underway. upon completion of the investigation, a supplemental report will be filed. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella; rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no issues pertaining to this failure mode were found during device history record review. the data logs were returned and confirmed that the pump was unable to start and alarm 119 was triggered. the product was returned and no biomaterial or abnormalities were found. the pump was run in a basin without purge and ran at a higher than expected motor current and flow rate, causing a high motor current pump stop. the pump run again with purge for 18 hours and operated within specification. the red handle was opened and no abnormalities were found. electrical testing was done and all resistances were within specification. the pump was run in a water both at body temp and the failure mode was not reproduced. the catheter was torqued and manipulated while the pump was running and the failure was not reproduced. the alarm 119 pump stop was unable to be reproduced in the lab. the root cause of the pump stop was most likely an intermittent open line on one of the motor phases as the alarm 119 accompanied by the motor current spike to approximately 500ma indicates. this report is being filed as part of a retrospective review of historical records.
18685 | 1220648-2024-19091 | 20361029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20361029&pc=OZD | 11/14/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 11/14/2022 | 10/03/2024 |  |  | 06/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236609 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing off pump coronary artery bypass was implanted with an impella cp device for mechanical circulatory support. upon initiating support, the motor stopped due to a high motor current alarm. the pump was restarted at p8, but the motor current was higher than normal. the pump was then explanted and replaced with another impella. there was no reported harm to the patient. | [5 empty] | The investigation for the reported temporary pump stop has been completed. the impella device has not been returned for evaluation. data logs showed the "impella stopped - motor current high" alarm occurred within a minute of activating the pump after a rapid increase in motor current to nearly 1500. the pump was successfully restarted. however, the motor current remained elevated at around 1000, and the pump flow waveform was saturated regardless of p-level. the left ventricle waveform was nearly double the value of the aortic placement signal; there was no evidence of an optical sensor malfunction. however, the cause of the pump stop was not determined since product was not returned. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella; rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18686 | 1220648-2024-19092 | 20361213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20361213&pc=OZD | 11/15/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 11/15/2022 | 10/03/2024 |  |  | 11/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168965 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, there were abnormal readings via the device's optical signal, though no alarms occurred. the patient attempted troubleshoot, but the issue persisted. eventually, the signal dropped back into the normal range on its own, and the procedure was completed on impella support. however, the device was removed during unloading due to this issue. there was no reported patient harm. | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the optical signal issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18687 | 1220648-2024-19093 | 20361232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20361232&pc=OZD | 08/30/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 09/01/2021 | 10/03/2024 |  |  | 04/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022077306 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient of unknown race/ethnicity noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. during placement, the pump was found to have a noticeable kink. the team attempted to insert and use the pump, but the flow rates were less than desirable. the pump was explanted and exchanged for a new pump. no patient harm was reported. | [5 empty] | The investigation for the reported product damage (cannula kink) has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the product damage (cannula kink) was due to improper technique used resulting in kinked cannula. this report is being filed as part of a retrospective review of historical records.
18688 | 1220648-2024-19094 | 20361556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20361556&pc=OZD | 09/01/2021 | 03/05/2025 |  | 10/03/2024 |  |  | 09/01/2021 | 10/03/2024 |  |  | 04/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022071624 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited low flow. under fluoroscopy a kink was noted, and repositioning was attempted. it was reported that this did not resolve the issue. the pump was explanted and replaced with a new impella cp. it was reported that the patient survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smart assist system section: warnings & cautions: warnings ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
18689 | 1220648-2024-19096 | 20361594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20361594&pc=OZD | 09/01/2021 | 09/30/2025 |  | 10/03/2024 |  |  | 09/01/2021 | 10/03/2024 |  |  | 05/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022088958 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | User facility | Company representation |  |  |  | 84 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the cp placed had lower than optimal support for pump flows. during support the device exhibited low flow. it was reported that flows never achieved prescribed performance level despite repositioning under fluoroscopy. due to hemodynamic instability the physician decided to explant the pump and replace with another impella cp. | [5 empty] | The investigation for the reported low flow has been completed. the product was returned, and per the results of the clinical review and investigation: the root cause of the cannula kink was determined to be wireless re-access which is not permitted with the use of impella cp. data logs from the case were not sent for investigation. no notable damage or abnormalities were found on the returned pump. when the pump was run, the low pump flow was not reproduced as the flow rate was within specification. in conclusion, the cause of the low pump flow was not determined since the failure mode could not be reproduced. this report is being filed as part of a retrospective review of historical records. | D7a was inadvertently omitted on the initial manufacturer device report 1220648-2024-19096. h6 event history log review was coded incorrectly in the initial manufacture report 1220648-2024-19096 as there was no return of logs. h10 the root cause of the cannula kink was determined to be wireless re-access which is not permitted with the use of impella cp was reported incorrectly in the initial manufacture report 1220648-2024-19096. there was no cannula kink issue for this complaint.
18690 | 1220648-2024-19095 | 20361595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20361595&pc=OZD | 11/15/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 11/15/2022 | 10/03/2024 |  |  | 02/29/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200757 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump had high purge pressure and low purge flow during support. the team made several troubleshooting attempts and replaced the purge cassette, but the problem continued. they then elected to remove and replace this pump with a second impella device. there was no reported harm to the patient. | [5 empty] | The investigation for the reported high purge pressure has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the high purge pressure could not be determined. this report is being filed as part of a retrospective review of historical records.
18691 | 1220648-2024-19114 | 20361641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20361641&pc=OZD | 12/30/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 12/30/2021 | 10/03/2024 |  |  | 09/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022136591 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump was unable to obtain flows greater than one liter per minute and the motor current remained flat. the decision was made to remove the pump. the device was explanted. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the low pump flow has been completed. the review of the data logs confirmed suction alarms and low flows. however, the root cause could not be determined without the returned device. no clinical details provided were sufficient to determine a root cause. this report is being filed as part of a retrospective review of historical records.
18692 | 1220648-2024-19098 | 20361712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20361712&pc=OZD | 09/13/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 09/13/2021 | 10/03/2024 |  |  | 06/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022094960 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the complainant reported that hemolysis occurred after insertion of impella. p level lowered; however, it was noted that the indwelling sheath was not locked, and position was adjusted. hemolysis improved and impella was weaned and explanted. patient is stable post explant. | [5 empty] | The returned logs showed no alarms and the returned product showed no damage which would have contributed to the hemolysis. the physician reported that repositioning the pump resolved the hemolysis. therefore, the root cause of the hemolysis was incorrect positioning. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
18693 | 1220648-2024-19117 | 20361836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20361836&pc=OZD | 01/03/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 01/03/2022 | 10/03/2024 |  |  | 10/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022153063 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the pump would not flow above 0.7 liters per minute. the pump was repositioned multiple times and volume given with no resolution. this pump was replaced, and a new pump placed. there was no patient harm reported. | [5 empty] | The investigation into the low pump flow has been completed. the impella pump was returned for investigation. the returned pump was visually inspected and a kink in the cannula was observed. the root cause investigation determined the low pump flow resulted from a kinked cannula. this report is being filed as part of a retrospective review of historical records.
18694 | 1220648-2024-19097 | 20361846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20361846&pc=OZD | 11/16/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 11/16/2022 | 10/03/2024 |  |  | 07/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254825 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the device produced a suction alarm and had low flows. repositioning attempts failed to resolve the issue, and it was later discovered that the cannula would kink while in the patient due to pressure being applied from a calcified aorta and shallow ventricle. however, the team chose to keep the device in the patient. the patient remained on impella support, and no harm was reported. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. it was determined that the cause of the low pump flow was most likely a kinked cannula due to the patient's calcified aorta and shallow ventricle. this report is being filed as part of a retrospective review of historical records.
18695 | 1220648-2024-19102 | 20362141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20362141&pc=OZD | 12/21/2022 | 03/27/2025 |  | 10/03/2024 |  |  | 12/21/2022 | 10/03/2024 |  |  | 09/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023277441 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported the impella pump was implanted, and placement was optimal. immediate suction alarms were received upon initiation of the device, and the pump flow would only reach 1.4 liters per minute. the impella was explanted and replaced with a new pump. no further action is available. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the device passed all post-sterile inspection checks. the evaluation of the device was visually inspected, and a cannula kink was observed. no broken blade was found. biomaterial was observed around the impella blades. since the pump started there were immediate suction alarms and pump would not flow more than 1.4l/min, the root cause most likely is biomaterial. data logs showed suction alarms and motor current spikes which also indicate biomaterial ingestion. the cannula kink may have happened during attempting to solve the issue. this report is being filed as part of a retrospective review of historical records.
18696 | 1220648-2024-19103 | 20362149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20362149&pc=OZD | 05/11/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/11/2022 | 10/03/2024 |  |  | 02/29/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193110 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during insertion, the peel away sheath was removed but the patient had profound bleeding around repo site. stitch was placed with minimal improvement. the physician decided to rewire the port and place a new pump and leave the peel away in place. no harm or injury was noted. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. after reviewing the clinical information, it was determined that the root cause of the access site bleeding was most likely a lack of vessel recoil onto the repositioning sheath. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18697 | 1220648-2024-19101 | 20362165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20362165&pc=OZD | 09/05/2024 | 08/25/2025 |  | 10/03/2024 |  |  | 09/05/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520879 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, that the impella cp pump stopped. the pump was attempted to be restarted without resolution. an alarm on the automated impella controller (aic) displayed impella stopped, controller failure. nurse was instructed via phone to retrieve another controller and was instructed on how to perform a controller change out. care was withdrawn and patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | B1 product problem was inadvertently omitted on the initial manufacturer device report number. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h6 investigation conclusions revised on manufacturer device report in accordance with updated procedures.
18698 | 1220648-2024-19099 | 20362169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20362169&pc=OZD | 12/21/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 12/21/2022 | 10/03/2024 |  |  | 06/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231940 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant had a successful cp support of 5 days when the pump was weaned and explanted. the pump was observed at the explant that there was a kink to the pump. no reported harm to patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the product damage was most likely a kinked catheter. as noted in the impella ifu: impella cp with smart assist system section: warnings, cautions, and precautions ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records.
18699 | 1220648-2024-19105 | 20362263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20362263&pc=OZD | 07/10/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 07/10/2023 | 10/03/2024 |  |  | 01/31/2025 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316552 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to being implanted with an impella cp device for mechanical circulatory support. a thrombus was discovered during the sheath placement. therefore, the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation into the delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. based on clinical description, the root cause of delivery issue was most likely due to patient condition. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18700 | 1220648-2024-19108 | 20362342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20362342&pc=OZD | 07/08/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 07/08/2023 | 10/03/2024 |  |  | 10/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024297577 |  | 048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, flows were low. multiple attempts to re-position catheter were not successful. the impella cp was explanted and replaced with a new impella cp device. clots were seen in the inlet and outlet windows. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records.
18701 | 1220648-2024-19106 | 20362351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20362351&pc=OZD | 01/04/2023 | 10/20/2025 |  | 10/03/2024 |  |  | 01/04/2023 | 10/03/2024 |  |  | 09/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278107 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon startup, the pump exhibited low flow, flat motor current, and suction issues, which did not resolve with repositioning. thrombus ingestion was suspected. the pump was successfully explanted and replaced, and a clot was visible in the cannula upon explant. the patient survived the procedure. | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18702 | 1220648-2024-19115 | 20362384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20362384&pc=OZD | 01/01/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 01/01/2022 | 10/03/2024 |  |  | 04/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022076524 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the impella experienced suction issues and reduced pressure. the clinician assisted the team with troubleshooting the issue and noted a kink in the proximal portion of the cannula distal to the outlet. reportedly, the patient had an extremely small left ventricle, short aorta root and type three arch. the physician pulled the device back to attempt to straighten the kink, without success. the device was removed, and a new device successfully inserted with a v-18 guide wire for more support. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the product damage cannula kink has been completed. the root cause of the low flow was determined to be a cannula kink that occurred during insertion due to the patient¿s condition of a tortuous vasculature. this report is being filed as part of a retrospective review of historical records.
18703 | 1220648-2024-19107 | 20362476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20362476&pc=OZD | 09/05/2024 | 08/22/2025 |  | 10/03/2024 |  |  | 09/05/2024 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1221443 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, that the impella cp pump stopped. the pump was attempted to be restarted without resolution. an alarm on the automated impella controller (aic) displayed impella stopped, controller failure. nurse was instructed via phone to retrieve another controller and was instructed on how to perform a controller change out. care was withdrawn and patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant reported the patient was on impella cp support. while on support, that the impella cp pump stopped. the pump was attempted to be restarted without resolution. an alarm on the automated impella controller (aic) displayed ¿impella stopped, controller failure¿. nurse was instructed via phone to retrieve another controller and was instructed on how to perform a controller change out. later care was withdrawn and patient expired. the death has been documented on manufacture report number 1220648-2024-19101. |  |  |  |  | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. b1 adverse event revised as it was erroneously entered on the initial manufacturer device report number. b2 death should not be on the initial submission as death was reported on another manufacture report. b5 revised from the initial submission as death was reported on another manufacture report. d2 common device name revised on manufacturer device report in accordance with updated procedures. d9 device returned to manufacturer date added g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h1 revised from the initial submission as death was reported on another manufacture report. h3 revised from the initial submission as device was returned h6 health effect impact code revised from the initial submission as death was reported on another manufacture report and the devised was replaced. h6 investigation conclusions revised from the initial submission in accordance with updated procedures. h10 additional manufacture narrative revised as product was returned.
18704 | 1220648-2024-19113 | 20362730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20362730&pc=OZD | 12/21/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 12/21/2021 | 10/03/2024 |  |  | 09/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022147132 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. during attempted implant, physician had trouble placing the pump and felt the pump was not flowing well. the physician made the decision to remove the pump as the patient was stable. there was no patient harm reported. | [5 empty] | The investigation into the low pump flow has been completed. the impella pump was returned for investigation. the pump was returned, and visual inspection revealed dried blood around the outflow and impeller, but a thrombus was not observed. the pump was run in a basin of water and operated within specification. the root cause of the low flow was most likely improper technique by the end user as when positioning the device, the physician mentioned that they were unable to get the pump into the correct position. this report is being filed as part of a retrospective review of historical records.
18705 | 1220648-2024-19109 | 20362772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20362772&pc=OZD | 12/21/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 12/21/2021 | 10/03/2024 |  |  | 09/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140109 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the pump experienced low pump flow issues and failed to flow over 0.8 liters per minute. the motor current was reported to be flat and the impella position was confirmed to be good. the complaint requested replacement due to failure and a new pump was inserted without issue. there was no patient harm reported. | [5 empty] | The investigation into the low pump flow issue has been completed. the impella pump was returned for investigation. a review of the provided data logs confirmed that the pump flow was approximately 0.8 lpm. evaluation of the returned pump found biomaterial around the impeller. the pump was run in the lab, and low flows were reproduced. upon removal of the biomaterial, the low flows resolved. the root cause of the low flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
18706 | 1220648-2024-19100 | 20362830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20362830&pc=OZD | 05/10/2022 | 04/07/2025 |  | 10/03/2024 |  |  | 05/20/2022 | 10/03/2024 |  |  | 01/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178489 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient endured an ischemic stroke approximately 10 days post impella explant with unknown relationship. it is unknown how the reported issue was resolved. no additional information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
18707 | 1220648-2024-19104 | 20362877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20362877&pc=OZD | 12/21/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 12/21/2022 | 10/03/2024 |  |  | 08/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267110 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant had an impella cp placed when the team discovered the pump was out of position and the team was unable to reposition. the pump was removed and replaced. no patient harm reported due to the issue. | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
18708 | 1220648-2024-19110 | 20362937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20362937&pc=OZD | 05/11/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/11/2022 | 10/03/2024 |  |  | 02/29/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196166 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with high-risk percutaneous intervention was implanted with an impella cp device for mechanical circulatory support. the impella was primed, inserted, aspirated, flushed sheath, and initiated with suction. in response to the event, p-level was decreased, and the device was repositioned, but device continued to have suction above p1. impella was turned off, explanted and replaced with a new impella cp pump. no patient harm was reported. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the pump was inspected. biomaterial was found in the cannula and adhered to the impeller. the pump was run at p1. switched the pump to p2 and suction alarms occurred after 1 minute of running. after running the pump, the biomaterial was still adhered on the impeller. the suction issue described in the case was able to be reproduced successfully. the root cause of the low pump flow was determined to be biomaterial. this report is being filed as part of a retrospective review of historical records.
18709 | 1220648-2024-19111 | 20363014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363014&pc=OZD | 12/16/2021 | 09/30/2025 |  | 10/03/2024 |  |  | 12/16/2021 | 10/03/2024 |  |  | 09/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022145990 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient had coagulopathy with bleeding from the mouth, nares (nose), endotracheal tube, foley and central line. reportedly, a pulmonologist branched the patient at bedside and removed massive clots. the patient remained on impella support and was successfully weaned. additionally, a thrombus was noted on the catheter post removal of the impella. patient survived the explant procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). impella cp with smartassist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. the returned pump was run in a basin of water and operated to specification. no thrombus was observed. the thrombus was most likely present on the inflow and pigtail at the time of removal per images provided from the field. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. corrected information provided in h6 and h10.
18710 | 1220648-2024-19112 | 20363247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363247&pc=OZD | 11/08/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 09/12/2024 | 10/03/2024 |  |  | 08/31/2025 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398005 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Delivered as Unsterile Product | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the impella's cp sheath packaging was compromised. during inspection of the product a very small tear was found on underside of sheath packaging on the left-hand side. there was no patient involvement therefore no patient information. | [5 empty] | The investigation has been completed. no product was returned. per the clinical investigation, the root cause of the packaging (sterility) issue was not determined since product was not returned for investigation.
18711 | 1220648-2024-19116 | 20363319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363319&pc=OZD | 01/01/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 01/01/2022 | 10/03/2024 |  |  | 08/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022134819 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. there were continuous alarms of placement signal unreliable. reportedly, troubleshooting was unsuccessful, and the pump placement was aborted. the pump was replaced with another pump which was successfully placed and noted to be functional. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the optical signal issue has been completed. the root cause of the error could not be determined as no product or data were returned. this report is being filed as part of a retrospective review of historical records.
18712 | 1220648-2024-19118 | 20363387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363387&pc=OZD | 01/04/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 01/04/2022 | 10/03/2024 |  |  | 10/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151899 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. during attempted implant, positioning issues were encountered due to the grey cannula kink. the physician made the decision to abort placement of this impella and replaced the device with another impella cp. there was no patient harm reported. | [5 empty] | The investigation into the product damage has been completed. the impella pump was returned for investigation. visual inspection of the returned pump confirmed a cannula kink near the motor housing. the root cause of the delivery issues was the observed cannula kink. this report is being filed as part of a retrospective review of historical records.
18713 | 1220648-2024-19122 | 20363533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363533&pc=OZD | 09/14/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 09/14/2021 | 10/03/2024 |  |  | 06/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094898 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hemolysis and thrombocytopenia. the patient's peri-procedural condition was critical and subsequently expired due to low cardiac function and circulatory collapse. the physician stated that there was no causal relationship with impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
18714 | 1220648-2024-19119 | 20363544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363544&pc=OZD | 05/12/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/12/2022 | 10/03/2024 |  |  | 01/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191054 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with high-risk surgery was attempted to be implanted with an impella cp device for mechanical circulatory support. during implant, there were delivery issues due to limitations. the impella cp device was not navigating the right axillary artery torturous into ascending aorta. it was noted the pigtail was bending and cannula was not going down descending aorta. the impella cp was replaced with an impella 5.5 pump. the physician was able to maneuver the impella 5.5 device into the ascending aorta and across the aorta valve. there was no harm or injury from insertion or delivery attempts. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the pump visually inspected and no abnormalities were observed on the returned product. the root cause of the delivery issue was determined to be patient condition related. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18715 | 1220648-2024-19125 | 20363558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363558&pc=OZD | 01/13/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 01/13/2022 | 10/03/2024 |  |  | 10/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022148628 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was taken to the cardiac catheterization laboratory (ccl) for impella placement. it was reported due to the patient¿s strong calcification and meandering of the iliac, the impella was not able to be implanted. reportedly, the catheter shaft kinked, and it was thought with any additional force the blood vessel could rupture. the impella placement was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18716 | 1220648-2024-19134 | 20363583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363583&pc=OZD | 05/13/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/13/2022 | 10/03/2024 |  |  | 02/29/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194745 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after impella insertion, pump flows were low between 1.2-1.6. activated clotting time (act) was 305 at the time of insertion and sheath was flushed well before insertion. the physician removed the pump, flushed it (no clots were noted to come out of device) and a new impella cp device was inserted. patient survived the procedure and no harm was reported. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs showed "suction" alarms and reduced motor pulsatility about 10 minutes into support. prior to the alarm, the motor current spiked and then rapidly decreased. no damage or abnormalities were found on the returned pump. the pump ran with no issues. therefore, it was determined that the root cause of the low pump flow was patient condition related. this report is being filed as part of a retrospective review of historical records.
18717 | 1220648-2024-19121 | 20363588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363588&pc=OZD | 03/29/2022 | 04/07/2025 |  | 10/03/2024 |  |  | 03/29/2022 | 10/03/2024 |  |  | 07/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110755 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patientnoted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia, no further information was available regarding the ischemia. the patient remained on impella support and was successfully weaned. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
18718 | 1220648-2024-19137 | 20363597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363597&pc=OZD | 01/23/2023 | 03/10/2025 |  | 10/03/2024 |  |  | 01/23/2023 | 10/03/2024 |  |  | 09/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275593 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the clinical team lost pump position. the pump was accidentally pulled from the left ventricle and was unable to be crossed back across the valve. the team explanted the pump and replaced it with another impella cp. no further information is available. | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on clinical data, the pump was accidentally pulled into the aorta. multiple attempts to re-cross the aortic valve failed. the root cause was use related due to improper technique. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ this report is being filed as part of a retrospective review of historical records.
18719 | 1220648-2024-19123 | 20363613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363613&pc=OZD | 11/22/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 11/22/2021 | 10/03/2024 |  |  | 08/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022134637 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the pump was place and started in auto mode and pump flows were reported to be .5 liters per minute. the pump appeared to be too deep, the physician attempted to pull back the pump, but it was kinked and attempts to pull back more to get the kink out and readvance were unsuccessful. this pump was removed and a new impella pump placed without any issues. there was no patient harm reported, | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the product damage has been completed. the site reported that pump flow was 0.5 lpm and it appeared that the pump was positioned too deep and was kinked. the root cause of the low pump flow was most likely a kinked cannula due to poor positioning resulting from improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
18720 | 1220648-2024-19120 | 20363630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363630&pc=OZD | 01/10/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 01/10/2022 | 10/03/2024 |  |  | 10/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022153512 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump pressure level seven was lower than expected and if the pressure was increased, suction alarms were generating. repositioning attempts failed. the patient was taken to the operating room for repositioning under fluoroscopy, as the latest echocardiogram showed the impella position to be suboptimal, several attempts to reposition were unsuccessful. the decision was made to remove the initial impella device and place a new impella to maintain sterility. the device was replaced. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18721 | 1220648-2024-19132 | 20363635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363635&pc=OZD | 01/23/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 01/23/2023 | 10/03/2024 |  |  | 09/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274578 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. despite active clotting time of 260 and flushed sheath, the impella cp was advanced successfully. flows were immediately reduced after initiation of support and the cardiac team made the decision made to removed impella cp. upon explant a clot was visualized. the pump was run under heparinized saline and flows returned to normal, however purge flow was severely reduced, and purge pressures was rising. the decision was made to prepare and insert a new impella cp. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was possible. in the clinical review it was noted that biomaterial was observed. when looking at the logs it was confirmed that a suction alarm was present with the placement signal trending downwards. the returned product was soaked and there was biomaterial that remained in the outflow cage. the pump was run in the bucket and low flows were reproduced. the cause of the low pump flow was ingested biomaterial. instructions for use for the related event are as follows: embolism/thrombus impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18722 | 1220648-2024-19124 | 20363769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363769&pc=OZD | 05/13/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/13/2022 | 10/03/2024 |  |  | 12/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176798 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp pump stopped after 8 days of support. the impella cp was explanted and replaced. no patient harm or injury noted. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smart assist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
18723 | 1220648-2024-19127 | 20363784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363784&pc=OZD | 01/03/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 01/03/2022 | 10/03/2024 |  |  | 10/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151775 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported during the initial setup of the impella in the field, a purge leak was noted at the red plug on the catheter side where it appeared to be leaking between the red and grey pieces. the pump was not inserted due to the leak and procedure was aborted. there was no patient harm reported. | [5 empty] | The product was returned, and a purge leak was visually confirmed in between the kink protector and the red handle on the catheter end. the kink protector was able to shift slightly but was securely glued to the red handle. the catheter was unable to freely twist inside of the kink protector. the red handle was opened, and no twisted wires were observed. the leak was originating behind the monkey fist and under the kink protector. the monkey fist was easily removed from the kink protector and the lines were pulled to expose a severed purge line. a product history review was performed which did not indicate any systemic issues. the root cause of the purge leak was determined to be a severed purge line, due to manufacturing operator error. as noted in the impella ifu: this report is being filed as part of a retrospective review of historical records. impella cp with smart assist for use during cardiogenic shock and high risk cpi section: cleaning as noted in the impella ifu ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18724 | 1220648-2024-19129 | 20363811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363811&pc=OZD | 01/21/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 01/21/2023 | 10/03/2024 |  |  | 08/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262264 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a purge pressure error when setting up pump, also the pump failed to prime. the physician decided to explant the device and replace with another impella cp. no additional information was provided | [5 empty] | The investigation for the reported issue has been completed. the field log was reviewed and found air in purge and the purge system was open. additionally, the purge line click-on was not detected and this triggered alarms during priming. the returned device was visually inspected, and no abnormalities were observed. the returned purge cassette was primed without issue, connected to the returned pump, and the purge exited the purge gap without issue. purge metrics were within specification. the cause of the priming issue was most likely use related since the issue was unable to be reproduced on the returned product and the logs indicated that the purge disc was not connected. this report is being filed as part of a retrospective review of historical records.
18725 | 1220648-2024-19126 | 20363821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363821&pc=OZD | 12/01/2021 | 10/03/2024 |  | 10/03/2024 |  |  | 12/01/2021 | 10/03/2024 |  |  | 05/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022088684 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was suspected that the patient had endured hemolysis during impella support confirmed via hemolyzed labs. no intervention was required. the hemolyzed labs improved after successful wean/removal of the device. patient survived the procedure. | [5 empty] | The investigation for the reported hemolysis issues has been completed. the impella device has been returned for evaluation. the product was returned and hemolysis was confirmed. the root cause of the hemolysis issue was patient condition. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
18726 | 1220648-2024-19128 | 20363822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363822&pc=OZD | 01/13/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 01/13/2022 | 10/03/2024 |  |  | 10/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022145924 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patientin acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella cp pump stopped. several unsuccessful attempts were made to restart the pump on a different pressure level and changing out the automated impella controller. it was reported the pump still would not start and showed a controller error after several attempts to restart. later it was discovered that the purge solution the clinic was using was not optimal, it was unclear how long the purge solution was used. the patient was reported to be stable and was successfully weaned. | [5 empty] | The investigation into the pump stop has been completed. the impella pump was returned for investigation. the system logs show a sudden increase in motor current followed by a high motor current pump stop. bearing wear was confirmed on the returned product. trending did not indicate any systemic issues related to pump stop. the cause of the pump stop issue was bearing wear. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi & impella rp with smart assist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smart assist system catheter unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
18727 | 1220648-2024-19142 | 20363827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363827&pc=OZD | 03/22/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 03/22/2022 | 10/03/2024 |  |  | 12/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176211 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was taken to the cardiac catheterization laboratory (ccl) for impella pre-operative placement. the impella was placed and upon starting in auto mode suction issues were noted, the pressure level was dropped to three to break the suction, however the suction alarm occurred again. this was reportedly repeated multiple times and eventually the pump would only sustain low flows without suction at pressure level one. an echocardiogram confirmed the correct positioning of the impella, the physician thought the issue might be volume, however after a 500cc bolus there was no improvement. additionally, the physician initially did not believe the suction events could be clot related, but having exhausted all other measures to resolve the suction issue, the physician made the decision to remove the impella. the impella was removed and replaced with a new impella pump. there was no patient harm reported, and this device was replaced. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the low pump flow has been completed. the root cause of the suction and low flow was determined to be ingested biomaterial as a thrombus was seen upon explant. this report is being filed as part of a retrospective review of historical records.
18728 | 1220648-2024-19138 | 20363877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363877&pc=OZD | 01/24/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 01/24/2023 | 10/03/2024 |  |  | 09/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273899 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was olled over by the medical staff for cleaning purposes and the catheter was pulled back into the aorta. multiple attempts were made to re-advance the pump; however, they were ultimately unsuccessful. the pump was subsequently explanted and replaced with an impella 5.5 to continue with patient support. there was no report of patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on the clinical details, root cause of this issue was patient movement for cleaning purposes. this action led to pull back from the impella catheter to the aorta. the physician was unable to reposition again due to improper use technique. this report is being filed as part of a retrospective review of historical records.
18729 | 1220648-2024-19154 | 20363894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363894&pc=OZD | 03/25/2022 | 07/09/2025 |  | 10/03/2024 |  |  | 03/25/2022 | 10/03/2024 |  |  | 10/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144077 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Restricted Flow rate | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a priming failure during preparation of the device. when priming was performed, no outflow of purge liquid from the pump was confirmed. the pump was replaced, and support proceeded. there was no harm to the patient. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation of priming issue has been completed. a field real time log was reviewed and purge flow was confirmed to operate to specification. the returned pump and returned purge cassette were connected to a console and priming fluid was confirmed to exit the purge gap. there was no issue found during the case. the device functioned/operated to specification. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h6 type of investigation, investigation findings, and, investigation conclusions codes were erroneously entered on the initial manufacturer device report number. the investigation was completed at the time of the initial report h10 investigation results were inadvertently omitted on the initial manufacturer device report.
18730 | 1220648-2024-19133 | 20363902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363902&pc=OZD | 01/23/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 01/23/2023 | 10/03/2024 |  |  | 10/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283815 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there was a positioning issue. following initial placement, it was reported that the device dropped in the left ventricle. repositioning was attempted; however, the pigtail was caught in the papillary muscle. the pump was eventually pulled back, however it jumped into the aorta. multiple attempts were made to re-advance the device, but the pump subsequently kinked and the decision was made to switch out the pump. the pump was explanted and replaced with another impella cp pump for continued support. it was reported that the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18731 | 1220648-2024-19130 | 20363928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363928&pc=OZD | 01/14/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 01/14/2022 | 10/03/2024 |  |  | 07/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121208 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the impella cp pump stopped on day eight of support. the team was unsuccessful in attempts to restart the pump at the previous pressure level and attempts to restart the pump by unplugging and re-plugging the catheter. the consoles were also changed with no success. the physician came to pull the catheter back across the valve. heparin was reported to have been maintained in the purge since insertion in the catheter laboratory and there were no breaks in the heparinized purge. the pump was removed, heparin stopped and activated clotting time (act) decreased to allow for the removal of the pump. impella was successfully weaned, and patient survived the procedure. | [5 empty] | The investigation into the pump stop has been completed. the impella pump was returned for investigation. a review of the data logs confirmed the pump stop. an analysis of the returned product found no kinks in the catheter, open lines were found on all three phases of electrical continuity testing. additionally, open lines were observed adjacent to the monkey fist by the catheter kink protector. the root cause of the pump stop was open lines due to excessive force. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi & impella rp with smart assist system section: using the automated impella controller with the impella catheter as noted in the impella ifu ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smart assist system catheter as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
18732 | 1220648-2024-19131 | 20363957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363957&pc=OZD | 01/23/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 01/23/2023 | 10/03/2024 |  |  | 05/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218842 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. as the team was preparing for implantation and the pump passed through the 14fr sheath, the console displayed a "change aic (automated impella controller)" alarm. the aic was changed, however the alarm persisted. the physician decided to explant the impella cp and replace it with a new impella cp. no further information is available. | [5 empty] | The investigation for the reported "pump not detected" alarm has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18733 | 1220648-2024-19162 | 20363974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363974&pc=OZD | 03/25/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 03/25/2022 | 10/03/2024 |  |  | 01/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022043579 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a pump stop during use of the device. the pump was removed and replaced with no harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. data logs were not sent. visual inspection of the returned pump revealed significant amounts of biomaterial in the outflow, around the impeller, and in the purge gap. therefore, the cause of the pump stop was ingested biomaterial. as noted in the impella ifu: impella cp with smart assist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
18734 | 1220648-2024-19136 | 20363978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363978&pc=OZD | 01/17/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 01/17/2022 | 10/03/2024 |  |  | 10/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022155890 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention with atherectomy was implanted with an impella cp device for mechanical circulatory support. it was reported low flows were reported after placement of the impella. reportedly, the procedure was able to be successfully completed with the device set to pressure level nine resulting in flows of 2.5 to 3.0 liters per minute. the team reportedly silenced the suction alarm to enable the completion of the procedure. the patient remained on impella support and was successfully weaned. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the low pump flow has been completed. the root cause of the low flow pump flow was determined to be due to a cannula kink as seen on imaging. the cannula kink was determined to be caused by patient condition as this secondary pump was placed and kinked in the same location. the root cause of the kink was the patient anatomy. this report is being filed as part of a retrospective review of historical records.
18735 | 1220648-2024-19141 | 20363993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20363993&pc=OZD | 01/24/2023 | 03/10/2025 |  | 10/03/2024 |  |  | 01/24/2023 | 10/03/2024 |  |  | 09/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276564 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp exhibited low flow, and the physician tried to reposition the device without success. the physician decided to explant the device and replace it with a new impella cp. it was reported that the patient survived the procedure. | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported low pump flow issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18736 | 1220648-2024-19143 | 20364009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364009&pc=OZD | 05/16/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/16/2022 | 10/03/2024 |  |  | 10/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023159504 |  | 0048-0047 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the device had high purge pressure alarms. troubleshooting was performed by the team by replacing the personal computer (pc) but the purge pressures were not resolved. the team decided to explant the pump prematurely. no patient harm was reported. | [5 empty] | The investigation for the reported high purge pressure issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs showed the purge pressure gradually increased from the beginning of support until the "purge pressure high" alarm triggered after about 5 hours. the motor current remained pulsatile. the impella was returned with two purge cassettes. a catheter kink was found next to the kink protector of the red pump handle. the catheter was stripped, and considerable kinking was observed on the purge lumen (along with the motor and optical cables). therefore, the root cause of the high purge pressure was a kinked purge lumen. this report is being filed as part of a retrospective review of historical records.
18737 | 1220648-2024-19144 | 20364036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364036&pc=OZD | 01/27/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 01/27/2023 | 10/03/2024 |  |  | 04/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023211688 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was bleeding from the repositioning sheath. blood came out of side of the sheath and not the puncture site/access site. the impella was removed as the patient was hemodynamically stable. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer, therefore, evaluation of the device was not possible. the root cause of the bleeding is unknown due to lack of clinical details regarding best practices. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18738 | 1220648-2024-19139 | 20364042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364042&pc=OZD | 05/13/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/13/2022 | 10/03/2024 |  |  | 11/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162396 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was an inability to proceed beyond the purge screen during preparation of the pump. a new impella pump was used with no harm to the patient. | [5 empty] | The investigation for the reported priming issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned product indicated the pump had gone through case start and was able to purge and run without issue. data logs also indicated the pump was able to be primed successfully. thus, the purge cassette failing to prime is most likely a use issue. the root cause of the priming issue was determined to be use related. this report is being filed as part of a retrospective review of historical records.
18739 | 1220648-2024-19140 | 20364057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364057&pc=OZD | 01/24/2023 | 10/13/2025 |  | 10/03/2024 |  |  | 01/24/2023 | 10/03/2024 |  |  | 08/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023271058 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | White | Injury | Yes | Yes | Break | Defective Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella delivery, the clear locking ring from the long dilator remained on the catheter and the impella catheter went through the locking ring. the physician had to utilize shears to cut the plastic ring off of the impella catheter to continue advancing the catheter. the repositioning sheath went into the right femoral artery. it was reported that the right femoral artery access site had bloody drainage on arrival to the unit. the site was angle matched and redressed. the medical staff added bicarbonate to the purge and administered blood products. purge initiated. additionally, the patient experienced bleeding at the right groin. manual pressure was held, bicarbonate was added to the purge, and blood products were administered. it was reported that the patient survived the procedure. | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella delivery, the clear locking ring from the long dilator remained on the catheter and the impella catheter went through the locking ring. the physician had to utilize shears to cut the plastic ring off of the impella catheter to continue advancing the catheter. the repositioning sheath went into the right femoral artery. it was reported that the right femoral artery access site had bloody drainage on arrival to the unit. the site was angle matched and redressed. the medical staff added bicarbonate to the purge and administered blood products. purge initiated. additionally, the patient experienced bleeding at the right groin. manual pressure was held, bicarbonate was added to the purge, and blood products were administered. it was reported that the patient survived the procedure. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: access site bleeding. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation introducer damage ¿ mechanical and access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-19140 in accordance with updated procedures. h6 component code 4742 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-19140.
18740 | 1220648-2024-19153 | 20364094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364094&pc=OZD | 04/15/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 04/15/2023 | 10/03/2024 |  |  | 12/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320843 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, low flows occurred. as a result, the decision was made to explant the device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the product was returned. the root cause of the low pump flow was due to biomaterial on the outflow cage. this report is being filed as part of a retrospective review of historical records.
18741 | 1220648-2024-19167 | 20364100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364100&pc=OZD | 09/19/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/25/2022 | 10/03/2024 |  |  | 09/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022137123 |  | 0048-0024-JP |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, gender and ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding. reportedly the bleeding occurred after the sheath was exchanged and was difficult to control, resulting in removal of the impella. the device was removed due to the bleeding, no additional information is available. | [5 empty] | The impella device was not received from the customer, therefore, evaluation of the device was not possible. the investigation into the access site bleeding has been completed; however, the cause of the issue was not determined since the product was not returned. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18742 | 1220648-2024-19150 | 20364110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364110&pc=OZD | 04/12/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 04/12/2023 | 10/03/2024 |  |  | 12/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316470 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the surgeon was unable to gain access to coronaries via radial approach. the physician felt anatomy in arch was preventing impella from flowing well, and also inhibiting his ability to cannulate coronary. therefore, the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18743 | 1220648-2024-19152 | 20364130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364130&pc=OZD | 04/15/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 04/15/2023 | 10/03/2024 |  |  | 10/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024296319 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, a spike in motor current followed by the pump stop occurred. the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the product was returned, and the pump stop was confirmed. the root cause of the pump stop was bearing wear. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18744 | 1220648-2024-19155 | 20364171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364171&pc=OZD | 03/25/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 03/25/2022 | 10/03/2024 |  |  | 11/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158069 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a pump stop during use of the device. the pump was restarted without any harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the product was returned, and the purge gap was found to be enlarged. no data logs were returned for this investigation. the root cause of the pump stop was most likely bearing wear. as noted in the impella ifu: impella cp with smart assist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18745 | 1220648-2024-19156 | 20364172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364172&pc=OZD | 03/26/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 03/28/2022 | 10/03/2024 |  |  | 01/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174652 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant had a case start issue which prompted the team to replace the pump. the team had easy guide lumen issues. there was no harm or injury from the delay in support initiation. | [5 empty] | The investigation for the reported issue has been completed. the pump was returned for analysis, but the red lumen was not. the root cause of the delivery issues was determined to most likely be red lumen issues caused by improper technique. this report is being filed as part of a retrospective review of historical records.
18746 | 1220648-2024-19158 | 20364175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364175&pc=OZD | 03/26/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 04/05/2022 | 10/03/2024 |  |  | 12/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023164191 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant had a successful cp support. approximately one month later, limb ischemia was noted in the case, the day of explant. patient recovered and is stable post issue. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the limb ischemia was most likely patient condition. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
18747 | 1220648-2024-19161 | 20364193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364193&pc=OZD | 05/17/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/17/2022 | 10/03/2024 |  |  | 03/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191406 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/17/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the medical doctor (md) reported that shortly after implant there was no flow. the impella cp device was replaced. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause of the low pump flow was most likely poor patient condition related. the returned product operated to specification and the patient had low volume. this report is being filed as part of a retrospective review of historical records.
18748 | 1220648-2024-19145 | 20364237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364237&pc=OZD | 01/27/2023 | 03/10/2025 |  | 10/03/2024 |  |  | 01/27/2023 | 10/03/2024 |  |  | 09/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276522 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Tachycardia | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis, as evidenced by red urine. the patient also experienced multiple instances of pulseless ventricular fibrillation and ventricular tachycardia. in addition, the device had a positioning issue as the pigtail was still across the valve. it is unknown how these issues were addressed; however it was reported that the patient survived the procedure. | [5 empty] | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: hemolysis: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hemolysis, vf/vt/af/at and positioning issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18749 | 1220648-2024-19170 | 20364239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364239&pc=OZD | 04/10/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 04/10/2022 | 10/03/2024 |  |  | 01/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183629 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump stopped during impella support. the pump was removed. the patient remained on extracorporeal membrane oxygenation (ecmo). no patient harm was reported. | [5 empty] | The investigation for the reported pump stop issue has been completed. the product was returned, and the pump stop issue was not confirmed. the cause of the pump stop was bearing wear. the device was used beyond product claims. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
18750 | 1220648-2024-19171 | 20364244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364244&pc=OZD | 09/29/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 09/29/2022 | 10/03/2024 |  |  | 06/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240389 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was noted that the patient had previous bilateral iliac stents so impella placement was difficult. after insertion an x-ray was performed, and the physician discovered a kink in the impella cannula. the physician decided to replace the device with another impella cp. it was reported that the patient survived the surgery. | [5 empty] | The investigation for the reported product damage (cannula kink) issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause was patient condition related. this report is being filed as part of a retrospective review of historical records.
18751 | 1220648-2024-19147 | 20364257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364257&pc=OZD | 04/11/2023 | 10/29/2025 |  | 10/03/2024 |  |  | 04/14/2023 | 10/03/2024 |  |  | 12/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024301940 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, flows were low and suction alarms occurred. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was returned, and the evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the -- left ventricle (lv) waveform reproduced after connecting to an automated impella controller (aic). the field long term log review confirmed the suction and position in aorta alarms and showed no issue with the optical sensor parameters. the root cause of the missing lv waveform is due to the patient¿s suction condition that resulted in a drop in p-level and the missing numerical value for the lv waveform. this report is being filed as part of a retrospective review of historical records.
18752 | 1220648-2024-19149 | 20364264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364264&pc=OZD | 04/12/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 04/12/2023 | 10/03/2024 |  |  | 11/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316001 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient experienced a hematoma at the impella access site. the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of access site bleeding. the device was not received for evaluation. the root cause of the access site bleeding was unable to be determined since product was not returned and sufficient clinical information was not provided. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18753 | 1220648-2024-19157 | 20364299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364299&pc=OZD | 03/26/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 03/28/2022 | 10/03/2024 |  |  | 12/31/2022 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022043520 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 24 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that impella cannula was found kinking, which in turn reduced flows with a possible pump stop that occurred during use of the device. the pump was removed and replaced with no harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the product was returned and a kink in the cannula was confirmed. the root cause of the low flow was determined to be a cannula kink as confirmed through the clinical account. the cause of the cannula kink was most likely a result of improper positioning and technique while removing the peel away sheath. this report is being filed as part of a retrospective review of historical records.
18754 | 1220648-2024-19180 | 20364305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364305&pc=OZD | 10/06/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/06/2022 | 10/03/2024 |  |  | 04/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218701 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that when support was initiated, impella flows were noted to be abnormally low. the physician decided to explant the device and replace with another impella cp. it was reported that the patient survived the procedure, no additional information is available. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has not been returned for evaluation. the data logs were reviewed and showed that suction alarms triggered within seconds of starting the pump and persisted for the remainder of support. the motor current had low pulsatility and the pump flow gradually decreased throughout support despite troubleshooting attempts. however, since no product was returned, the cause of the low pump flow was not determined. this report is being filed as part of a retrospective review of historical records.
18755 | 1220648-2024-19146 | 20364310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364310&pc=OZD | 03/23/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 03/23/2022 | 10/03/2024 |  |  | 10/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022153516 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the impella developed a kink in the cannula just below the outlet cage while being placed into the ventricle. the impella was placed regardless of the product damage. reportedly the physician was aware of the kink in the cannula and attributed the kink to the patients' anatomy being extremely small and with little aortic arch. the procedure was completed with flows at 1.41 lpm and pressure level 2, the patient was reported to have tolerated the procedure well and the impella was removed upon completion of the procedure. there was no patient harm reported | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the investigation into the product damage has been completed. per clinical review the patient had very small and tortuous vessels. the root cause of the cannula kink was determined to be the patient¿s condition. this report is being filed as part of a retrospective review of historical records.
18756 | 1220648-2024-19163 | 20364330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364330&pc=OZD | 09/15/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 09/15/2022 | 10/03/2024 |  |  | 05/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227322 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported the impella was successfully placed and upon start up there was a complete decoupling of the arterial left ventricle wave forms. the flows never got above 1 liter. the physician was concerned there may be a clot and decided to remove this impella and place a new impella. there was no patient harm reported, and this device was replaced. | [5 empty] | The investigation into the low pump flow has been completed. the impella pump was returned for investigation. the returned pump was visually inspected, and a cannula kink was observed near the inflow. the cause of the issue was a kinked cannula. the cause of the kink was unknown. this report is being filed as part of a retrospective review of historical records.
18757 | 1220648-2024-19172 | 20364336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364336&pc=OZD | 07/21/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 07/21/2022 | 10/03/2024 |  |  | 04/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217321 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. shortly after beginning support, the optical sensor began displaying incorrect readings, and a "placement signal unavailable" alarm was produced. the pump was explanted and replaced, and no patient harm was reported. | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the optical signal issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18758 | 1220648-2024-19148 | 20364350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364350&pc=OZD | 04/11/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 04/11/2023 | 10/03/2024 |  |  | 11/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024308560 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced hematoma at the impella access site and the pump was explanted. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18759 | 1220648-2024-19164 | 20364359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364359&pc=OZD | 09/19/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 09/19/2022 | 10/03/2024 |  |  | 03/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2023192546 |  | 0048-0014 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient, demographics unknown, undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported the pump stopped during the night. reportedly the pump was repositioned and then was started again. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, evaluation of the device was not possible. the investigation into the pump stop issue has been completed. no product or logs were returned. the pump had flows of 2.0 l/min which is below the normal range of flows at p6. based on the clinical information, the cause of the pump stop was most likely ingested biomaterial. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi & impella rp with smart assist system section: using the automated impella controller with the impella catheter as noted in the impella ifu ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smart assist system catheter as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smart assist & impella cp with smart assist for use during cardiogenic shock section: troubleshooting the purge system as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped as noted in the impella ifu ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18760 | 1220648-2024-19151 | 20364364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364364&pc=OZD | 04/14/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 04/14/2023 | 10/03/2024 |  |  | 12/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318012 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp pump could not be fully advanced into the patient because of challenging anatomy; therefore, the insertion was aborted. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18761 | 1220648-2024-19176 | 20364370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364370&pc=OZD | 10/01/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/01/2022 | 10/03/2024 |  |  | 05/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228962 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. pump was not able to be advanced into left vertical patient anatomy and or wire position. the pump was removed, and new pump was opened and inserted without issue. no further information is available. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause of the delivery issue was patient condition related. this report is being filed as part of a retrospective review of historical records.
18762 | 1220648-2024-19166 | 20364382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364382&pc=OZD | 09/19/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 09/19/2022 | 10/03/2024 |  |  | 04/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223276 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 21 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a pump stop occurred during support, attempts to restart through troubleshooting were unsuccessful. it is unclear if the pump was replaced as it appears to have occurred at the end of support, which was 17 days. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, evaluation of the device was not possible. the investigation into the pump stop has been completed. the cause of the issue was not determined since the product was not returned. the pump stop occurred beyond the indication for use, at 17 days. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi & impella rp with smart assist system section: using the automated impella controller with the impella catheter as noted in the impella ifu ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smart assist system catheter as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smart assist & impella cp with smart assist for use during cardiogenic shock section: troubleshooting the purge system as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped as noted in the impella ifu ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18763 | 1220648-2024-19159 | 20364423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364423&pc=OZD | 05/06/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/06/2022 | 10/03/2024 |  |  | 08/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022129446 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/06/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was replaced due to low flows. medical doctor (md) could not resolve consistent low flows and waveforms were not able to get a flow performance of 2.5l at p5 without suction. the device was placed shallow, about 2cm. the md wanted to remove and place a 2nd device. no additional information was provided. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18764 | 1220648-2024-19160 | 20364425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364425&pc=OZD | 05/10/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 05/10/2022 | 10/03/2024 |  |  | 02/29/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193099 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp had a cannula kink. during insertion there were no issues placing the pump into the groin or across the valve. it is unknow what caused the kink. the pump was replaced due to kinked cannula. no injury or patient harm was noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18765 | 1220648-2024-19169 | 20364450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364450&pc=OZD | 09/27/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 09/27/2022 | 10/03/2024 |  |  | 04/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2023216354 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, ethnicity unknown, with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported the pump stopped and was unable to be restarted. the physician removed the impella, and the patient remained stable under extra-corporeal membrane oxygenation. there was no patient harm reported. no additional information is available. | [5 empty] | The investigation into the pump stop issue has been completed. the impella pump was returned for investigation. the product was visually inspected and dried blood was observed in the outflow cage. the pump was soaked in warm water for 5 minutes, after soaking, biomaterial was observed around the impeller. the root cause for the pump stop is biomaterial. impella cp with smart assist system instructions for use for the related event are as follows: as noted in the impella ifu: impella cp with smart assist for use during cardiogenic shock and high risk cpi & impella rp with smart assist system section: using the automated impella controller with the impella catheter as noted in the impella ifu ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smart assist system catheter as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
18766 | 1220648-2024-19179 | 20364451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364451&pc=OZD | 09/12/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 09/12/2022 | 10/03/2024 |  |  | 05/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225210 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the medical team observed that the patient was bleeding from the cp optical placement. when the access site of the patient was confirmed, the team fixed the sheath and remedied the issue. the sheath was 1cm protruding from the site. the fixation angle was adjusted, but the bleeding did not stop, so the stitches were once removed and the nose was attached to the skin for fixation, and then the bleeding stopped. it was reported that the patient was normally weaned and survived the procedure. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. no product was returned for investigation. based on the clinical information, the cause of the access site bleeding was use related. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18767 | 1220648-2024-19182 | 20364455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364455&pc=OZD | 10/07/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/07/2022 | 10/03/2024 |  |  | 04/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED. INC. |  |  | US |  | 2023210673 |  |  |  | Yes |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the medical staff observed that the device was exhibiting high purge pressure. the resolution for this issue is unknown, the device was normally weaned. no further information is available. | [5 empty] | The investigation for the reported high-pressure purge has been completed. the product was returned, and the high pressure was confirmed. per the results of the clinical review and investigation: the returned product was visually inspected, and a kink was observed in the cannula and biomaterial was observed in the purge gap and in the outflow. the cause of the issue was biomaterial. as noted in the impella ifu: impella cp with smart assist system impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
18768 | 1220648-2024-19165 | 20364479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364479&pc=OZD | 03/25/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 03/25/2022 | 10/03/2024 |  |  | 10/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158498 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant had a cp replaced due to low flows from a kink. there was no harm or injury from the interruption. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. no product was returned. data log showed "suction" alarms occurred upon initial pump activation, followed by a rapid decrease in motor current and pump flow to 0; flows eventually restored by the "suction" alarms persisted. after reviewing the clinical information, it was determined that the cause of the low pump flow was most likely a kinked cannula. this report is being filed as part of a retrospective review of historical records.
18769 | 1220648-2024-19178 | 20364498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364498&pc=OZD | 10/05/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/05/2022 | 10/03/2024 |  |  | 06/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235834 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. a pump stop issue with the impella device was reported. while attempting to transfer the patient to a new aic upon arrival at the hospital, an impella defective alarm appeared. attempts were made to switch back to the original aic but the pump continued to remain inactive. the pump was explanted, and the access site was tended to with manual pressure and perclose. there was no reported harm to the patient. no additional information is available. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. the cause of the issue was not determined since product was not returned. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
18770 | 1220648-2024-19168 | 20364503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364503&pc=OZD | 09/23/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 09/23/2022 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the physician had to re-insert the impella cp several times. during the process the device exhibited a placement signal failed error message. the physician decided to replace the catheter and the device was successfully placed. it was reported that the patient was normally weaned and survived the procedure. | [5 empty] | The investigation into the optical signal issue has been completed. the data logs were reviewed and confirm that the placement signal and optical sensor were functional at the start of the case. during the case there was a loss of placement signal and the five signal parameters were consistent with a broken optical fiber. it is likely the fiber was broken during the difficult insertion attempts. the root cause of the optical signal issue was most likely a broken fiber line due to excessive force during a difficult insertion through the patient's tortuous vasculature. the cause of the issue was a damaged optical fiber line. this report is being filed as part of a retrospective review of historical records.
18771 | 1220648-2024-19183 | 20364508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364508&pc=OZD | 10/10/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/10/2022 | 10/03/2024 |  |  | 07/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250493 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp had low pump flow with a kink in the catheter. the physician was unable to manually manipulate catheter. the pump was removed and replaced with another impella cp with no further issue. no further information was available. | [5 empty] | The investigation for the reported low pump flow and cannula kink has been completed. the product was returned, and the cannula kink was confirmed. per the results of the clinical review and investigation: the rt log confirms low pump flow. a visual inspection of the pump revealed a kink in the cannula. no other abnormalities were observed. the cause of the low pump flow was a kinked cannula. this report is being filed as part of a retrospective review of historical records.
18772 | 1220648-2024-19173 | 20364545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364545&pc=OZD | 09/30/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 09/30/2022 | 10/03/2024 |  |  | 08/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022131985 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported a pump stopped during the use of the device. the pump was removed and replaced. no patient injury or harm noted, no further information is available. | [5 empty] | The investigation for the reported pump stop issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause of the pump stop. the reported issue was use related as the guidewire was not removed before starting the pump. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
18773 | 1220648-2024-19174 | 20364548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364548&pc=OZD | 09/30/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 09/30/2022 | 10/03/2024 |  |  | 06/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236292 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was going to be implanted with an impella cp device for mechanical circulatory support. a delivery issue with the impella device was reported. pump placement was attempted during ongoing patient cardiopulmonary resuscitation (cpr) and defibrillation; however, the wire was pulled too early which contaminated the device. the pump insertion was aborted because of the contamination. no patient harm or injury noted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause of the delivery issue was use related. this report is being filed as part of a retrospective review of historical records.
18774 | 1220648-2024-19175 | 20364550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364550&pc=OZD | 10/01/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/01/2022 | 10/03/2024 |  |  | 06/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249519 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the device exhibited low pump flow. normal troubleshooting attempts were unsuccessful in resolving the issue. the pump was removed and intra-aortic balloon pump (iabp) was placed. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has been returned for evaluation. the log confirms low pump flow and suction alarms. a visual inspection of the pump did not reveal any abnormalities. the pump was attached to a console and run, and the low flow could not be reproduced. the cause of the low pump flow was most likely patient condition related. this report is being filed as part of a retrospective review of historical records.
18775 | 1220648-2024-19184 | 20364556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364556&pc=OZD | 03/01/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 03/01/2023 | 10/03/2024 |  |  | 09/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275746 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the implant was unable to pass over the arch and across the aortic valve due to the frozen elephant trunk graft in the patient. the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient frozen elephant trunk anatomy, most likely due to patient condition. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18776 | 1220648-2024-19177 | 20364559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364559&pc=OZD | 10/05/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/05/2022 | 10/03/2024 |  |  | 08/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022126412 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the impella cp device could not pass the femoral artery, due to stent in place and peripheral arterial occlusive diseases (pavk) as well. delivery attempts were made although contraindications present. it was decided to replace the device. patient was reported to be stable; no further information is available. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was patient condition related, specifically the patient's calcified vessels. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18777 | 1220648-2024-19186 | 20364566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364566&pc=OZD | 03/01/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 03/01/2023 | 10/03/2024 |  |  | 11/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295716 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to being implanted with an impella cp device for mechanical circulatory support. the complainant reported a failure to deliver the device due to patient's anatomy. as a result of this complication; therefore, the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient's had iliac lesion. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18778 | 1220648-2024-19188 | 20364576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364576&pc=OZD | 04/18/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 04/18/2023 | 10/03/2024 |  |  | 12/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320273 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there was high purge pressure and one day later impella cp had a pump stop with high motor current. the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18779 | 1220648-2024-19185 | 20364580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364580&pc=OZD | 04/16/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 04/16/2023 | 10/03/2024 |  |  | 10/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268569 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the impella stopped on the 13th day of support. normal troubleshooting attempts to restart the device were unsuccessful. the impella cp was explanted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the device was not received for evaluation. the root cause of the pump stop was unable to be determined as no product or logs were returned for analysis. this report is being filed as part of a retrospective review of historical records.
18780 | 1220648-2024-19181 | 20364581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364581&pc=OZD | 10/06/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/06/2022 | 10/03/2024 |  |  | 03/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204833 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. a kinked cannula was reported after delivering the device into calcified vasculature. low pump flows were exhibited; therefore, the pump was removed and replaced with another impella cp. no patient harm or injury noted, no additional information is available. this report is being filed as part of a retrospective review of historical records. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the device was returned, and the cannula kink was confirmed. visual inspection revealed a cannula kink near the outflow cage. no other damage or abnormalities were found. per the results of the clinical review and investigation: data logs showed there were persistent "suction" alarms and low flows after the pump was activated. visual inspection revealed a cannula kink near the outflow cage. no other damage or abnormalities were found. therefore, it was determined that the cause of the low pump was a kinked cannula. after reviewing the clinical information, it was determined that the cause of the delivery issues was patient condition related, specifically the patient's calcified vessels. this report is being filed as part of a retrospective review of historical records.
18781 | 1220648-2024-19190 | 20364582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364582&pc=OZD | 10/11/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/11/2022 | 10/03/2024 |  |  | 05/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229094 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Fracture |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the physician placed the impella cp per protocol, and when the physician removed the abiomed 0.018 wire from the body, the distal portion of the wire broke off inside the peel-away sheath and was retained in the sheath. this issue was visible under fluoroscopy. the physician exchanged the sheath for the unused peel-away and reinserted the pump per protocol using a replacement platinum plus 0.018 wire. it was confirmed that no fragments from the wire was present in the patient's body. additionally, during beside repositioning attempts the ao waveform became erratic. teh medical staff observed that the left ventricle waveform became very small and sometimes non-existent. also, flows dropped dramatically, and a suction alarm occurred. the pump was put on p-0 range for several seconds and then restarted, this action alleviated the waveform and flow issues. however, 10 minutes, the same problem occurred. the physician decided to explant the pump and replace it with another impella cp. it was reported that the replacement was performed without issue and the patient remained stable. | [5 empty] | The investigation for the reported delivery issue has been completed. the product was returned and confirmed that the guidewire was broken. the distal portion of the guidewire was not returned. the pump was returned and examined. no abnormalities were found. the pump was run, and flows were within specification. the root cause of the guidewire damage was determined to be use related. the root cause of the low flows was determined to be patient condition as the pump preformed without issue when returned. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18782 | 1220648-2024-19187 | 20364583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364583&pc=OZD | 04/17/2023 | 10/03/2024 |  | 10/03/2024 |  |  | 04/17/2023 | 10/03/2024 |  |  |  |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 0048-0045 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the impella was unable to advance as physician felt the pigtail was entrapped in papillary muscle. it was also noted that patient went into atrial fibrillation. one unit of packed red blood cells was given the following morning. the impella was explanted and replaced with a new impella cp. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. d4 is unknown. this report is being filed as part of a retrospective review of historical records.
18783 | 1220648-2024-19191 | 20364589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364589&pc=OZD | 10/11/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/11/2022 | 10/03/2024 |  |  | 04/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217493 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported a delivery issue and cannula kink with the impella device. the pump insertion was aborted and a new impella was used without further issue. no further information was available. | [5 empty] | The investigation for the reported delivery issue has been completed. no damage or abnormalities were found on the returned pump. after reviewing the clinical information, it was determined that the cause of the delivery issues was patient condition related, specifically the patient's diseased vessels. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18784 | 1220648-2024-19189 | 20364648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364648&pc=OZD | 09/12/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 09/12/2022 | 10/03/2024 |  |  | 04/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215635 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported there was a purge alarm that the medical team was unable to resolve. reportedly the impella was removed and the team unsuccessfully attempted to re-prime the pump. the decision was made to remove the pump and replace the pump with another impella, which primed and was placed without issue. | [5 empty] | The investigation into the priming issue has been completed. the impella pump was returned for investigation. the product was visually inspected, and no abnormalities were observed. the pump was connected to an automated impella controller and purge fluid was observed to be exiting the purge gap. per data log review, the purge alarm observed was "open purge" indicating that the cause of the alarm was likely an improper purge connection since the returned product functioned as intended. the root cause of the issue was improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
18785 | 1220648-2024-19192 | 20364657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364657&pc=OZD | 09/13/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 09/13/2022 | 10/03/2024 |  |  | 04/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2023212878 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the decision was made to reposition the pump due to low p ump flow. the pump was repositioned under echocardiogram and fluoroscopy, reportedly during repositioning the pump was dislocated in the aorta, and repositioning was not possible. the device was removed because it had been running for more than two hours. a new impella was placed. there was no patient harm reported. | [5 empty] | The investigation into the low pump flow has been completed. the impella pump was returned for investigation. the data logs were reviewed and confirmed low flow and suction throughout the case. the product was returned, and slight damage was seen on the cannula. no biomaterial was seen. the pump was run in a bucket of water and and were unable to reproduce low flow. based on the clinical details the root cause of the low pump flow was most likely improper technique by the end user when positioning the device. this report is being filed as part of a retrospective review of historical records.
18786 | 1220648-2024-19193 | 20364723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364723&pc=OZD | 10/11/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/11/2022 | 10/03/2024 |  |  | 07/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported blood dripping from the pump's red plug after initiation. the pump continued to be used without any harm to the patient. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported product damage has been completed. the returned product was visually inspected and a hole in the catheter at the 83 cm mark was visually inspected. the cause of the issue was blood ingress due to a hole in the catheter resulting from improper use. this report is being filed as part of a retrospective review of historical records.
18787 | 1220648-2024-19194 | 20364727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364727&pc=OZD | 10/12/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/12/2022 | 10/03/2024 |  |  | 05/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218215 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported sudden pump stop in the morning (runtime day eight). several attempts to restart failed, the pump was successfully weaned, and the device was explanted. no further information is available. | [5 empty] | The investigation for the reported issue has been completed. the lt log confirms high motor current pump stops. purge pressure high and purge system blocked alarms also occurred the day before the pump stop. a visual inspection of the pump revealed a large impeller purge gap. no other abnormalities were observed. the cause of the pump stop was bearing wear. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18788 | 1220648-2024-19195 | 20364733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364733&pc=OZD | 10/13/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/13/2022 | 10/03/2024 |  |  | 07/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237996 |  | 0048-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported suction issues after pump initiation. a kink was seen under fluoroscopy; therefore, the pump was removed and replaced due to the low flows. no further information is available. | [5 empty] | The investigation for the reported low flow issue has been completed. product was visually inspected, and a kink was observed on the cannula. it did not appear that excessive force was used on the cannula. root cause is a kinked cannula. the field log was reviewed and suction with low pump flow occurred upon pump activation. no motor current spike was observed. the cannula kink likely occurred during insertion. this report is being filed as part of a retrospective review of historical records.
18789 | 1220648-2024-19196 | 20364735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364735&pc=OZD | 10/13/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/13/2022 | 10/03/2024 |  |  | 06/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237996 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant had a cp pump stop shortly after insertion. the team replaced the cp without any harm or injury from case support initiation delay. the cp was not a full support pump. no further information is available. | [5 empty] | The investigation for the reported issue has been completed. product was visually inspected. the red lumen was observed to be still inside the outflow cage. root cause is use related. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18790 | 1220648-2024-19197 | 20364754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364754&pc=OZD | 10/14/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/14/2022 | 10/03/2024 |  |  | 06/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023242199 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. sixteen minutes after impella support started, the catheter was found to be kinked near impeller with suboptimal flows of 0.9l/min observed. successfully replaced pump before the procedure. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. no product was returned. the rt log confirms low pump flow and shows a suction alarm occurring throughout the duration of support. the cause of the low pump flow was most likely a cannula kink based on the provided clinical information. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18791 | 1220648-2024-19198 | 20364767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364767&pc=OZD | 10/14/2022 | 10/08/2025 |  | 10/03/2024 |  |  | 10/17/2022 | 10/03/2024 |  |  | 07/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110369 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported tortuous anatomy that made placement difficult, and the long sheath was kinked during implant. the impella cp device was removed to exchange for short introducer and the purge system open alarm sounded. a diagnostic catheter was performed and revealed that the patient has severe triple vessel disease. the physician decided to replace the device with another impella cp. the replacement device was successfully placed. no further information is available. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smart assist system section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | The returned pump was connected to a lab purge cassette, system was purged, and no open purge or purge disc failure alarms triggered. the purge values were within specification at 5.8 ml/hr and 577 mmhg. the cause of the purge issue was not determined since the purge cassette was not returned and lack of clinical information. the issue was unable to be reproduced on the returned pump and the clinical reported that replacing the aic did not resolve the issue. delivery issue/introducer damage: this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-19198. g1 reporting contact email revised on manufacturer device report 1220648-2024-19198 in accordance with updated procedures.
18792 | 1220648-2024-19199 | 20364768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364768&pc=OZD | 10/15/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/15/2022 | 10/03/2024 |  |  | 04/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215043 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a delivery issue with the impella device. upon successful insertion, the pump failed to cross the aortic arch due to extensive calcium. after multiple failed attempts, the catheter kinked, and the pump was removed. a new impella cp pump was subsequently implanted without further issue. there was no reported harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. data logs indicate a kink in the purge line causing sudden high purge pressure and purge system blocked alarms (attached). analysis of the returned product revealed a large catheter kink and kinked purge line around the 43 cm marking. based on the clinical account that the patient had a highly calcified aorta, the root cause of the high purge pressure is due to a kinked purge line as a result of patient condition. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18793 | 1220648-2024-19200 | 20364770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364770&pc=OZD | 10/20/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/20/2022 | 10/03/2024 |  |  | 12/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178618 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, it was reported that there was a leaking from the purge sidearm after changing the purge composition from sodium bicarbonate to heparin. the location of the leaking was not specified, however no obvious damage or abnormalities were observed on the sidearm or purge cassette. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of purge leak - pump. the device was not received for evaluation. the cause of the purge leak was not determined since no product was returned. the issue occurred beyond the pump's 8-day design life as noted in the impella ifu: rp smart assist system with automated impella controller section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (eg, infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.¿ this report is being filed as part of a retrospective review of historical records.
18794 | 1220648-2024-19201 | 20364777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364777&pc=OZD | 10/20/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/20/2022 | 10/03/2024 |  |  | 07/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253079 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during placement of the device, a torn sterile sleeve occurred. the impella was explanted and replaced with a new impella cp. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of product damage (sterile sleeve damage). the device was not received for evaluation. the root cause of the product damage (sterile sleeve damage) was unable to be determined as no product was returned for analysis. this report is being filed as part of a retrospective review of historical records.
18795 | 1220648-2024-19202 | 20364786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364786&pc=OZD | 10/22/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/22/2022 | 10/03/2024 |  |  | 05/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228576 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on day 12 of support, the physician observed a pump stop due to mc high alarm. several attempts were made to restart the device, none were successful. the physican decided to explant the device. it was reported that the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the device was not received for evaluation. the root cause of the pump stop was unable to be determined as no product was returned for analysis. it was noted that the pump stop occurred beyond the 8-day impella cp indication for use. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18796 | 1220648-2024-19204 | 20364981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364981&pc=OZD | 10/23/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/23/2022 | 10/03/2024 |  |  | 07/31/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252368 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. before inserting the impella there was a ¿placement signal unreliable¿ alarm. the impella was started, and the motor current appeared normal, but the placement signal waveform (ps) was extremely high, and the left ventricle signal appeared more like an aortic pressure waveform but with its value significantly lower than the ps. the flows also were very higher than the average flows on the cp. the average flows are reported to be > 7 l/min. the medical staff was concerned about the alarm and values, the device was explanted and a replacement impella cp was placed. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of optical signal. the device was returned for evaluation. the returned pump was visually inspected. cannula and catheter kinks were observed. the pump was connected to the aic in the lab and ran, no alarms occurred. the placement signal not reliable alarm was not reproduced. the saturated flow was also not reproduced. the pump ran within specification. the 5s parameters were tested in the lab and found to be within normal range. the rt log confirms a ps not reliable alarm. the alarm resolves once the pump is started. the rt log also confirms the saturated flow and shows that there was no heparin used. the cause of the optical signal issue was most likely biomaterial interacting with the sensor since the 5s parameters are unaffected and the ps not reliable alarm resolves once the pump is started. a pump tear down was completed, and biomaterial was found in the motor. the cause of the saturated flow was biomaterial. this report is being filed as part of a retrospective review of historical records.
18797 | 1220648-2024-19203 | 20364983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20364983&pc=OZD | 10/23/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/23/2022 | 10/03/2024 |  |  | 06/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249440 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, flows were low and suction alarms occurred. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was returned for evaluation. product was visually inspected, and buildup was observed around the impeller blade. product was soaked in water. no observable buildup around the impeller. cannula was inspected and a kink was observed. patient was diagnosed with mitral valve disease. root cause is most likely patient condition. this report is being filed as part of a retrospective review of historical records.
18798 | 1220648-2024-19205 | 20365000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20365000&pc=OZD | 10/25/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 10/25/2022 | 10/03/2024 |  |  | 06/30/2024 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247265 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, the device exhibited flow saturation. the physician decided to explant the device and replace it with a new impella cp. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of saturated flow. the device was not received for evaluation. returned data logs (attached) show the pump operating within specification for the first 40 minutes of support. suddenly, there is a large jump in motor current and a slight drop in motor speed, which immediately causes saturated flow and a subsequent rise in purge pressure. without the returned pump, the cause of the flow saturation cannot be determined. this report is being filed as part of a retrospective review of historical records.
18799 | 1220648-2024-19206 | 20365133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20365133&pc=OZD | 01/14/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 01/14/2022 | 10/03/2024 |  |  | 10/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022155385 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. while on support, the catheter seemed to have less torque and structure then it normally does. the impella was ramped up to pressure level nine with flow of 3.2 liters per minute. reportedly, during attempts to reposition the jr4 catheter hooked onto the impella catheter inadvertently causing the impella to be pulled entirely out of the left ventricle. multiple unsuccessful attempts were made to push the pump back across the valve into the aorta. the impella cp was eventually explanted and the impella long sheath was inserted and a replacement impella cp was prepped for insertion. no patient harm was reported. | [5 empty] | The investigation into the impella cp positioning issue has been completed. the impella pump was returned for investigation. the investigation revealed a mangled jr4 catheter, which was found fed into the introducer alongside the impella. there were several kinks along the length of the impella catheter. a cannula kink was also found near the pump's motor housing. no other damage or abnormalities were found. after reviewing the clinical information, it was determined that the root cause of the positioning issues was use related, specifically the improper interaction between the impella and jr4 catheter that led to the impella being pulled out of the left ventricle. the "loss" of mobility was due to the catheter kinks, which resulted from repeated wireless re-access attempts. this report is being filed as part of a retrospective review of historical records.
18800 | 1220648-2024-19207 | 20365145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20365145&pc=OZD | 01/17/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 01/17/2022 | 10/03/2024 |  |  | 05/31/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089328 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during placement of the impella, the cannula was noted to be kinked when advancing across the aortic valve (av), causing low flows and suction events when placed in pressure level mode. the impella was removed and replaced with another impella cp pump. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18801 | 1220648-2024-19208 | 20365148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20365148&pc=OZD | 01/19/2022 | 10/03/2024 |  | 10/03/2024 |  |  | 01/19/2022 | 10/03/2024 |  |  | 09/30/2023 |  | 10/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140777 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, race unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the impella was pulled into the ascending aorta while placing a bio prosthetic valve. the surgeon attempted multiple times to cross the aortic valve (av) but kept getting stuck on the valve apparatus. the surgeon decided to remove the pump and replaced it with a new impella cp pump. no patient harm was reported. | [5 empty] | The investigation into the positioning issue has been completed. the impella pump was returned for investigation. the returned pump was visually inspected, and no abnormalities were observed. no cannula kinks were observed. the root cause indicated the impella device is not indicated for wireless re-access, resulting in improper technique and use related. this report is being filed as part of a retrospective review of historical records.
18802 | 1220648-2024-18846 | 20350737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20350737&pc=OZD | 09/07/2024 | 11/14/2024 |  | 10/02/2024 |  |  | 09/07/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525220 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and their hemoglobin continued to drop despite giving blood products. the doctor wanted to remove the impella cp as they thought the pump was causing harm to the patient. it was thought that the impella leg (right) was bleeding internally and externally. the right leg was addressed and flows were present with no apparent signs of bleeding internally. the access site was addressed using a side closure technique with two perclose used. it was later reported that no bleeding or ischemia was due to the impella and the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. the root cause of access site bleeding was determined to be patient condition because the bleeding happened from non-impella site. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18803 | 1220648-2024-18847 | 20350742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20350742&pc=OZD | 09/09/2024 | 11/11/2024 |  | 10/02/2024 |  |  | 09/11/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525111 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The medical center had a st elevation myocardial infarction patient admitted. culprit vessel treated with a balloon inflation and then the patient coded. impella cp attempted to be placed during cardiopulmonary resuscitation. the medical doctor was unable to cross atrioventricular despite multiple attempts. the impella cp placement aborted, return of spontaneous circulation was never obtained and the patient expired. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because of the stenosis in the vessels and the cp was unable to implanted. added- h6 med dev prob code 2682. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci, section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿.
18804 | 1220648-2024-18849 | 20350754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20350754&pc=OZD | 09/04/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/04/2021 | 10/02/2024 |  |  | 06/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022103343 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were unresolvable suction alarms during use of this device. the pump was removed and replaced. there was no patient harm reported. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18805 | 1220648-2024-18848 | 20350759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20350759&pc=OZD | 09/07/2024 | 11/12/2024 |  | 10/02/2024 |  |  | 09/07/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025503051 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was to be implanted with an impella 5.5 implant but the insertion of the wire was very difficult through the right subclavian and into the aortic root. the wire made it into the left ventricle and the pump could not advance through the innominate artery due to patient anatomy. the implant was aborted and no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related. the impella was unable to pass through the vasculature as per the clinical description provided.
18806 | 1220648-2024-18850 | 20350767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20350767&pc=OZD | 09/07/2024 | 11/13/2024 |  | 10/02/2024 |  |  | 09/07/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532307 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after an impella cp was placed, it was unable to obtain flows greater than 1.5 liters per minute despite confirmed positioning. the impella was removed and a new cp was inserted. no patient harm has been reported. | [5 empty] | A.5 is unknown. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the low pump flow has been completed. the complaint device was not returned for analysis. data log was reviewed and founda motor current spike observed after ten minutes of impella running followed by a drop in flow. positioning issue alarms were observed that resolved with the trouble shooting. the root cause of the low pump flow issue most likely was biomaterial ingestion.
18807 | 1220648-2024-18852 | 20350838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20350838&pc=OZD | 09/11/2024 | 11/12/2024 |  | 10/02/2024 |  |  | 09/11/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531671 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support due to the patient having ventricular tachycardia (vt) ablation pre-ablation. after the impella was placed, the pump positioning was was too deep then too shallow. upon repositioning, the pump got caught on mitral valve apparatus and had to be explanted. care team prepped device in bowl of hep saline then readvanced into patient, turned on with good flows of 3.4lpm at p-8. the patient then became hypotensive due to the induced vt. the impella had a motor alarm and the impella flows dropped from 3.8 down to .7. position was confirmed and the staff tried to adjust the p level with no success therefore the pump was explanted. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the low pump flow and the positioning has been completed. pump was returned for investigation. pump was soaked in water and there was heavy biomaterial found in the cannula and around the impeller. clinical mentions that there was a motor alarm which was resolved but flows were reduced from 3.8l/m to 0.7l/m and lowering p-levels was of no help. therefore, the root cause of the low pump flow issue was biomaterial. data logs were not returned for investigation. clinical mentions explicitly that pump was initially shallow and deep. after readvancing the pump, it was caught in the mitral valve apparatus and had to be explanted. therefore, the root cause of the positioning issue was use related to improper technique. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
18808 | 1220648-2024-18851 | 20350840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20350840&pc=OZD | 09/04/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/04/2021 | 10/02/2024 |  |  | 03/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022068117 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the physician placed the cp, and the pump had lower than optimal pump flows. the impella was removed, and a new impella pump was successfully placed with good flows. there was no patient harm reported. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs from the field were returned and confirmed that the pump only produced 0.1l/min of flow. a motor current spike can be seen halfway through the run time. the pump was returned, and no biomaterial was found in the cages, cannula, or around the impeller. the pump ran within specification. the root cause of the low flow was determined to not be a product malfunction as the product operated within specification within the lab. the root cause of the low flow was most likely a result of ingested biomaterial that was removed during explant as the high motor current is indicative of biomaterial interacting with the impeller. this report is being filed as part of a retrospective review of historical records.
18809 | 1220648-2024-18853 | 20350901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20350901&pc=OZD | 09/05/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/05/2021 | 10/02/2024 |  |  | 05/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022085756 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the pump was unable to pass through the iliac artery due to the patient's anatomy. the pump and sheath were removed and a long sheath inserted. the pump was replaced. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. clinical details provided were sufficient to determine a root cause. the reported delivery issues were determined to be related to the patient¿s condition. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18810 | 1220648-2024-18854 | 20350910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20350910&pc=OZD | 03/31/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 03/31/2022 | 10/02/2024 |  |  | 01/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174007 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon the transitioning from the 14fr to the repositioning sheath, an observation was made of a kinked cp cannula. the pump was explanted and replaced, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported cannula kink has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was most likely a catheter kink from excessive force. this report is being filed as part of a retrospective review of historical records.
18811 | 1220648-2024-18855 | 20350916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20350916&pc=OZD | 09/07/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/07/2021 | 10/02/2024 |  |  | 10/31/2022 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022042351 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the team noted a kink while explanting the 14fr peel away. the pump was replaced by extracorporeal membrane oxygenation (ecmo). there was no patient harm reported. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump was visually inspected, and the inflow cage was found to be detached. the delo was intact and the cannula resin was torn near the delo. a cannula kink was confirmed at the bend. the root cause of the cannula kink was determined to be that the impella device is not indicated for wireless re-access/improper technique. this report is being filed as part of a retrospective review of historical records.
18812 | 1220648-2024-18856 | 20350939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20350939&pc=OZD | 09/06/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/06/2021 | 10/02/2024 |  |  | 02/28/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022070798 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella was removed to restore cardiac function when a kink on the catheter shaft was noted. there was no health hazard impact. there was no particular problem with the insertion procedure and the patient's condition was stable after removal. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined as no clinical abnormalities were noted. this report is being filed as part of a retrospective review of historical records.
18813 | 1220648-2024-18857 | 20350990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20350990&pc=OZD | 09/11/2024 | 11/06/2024 |  | 10/02/2024 |  |  | 09/11/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1160668 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) lost impella connect and will not reactivate while patient was on support. the blue and yellow airplane light continued to blink simultaneously. the aic was changed when new cp was placed. no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) ¿ issue has been completed. after review of data log and device analysis, the cause of the issue was not determined since issue could not be reproduced
18814 | 1220648-2024-18859 | 20351039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351039&pc=OZD | 09/06/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/06/2021 | 10/02/2024 |  |  | 06/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022106003 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, the cp was not able to advance the guidewire using the easy guide lumen during case preparation. a new pump was used without issue. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issues issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the red lumen was still in place threaded through the pigtail. the impella was coated in blood suggesting that it had been placed in the body or a significant amount of blood had gotten on the pump during prep. a sample guidewire was used from the lab and the pigtail was used to backload the impella without any issues. the root cause of the backloading issues was most likely use related as the issue could not be reproduced. no guidewire was returned for this investigation. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18815 | 1220648-2024-18858 | 20351087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351087&pc=OZD | 09/06/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/06/2021 | 10/02/2024 |  |  | 02/28/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022051850 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella stopped suddenly, at the end of day four, and was unable to restart. they switched the console and the pump still would not start. the pump was explanted. there was no patient harm reported. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that no logs were provided for analysis. the returned console showed that the pump would not restart after being connected to the replacement console and triggered 'impella stopped' alarm 119. a kink in the catheter was noted at the 85 mark. electrical testing confirmed an open line. the pump would not run on a test console. the root cause of the pump stop was determined to be an open line caused by the catheter kink. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18816 | 1220648-2024-18860 | 20351096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351096&pc=OZD | 09/09/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/09/2021 | 10/02/2024 |  |  | 06/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022107846 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a hematoma in cardiac catheterization laboratory which was addressed by placement of suture. bleeding around the repositioning sheath was also noted. pump was removed due to the bleeding, and manual compression was applied to achieve hemostasis. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on the clinical account, the root cause of the access site bleeding was due to improper anticoagulation, as the physician mentioned the use of integrillin may have exacerbated the bleeding issues. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records.
18817 | 1220648-2024-18862 | 20351122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351122&pc=OZD | 09/13/2024 | 11/06/2024 |  | 10/02/2024 |  |  | 09/13/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536665 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had placed a cp for the support of the patient admitted in cardiogenic shock. the pump was placed and the site of insertion had a hematoma develop. the bleed was treated by one unit of red blood cells, no anticoagulation and the addition of the sodium bicarbonate rather than heparin in the purge. the patient survived the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 outcomes attributed to adverse event missing from manufacturer device report. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
18818 | 1220648-2024-18861 | 20351149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351149&pc=OZD | 09/12/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/12/2021 | 10/02/2024 |  |  | 07/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110499 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the aorta (ao) and left ventricle (lv) signals were not correlating at all with patient's hemodynamics. the physician acknowledged she could have damaged the sensor on insertion and the pump was exchanged. the new pump worked without issue, and there was no harm to the patient. | [5 empty] | The investigation for the reported placement signal issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18819 | 1220648-2024-18865 | 20351194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351194&pc=OZD | 09/10/2024 | 10/22/2024 |  | 10/02/2024 |  |  | 09/10/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025509747 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, left ventricle perforation was noted during coronary artery bypass graft. impella 5.5 was explanted and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the ventricle perforation has been completed. no product was returned for analysis. based on the provided clinical details, the root cause of the ventricle perforation is most likely due to use. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
18820 | 1220648-2024-18863 | 20351195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351195&pc=OZD | 03/31/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 03/31/2022 | 10/02/2024 |  |  | 11/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169876 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device was pulled back into the aorta for hours during coronary artery bypass graft surgery due to significant left ventricle (lv) distention that was unresolved with p1/p2 venting. the device was unable to safely traverse back into lv once grafts were placed. the pump was then removed and replaced with another impella cp. no patient harm was reported. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the root cause of the positioning issue was determined to be related to the patient condition. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ this report is being filed as part of a retrospective review of historical records.
18821 | 1220648-2024-18866 | 20351219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351219&pc=OZD | 09/14/2024 | 11/07/2024 |  | 10/02/2024 |  |  | 09/14/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531590 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 55-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. the pump was explanted after two days of support due to unresolved suction/low flows. the team attempted a 5.5 pump delivery but due to anatomy another, second impella cp was instead placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. impella cp returned for evaluation. upon visual inspection, biomaterial observed around impeller. no broken blade or cannula kink observed. no other abnormality. data logs were reviewed and lt logs from case show that pump was unable to flow higher than ~2.0l/min when turned down to p-6. suction alarms also present at that time. the cause of the low pump flow was biomaterial. d9 date device returned to manufacturer added. e4 should have been left blank on manufacturer device report 1220648-2024-18866.
18822 | 1220648-2024-18864 | 20351239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351239&pc=OZD | 09/10/2024 | 11/05/2024 |  | 10/02/2024 |  |  | 09/10/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471202 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. the surgeon believed the outflow of the 5.5 was causing the aortotomy to bleed. the impella 5.5 was explanted and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-18864.
18823 | 1220648-2024-18867 | 20351242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351242&pc=OZD | 09/11/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/14/2021 | 10/02/2024 |  |  | 03/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022059894 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  |  | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during the process of implanting the impella pump, a hole in the pigtail catheter was observed. after numerous efforts to advance, the original impella device was explanted and replaced. there was no patient harm reported. | [5 empty] | The investigation for the reported hole in catheter issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that an indent was found on both the pigtail and the lumen indicating that the guidewire had pressed into them with significant force. the root cause of the case start issue was most likely a use related as this can occur when the pigtail and lumen are not sufficiently straightened during the backloading of the guidewire. this report is being filed as part of a retrospective review of historical records.
18824 | 1220648-2024-18868 | 20351280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351280&pc=OZD | 04/01/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/01/2022 | 10/02/2024 |  |  | 12/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181503 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella cp device for mechanical circulatory support. during preparation, the device produced a "defective catheter" alarm when the pump was connected. the pump implant was aborted, and it was reported that there was no patient harm. | [5 empty] | The investigation for the reported placement signal issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18825 | 1220648-2024-18869 | 20351360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351360&pc=OZD | 09/13/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/10/2021 | 10/02/2024 |  |  | 07/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022103886 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the left ventricle was small and therefore was not inserted. the patient showed signs and symptoms of hemolysis. as a result, the auxiliary level was lowered. | [5 empty] | The investigation for the reported hemolysis issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on the clinical account, the root cause of the suspected hemolysis was due to patient condition, as the left ventricle was small causing suction. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
18826 | 1220648-2024-18870 | 20351392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351392&pc=OZD | 04/02/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/02/2022 | 10/02/2024 |  |  | 10/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158496 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump unexpectedly stopped after 16 days on support due to high motor current. the pump was removed, and no patient harm was reported. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the data logs show a sudden, sharp jump in mc corresponding to saturated flows and a pump stop after 16 days of use. although product was not returned, based on the motor current jump in the data logs the root cause of the pump stop is most likely bearing wear after 16 days. this is beyond the cp optical design life of 8 days. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18827 | 1220648-2024-18871 | 20351396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351396&pc=OZD | 09/13/2024 | 11/26/2024 |  | 10/02/2024 |  |  | 09/13/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534022 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Pump | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | In an attempt to implant an impella cp, the automated impella controller (aic) displayed a retrograde flow alarm even though the pump had not yet been started or implanted. aic was changed and the issue persisted. when the pump was connected, the alarm pump stopped, retrograde flow appeared. a new pump was used. no patient harm has been reported. | [5 empty] | The device has been received for evaluation. upon completion of the investigation, a supplemental mdr will be filed. | The investigation into pump stop has been completed. the product was returned and evaluated. no abnormalities were found with the pump. the pump ran in a bucket of water without complications. no logs were returned, so the timing and cause of alarms could not be determined. the root cause of the pump not starting was not determined since logs were not returned and insufficient clinical detail was provided. issue was unable to be reproduced on the returned product. the failure mode of pump stop was updated to pump did not start. instruction for use as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act greater than 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention with automated impella controller section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ h10 instructions for use was missing on manufacturer device report 1220648-2024-18871
18828 | 1220648-2024-18872 | 20351428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351428&pc=OZD | 09/07/2024 | 11/13/2024 |  | 10/02/2024 |  |  | 09/07/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025484294 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Embolism/Embolus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2024 |  |  |  |  | The complainant reported a patient on impella 5.5 support had increased bleeding from the impella 5.5 site from the jackson-pratt drain. this was managed by holding systemic heparin and placing a sandbag on the left chest for about an hour. the patient continued with oozing and one unit of packed red blood cells was given. it was further reported that the impella 5.5 was exchanged for concerned of micro-emboli on the device possibly contributing to hemolysis. the physician stated that there was some clot on the catheter on explant but said there was no way of knowing whether it was picked up coming back through the graft. the patient is stable. | [5 empty] | The investigation has been completed. the data logs showed the pump ran without issue during support. there were multiple alarms triggered but resolved quickly. the product was not returned for review. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. as the necessary information was not provided, the root cause of the thrombus was not determined f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-18872 and should not have been. | A.5 ethnicity is unknown. the impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
18829 | 1220648-2024-18873 | 20351508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351508&pc=OZD | 09/13/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/13/2021 | 10/02/2024 |  |  | 06/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022098172 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were two instances of a pump stop that occurred. each stop was successfully restarted; however, the decision was made to replace the pump. there was no adverse impact to the patient. | [5 empty] | The investigation for the reported temporary pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump was visually inspected, and no damage was found to the catheter such as a kink; however, blood was noted under the impeller. the pump was plugged into a console and ran at p-9 during which a brief impella stopped - motor current high alarm was triggered. the pump stopped briefly but restarted on its own. it was noted that purge flow and motor current were high throughout the lab run. further inspection of the bearing under a microscope revealed evidence of blood ingress into the motor. therefore, the root cause of the pump stops was determined to be blood ingress; however, as no data logs were provided, the cause of the blood ingress could not be determined. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18830 | 1220648-2024-18875 | 20351610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351610&pc=OZD | 04/07/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/07/2022 | 10/02/2024 |  |  | 12/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167873 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the device had lower than expected flow. after percutaneous coronary intervention (pci) was completed, the position was adjusted again. however, there was no pulse pressure in the motor waveform. the pump was removed, and biomaterial was found in the device. no patient harm or injury reported. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the root cause of the low pump flow was determined to be most likely biomaterial. this report is being filed as part of a retrospective review of historical records.
18831 | 1220648-2024-18874 | 20351626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351626&pc=OZD | 04/05/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/05/2022 | 10/02/2024 |  |  | 12/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163599 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, a kinked cannula with low pump flow was noted. the issue was resolved with removing and replacing the impella cp device. no patient harm or injury reported. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the root cause of the low pump flow was determined to be a kinked cannula. this report is being filed as part of a retrospective review of historical records.
18832 | 1220648-2024-18877 | 20351723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351723&pc=OZD | 09/13/2024 | 10/11/2024 |  | 10/02/2024 |  |  | 09/13/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025491651 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information received that the patient developed renal failure while on support. continuous renal replacement therapy (crrt) was started on (b)(6)2024 as treatment. the kidneys recovered on (b)(6)2024 and the crrt was no longer necessary and the patient recovered with no sequelae. | Us complainant reported the patient was on impella 5.5 support and the patient went into ventricular tachycardia for about two minutes and was not self-limiting. an echocardiogram was ordered and performed, however, readings were not confirmed. the patient survived the event. |  |  |  |  | B5 additional information received. h6 updated health effect - clinical code to reflect new information. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
18833 | 1220648-2024-18876 | 20351816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351816&pc=OZD | 09/08/2024 | 11/19/2024 |  | 10/02/2024 |  |  | 09/13/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025509313 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and there was a high purge pressure issue. the purge pressure was >1000 mmhg and purge flow <1 milliters per hour. the staff checked for any kinks, and none were found. the purge pressure was resolved with lysis treatment. the weaning of the impella 5.5 was successful and there were no reported complications. the patient survived the explant. | [5 empty] | A.5 is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. there was no product or data logs returned. in the clinical details, it was stated that lysis resolved the high purge pressure. lysis is a therapy that uses a medication to deliver medication to blood clots. based on the clinical details, the root cause of the high purge pressure is most likely due to biomaterial. b.1 and h.1 revised as intervention was required which makes this issue an adverse event/serious injury with a product problem and was incorrect on initial manufacturer device report number: 1220648-2024-18876.
18834 | 1220648-2024-18878 | 20351869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20351869&pc=OZD | 04/06/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/06/2022 | 10/02/2024 |  |  | 11/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022153283 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, device had unresolvable suction alarms. the pump was removed and replaced with no harm done to the patient. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs showed ventricular placement signal waveforms prior to pump activation. upon activation, "suction" alarms occurred along with the "impella position in ventricle" alarm. the positioning alarm resolved, however the "suction" continued with low flows and low motor current pulsatility. visual inspection of the returned pump revealed biomaterial on and around the impeller. no other damage or abnormalities were found. therefore, the root cause of the low pump flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
18835 | 1220648-2024-18879 | 20352012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352012&pc=OZD | 04/06/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/19/2022 | 10/02/2024 |  |  | 12/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165785 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, when the extracorporeal membrane oxygenation (ecmo) was removed, an access site bleed at the impella site was noted. this prompted the impella pump explant and the patient remained in stable condition. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the access site bleeding was patient movement. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18836 | 1220648-2024-18880 | 20352053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352053&pc=OZD | 09/13/2024 | 06/26/2025 |  | 10/02/2024 |  |  | 09/13/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474803 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Arrhythmia | Tachycardia | Ischemia Stroke | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Us complainant reported the patient was on impella 5.5 support and they complained of having blurry vision. a computed tomography scan showed an ischemic stroke. support continued. furthermore, the patient was very tachycardic and developed arrhythmias. the echocardiogram showed that the impella was near the septum and was thought to be causing the arrhythmias. the pump was repositioned three times. additionally, the pump was having frequent suction alarms and the automated impella controller (aic) showed placement signal not reliable alarms. the patient was very dependent on the impella support, so the staff chose to replace the pump. the patient survived the event. | [5 empty] | The investigation for the optical sensor issue, stroke, low pump flow, and arrhythmias has been completed. the device was not returned for evaluation. data logs were reviewed which showed that a placement signal not reliable (psnr) alarm is triggered when placement signal suddenly increases to 3000. before psnr alarm, placement signal was noted to step down ~50mmhg for approximately 3 hours before drifting down then spiking to 3000. clinical details noted that patient was experiencing suction alarms and psnr alarm was triggered on day 22 of support. the root cause of the optical signal issues was most likely due to sensor wear. data logs also showed low pump flow and suction alarms were triggered at that time. pump flows at time of alarms were all within normal ranges for p-levels. clinical details noted that patient was 100% pump dependent while on support. the root cause of the low pump flow was most likely patient condition related. the root cause of the arrhythmia and stroke could not be determined without additional clinical details. | G1 reporting contact email revised on manufacturer device report 1220648-2024-18880 in accordance with updated procedures. h6 investigation conclusions code 50 was erroneously entered on the initial manufacturer device report number 1220648-2024-18880-1. h6 investigation findings code 140 and investigation conclusions code 4307, 4310, and 4315 were inadvertently omitted on the initial manufacturer device report number 1220648-2024-18880-1. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
18837 | 1220648-2024-18881 | 20352093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352093&pc=OZD | 04/08/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/08/2022 | 10/02/2024 |  |  | 02/29/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182964 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped. the pump was removed and replaced without any harm to the patient. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs showed the occurrence of the "impella stopped - reverse flow" alarm with the pump failing to start. all pump performance metrics were at 0. no external damage or abnormalities were observed on the returned pump. electrical resistance testing revealed that the motor phases were all within specification. the pump was connected to the aic, and it successfully started. therefore, the root cause of the pump failing to start was not determined since the issue could not be reproduced. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18838 | 1220648-2024-18882 | 20352215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352215&pc=OZD | 09/08/2024 | 11/21/2024 |  | 10/02/2024 |  |  | 09/08/2024 | 10/02/2024 |  |  | 04/30/2025 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346053 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella cp implant and it was aborted due to patient anatomy. the impella cp failed to pass the aortic arch. the staff tried to use a more supportive guide, but the implant was not successful. no patient harm has been reported. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy.
18839 | 1220648-2024-18884 | 20352262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352262&pc=OZD | 04/12/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/12/2022 | 10/02/2024 |  |  | 11/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023160817 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient had impella cp support for over 14 successful days when it was explanted. at the explant a biomaterial/clot was observed. it was noted that there was no effect on the lower limb hemorrhage or brain. no harm or injury noted from the thrombosis. | [5 empty] | The investigation for the reported thrombosis issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the issue was related to the patient¿s condition. this report is being filed as part of a retrospective review of historical records.
18840 | 1220648-2024-18883 | 20352330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352330&pc=PYX | 09/14/2024 | 10/31/2024 |  | 10/02/2024 |  |  | 09/14/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529738 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 53-year-old male patient with right heart failure was implanted with an impella rp flex device for mechanical circulatory support. while on support, the patient developed hemolysis. additionally, flows decreased when started pulling on continuous renal replacement therapy. the pull was decreased to see if flows would improve. flows did not improve. bleeding was noted. the source of the bleed is unclear. three units of packed red blood cells were given. the pump was ultimately explanted and the patient procedure outcome was noted as survived. | [5 empty] | The investigation of low pump flow, vascular damage peripheral vessel, and hemolysis has been completed. visually inspected and no physical damage or abnormalities seen. borescope view of cannula obtained and biomaterial observed at inflow and wrapped around impeller. no issues seen with pump sensor and electrical resistance signals. visually inspected and no physical damage or abnormalities seen. borescope view of cannula obtained and biomaterial observed at inflow and wrapped around impeller. no issues seen with pump sensor and electrical resistance signals. the root cause of vascular damage - peripheral vessel could not be determined. the cause of the hemolysis issue was likely patient condition related. the cause of the low pump flow issue was biomaterial. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
18841 | 1220648-2024-18886 | 20352446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352446&pc=OZD | 09/08/2024 | 11/13/2024 |  | 10/02/2024 |  |  | 09/08/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447912 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support a patient with a cardiac and systemic history and comorbidities. the pump was placed and the limbs were bilaterally ischemic. additionally, there were signs of hemolysis. the team explanted and escalated support to an impella 5.5 pump. | [5 empty] | The investigation of limb ischemia and hemolysis has been completed. complaint device was not returned for analysis. data logs were reviewed and no motor current spike or suction alarms were observed. no positioning issue alarms were observed. no issue was observed that indicated to the reported hemolysis issue. the cause of the hemolysis issue could not be determined since the complaint device was not returned for investigation and insufficient clinical data was provided. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
18842 | 1220648-2024-18885 | 20352518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352518&pc=OZD | 09/13/2024 | 11/15/2024 |  | 10/02/2024 |  |  | 09/13/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489047 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 51-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient had access site bleeding. manual pressure was held and new sutures were placed. the bleeding subsided and support continued. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation the root cause of the access site bleeding issue was most likely use related, manual pressure and new sutures placed to achieve hemostasis as per the clinical description provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18843 | 1220648-2024-18889 | 20352641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352641&pc=OZD | 09/08/2024 | 11/14/2024 |  | 10/02/2024 |  |  | 09/08/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025530564 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was to be implanted with an impella 5.5 implant but there was difficulty advancing the pump. the staff used excessive force and the pump was removed. upon removal, a kink was noted behind the motor housing. a new impella was attempted. | [5 empty] | A.5 ethnicity is unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the root cause of the delivery issue was most likely use issue since the pump was noted with excessive force noted in attempts to advance into place noting kink behind the motor housing.
18844 | 1220648-2024-18887 | 20352694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352694&pc=OZD | 09/13/2024 | 11/15/2024 |  | 10/02/2024 |  |  | 09/13/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025518051 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support for 66-year-old male patient, the impella cp was not reaching adequate levels of flow. the pump was therefore explanted and a new pump was used. there was no harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The investigation of low pump flow has been completed. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review.
18845 | 1220648-2024-18893 | 20352734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352734&pc=OZD | 09/15/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/15/2021 | 10/02/2024 |  |  | 06/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022102841 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, an uncontrollable bleed was observed after the peel away sheath was removed. the impella pump was removed, and manual pressure was held for a new impella pump to be placed. the procedural outcome is the process was aborted and a new device was utilized. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding was use related as the introducer sheath was not peeled away outside the body. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18846 | 1220648-2024-18888 | 20352771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352771&pc=OZD | 04/18/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/18/2022 | 10/02/2024 |  |  | 12/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177717 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, suction alarms that were resolved by removing, rinsing the device, and reinserting. there was no harm to the patient and suction resolved. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that biomaterial was found in the outflow. the data logs confirm the low flow and suction that was resolved by removing the pump, running it in a bucket and reinserting it. the most likely root cause of the low flow was ingested biomaterial that was dislodged upon explant or while running the pump in a bucket allowing it to reach normal flows. this report is being filed as part of a retrospective review of historical records.
18847 | 1220648-2024-18897 | 20352847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352847&pc=OZD | 07/25/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 07/25/2022 | 10/02/2024 |  |  | 11/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161971 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | American Indian Or Alaskan Native | Malfunction | No | Yes | Fluid/Blood Leak | Fracture | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a purge leak and sidearm issues/crack. the pump was replaced. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the issue was not determined since product was not returned. as noted in the impella ifu: impella cp with smart assist system section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18848 | 1220648-2024-18890 | 20352860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352860&pc=OZD | 04/19/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/19/2022 | 10/02/2024 |  |  | 01/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023175021 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. while prepping to insert, a purge cassette was generating no pressure during preparation of the device. it is unknown how the priming issue was resolved. no patient harm was reported. | [5 empty] | The investigation for the reported priming issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18849 | 1220648-2024-18895 | 20352891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352891&pc=OZD | 09/09/2024 | 11/21/2024 |  | 10/02/2024 |  |  | 09/09/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025537299 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. according to the doctor, impella bleed heavily between the white and blue repositioning sheaths. the pump was explanted. the patient¿s outcome is unknown. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | A.5 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18850 | 1220648-2024-18896 | 20352892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352892&pc=OZD | 04/20/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/20/2022 | 10/02/2024 |  |  | 12/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023164391 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. after implantation and return the wire, the pump would not start. the cp was replaced and there was no harm or injury reported. | [5 empty] | The investigation for the reported pump did not start issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that only the lt log was returned which showed the pump ran for 3 minutes. no abnormalities were found on the returned pump, and the pump was able to start in the lab. thus, the cause of the pump stop was not determined as it could not be reproduced. this report is being filed as part of a retrospective review of historical records.
18851 | 1220648-2024-18891 | 20352903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352903&pc=OZD | 09/13/2024 | 11/20/2024 |  | 10/02/2024 |  |  | 09/13/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529693 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella cp placed for a 51-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs). the pump was placed but after 28 hours the pump was explanted due to low pump flows. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into low pump flow has been completed. the impella device was not returned for evaluation. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review.
18852 | 1220648-2024-18892 | 20352906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20352906&pc=OZD | 09/14/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/14/2021 | 10/02/2024 |  |  | 06/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022097544 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped suddenly with a high purge pressure alarm. numerous efforts were made to restart the device however, all were unsuccessful. it was also noted that the purge cassette leaked during this period. the impella was withdrawn into the descending aorta and the device was removed and replaced with a new pump. the procedural outcome is the patient survived. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the purge system blocked alarm was reproduced in the lab, with pp at 1063 mmhg/pf at 1.5 mmhg. the pump would not start in the lab. blood was found in the purge lumen. potential evidence of a slight kink was found on the catheter. data log evaluation indicates a sudden increase in purge pressure, followed by a sharp rise in mc and then a pump stop. thus, the root cause of the pump stop is most likely due to blood ingress as a result of high purge pressure caused by a kink. the purge cassette was not returned and thus the root cause could not be determined. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18853 | 1220648-2024-18894 | 20353030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353030&pc=OZD | 09/07/2024 | 11/14/2024 |  | 10/02/2024 |  |  | 09/07/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521720 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was escalated to impella 5.5 due to hemolysis. the patient survived the event. | [5 empty] | A.5 race is unknown. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. | The investigation for the reported hemolysis has been completed. the case of the hemolysis issue could not be determined since the complaint device was not returned for analysis and insufficient clinical data was provided. data logs were unavailable for review.
18854 | 1220648-2024-18900 | 20353112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353112&pc=OZD | 09/09/2024 | 11/18/2024 |  | 10/02/2024 |  |  | 09/09/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533736 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that soon after impella cp placement, a large hematoma was noted in the right groin. blood products were given and femostop was placed. bleeding continued. the patient returned to the catheterization lab for impella cp removal and covered stent. at beginning of the case, the right femoral arteriotomy was pre-closed with two percloses, one of which failed as the knot was externalized. it was also very difficult to deploy and got stuck trying to remove the needles and the device that deploys the needles. upon removal of the impella cp, covered stent was placed over the right femoral arteriotomy site along with balloon tamponade and manual pressure. the patient survived the event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed and vascular damage has been completed. impella cp, dilators, guidewire, and 14 french x 23 centimeter introducer were returned for evaluation. the 14 french x 13 centimeter introducer was noted to be used to insert the impella and was not returned for evaluation. no damage to the repositioning sheath was noted on the returned product. clinical details noted that the perclose suture failed as the knot was externalized and there was difficulty in deploying sutures and removing the suture device. the patient had received multiple blood thinners while on support and noted bleeding from multiple access sites while in the cath lab. a large hematoma and oozing was noted at the access site. additionally, arterial blood was present just inferior of the sheath site on angiogram during removal. the root cause of the vascular damage was most likely due to a use issue pertaining to the perclose sutures and was not related to the impella pump. the root cause of the access site bleeding - major was most likely patient condition related as it was noted that patient was bleeding from multiple access sites while on support.
18855 | 1220648-2024-18658 | 20353153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353153&pc=PYX | 01/24/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 01/24/2022 | 10/02/2024 |  |  | 12/31/2022 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2023160855 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. after initial successful insertion and ramp up, the device moved back across the valve due to issue with scrub holding manual pressure and pushing on catheter during the reposition sheath insertion. multiple explant and insertions were attempted, and the physician decided to replace the device with another impella rp. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, the details provided were sufficient to determine a root cause. the root cause of the positioning issues was determined to be use related as the rp was accidently pulled back across the valve. as noted in the impella ifu: impella cp with smart assist system impella rp smart assist system with automated impella section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
18856 | 1220648-2024-18902 | 20353231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353231&pc=OZD | 09/15/2024 | 11/14/2024 |  | 10/02/2024 |  |  | 09/15/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527273 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support, and after the implant, they were unable to doppler pulses to the right foot. it was found that the repositioning sheath was occlusive. patient had less than 1 l/min of native output, so decision was made to perform external femoral-femoral bypass. additionally, the patient had suspected hemolysis on support. the patient's urine was pink. the pump's position was verified. support continued. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of hemolysis and limb ischemia has been completed. the root cause of hemolysis was most likely due to patient condition. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. h6 code 2558, 4641, and 4625 was reported incorrectly on manufacturer device report 1220648-2024-18902.
18857 | 1220648-2024-18898 | 20353279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353279&pc=OZD | 09/14/2024 | 11/15/2024 |  | 10/02/2024 |  |  | 09/14/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539570 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella 5.5 would not advance into subclavian artery due to anatomy. the vessel was borderline in size. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, 5.5 would not advance into subclavian artery due to anatomy as per the clinical description provided.
18858 | 1220648-2024-18901 | 20353302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353302&pc=OZD | 07/19/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 07/19/2022 | 10/02/2024 |  |  | 03/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210529 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during insertion, the physician was unable to advance the impella into the abdominal aorta. the impella implant was aborted and replaced with an intra-aortic balloon pump, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was patient condition. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18859 | 1220648-2024-18904 | 20353311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353311&pc=OZD | 09/13/2024 | 11/13/2024 |  | 10/02/2024 |  |  | 09/13/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539079 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 49-year-old male patient with caridomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, hematoma was noted. hematoma resolved by washout. additionally, the patient had a run of supraventricular tachycardia (svt). lidocaine was given for anti-arrhythmia management. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
18860 | 1220648-2024-18903 | 20353332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353332&pc=OZD | 04/20/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/20/2022 | 10/02/2024 |  |  | 01/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189357 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the device cannula kinked after being placed in the left ventricle. a new impella was placed without further issue or harm to the patient. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that evidence of a cannula kink was confirmed. the root cause of the delivery issues was most likely the patient¿s condition. this report is being filed as part of a retrospective review of historical records.
18861 | 1220648-2024-18908 | 20353349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353349&pc=OZD | 07/26/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 07/29/2022 | 10/02/2024 |  |  | 01/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183298 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient developed hemolysis. no intervention was specified, and the patient remained on impella support. it was reported that the patient survived. | [5 empty] | The investigation for the reported hemolysis has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the hemolysis could not be determined. as noted in the impella ifu: hemolysis: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
18862 | 1220648-2024-18899 | 20353356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353356&pc=OZD | 07/01/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 07/01/2022 | 10/02/2024 |  |  | 12/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170752 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp pump did not start. a portion of the easy guide lumen was separating and becoming lodged within the impella cp blades. the pump was replaced and there was no further issue. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. visual inspection of the returned product revealed the presence of the red lumen in the pump. the lumen was fractured at the outflow and twisted around the impeller. after reviewing the clinical information, it was determined that the cause of the pump not starting was a red lumen removal caused by user technique. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
18863 | 1220648-2024-18909 | 20353364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353364&pc=OZD | 07/26/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 08/03/2022 | 10/02/2024 |  |  | 03/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207557 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. however, during implantation of the device, the physician had difficulty removing the easy guide lumen. the physician elected to discard the pump, and another device was used for the procedure. no patient harm was reported in relation to this issue. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely related to removal of the red lumen. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18864 | 1220648-2024-18907 | 20353397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353397&pc=OZD | 07/23/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 07/23/2022 | 10/02/2024 |  |  | 12/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171279 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella began to stop. after being restarted multiple times, the pump was removed and replaced. there was no reported harm to the patient. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause could not be determined. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18865 | 1220648-2024-18906 | 20353442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353442&pc=OZD | 09/09/2024 | 11/18/2024 |  | 10/02/2024 |  |  | 09/09/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525317 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was inserted per protocol via right femoral artery using short sheath. single access was attempted with peel away sheath but failed due to continuous bleeding at the hemostatic valve. single access was aborted and an alternate site was obtained for percutaneous coronary intervention in left femoral artery. hemostatic valve bleeding ceased with single access removed. the patient survived the event. | [5 empty] | D.9 updated as the pump and introducer were returned for investigation. the investigation into introducer damage - mechanical has been completed. the returned introducer was investigated and there was not any clear evidence of damage to the introducer. the returned introducer was put through a leak reproduction test by plugging the end of the introducer with a dilator and connecting the flushing lumen to a syringe and pressure gage by means of a three-way lure. the flushing lumen had to be cut down because there was so much dried blood/biomaterial that would not allow fluid to get into the introducer. the colored water in the syringe was slowly pushed into the introducer, and then the lure release was switched to the pressure gage to see how much pressure had been applied to the introducer. this showed that the introducer was under at least 264.4 mmhg of pressure, and a leak did not reproduce. therefore, the returned introducer met specification. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
18866 | 1220648-2024-18905 | 20353492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353492&pc=OZD | 07/22/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 07/22/2022 | 10/02/2024 |  |  | 03/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202253 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during delivery, the pump was not able to be fully inserted into the left ventricle due to patient anatomy, even when using a new wire. impella insertion was aborted, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was patient anatomy, related to patient condition. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18867 | 1220648-2024-18910 | 20353575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353575&pc=OZD | 04/20/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/20/2022 | 10/02/2024 |  |  | 08/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022129510 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a loss of optical signal after the placement in the left ventricle. the impella device was replaced and a new impella was implanted and supported successfully. no patient harm was reported. | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the red handle was opened the optical fiber was found to be not fully seated in the ferrule. the root cause of the optical failure was determined to be a damaged optical fiber due to an external manufacturing process issue. this report is being filed as part of a retrospective review of historical records.
18868 | 1220648-2024-18911 | 20353679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353679&pc=PYX | 09/16/2024 | 11/18/2024 |  | 10/02/2024 |  |  | 09/16/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025517118 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Hematuria | Cardiovascular Dysfunction/Insufficiency | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella rp had suction and acute drop in impella flow with increase in central venous pressure (cvp) and pulmonary artery pressures. labs were hemolyzing, platelet drop, and patient with acute hematuria. the impella position was confirmed. albumin was given for volume resuscitation with no improvement. systemic heparin started overnight with partial thromboplastin time (ptt) 55, and bicarbonate purge running. one unit of packed red blood cells was transfused. the impella rp was removed. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the low pump flow and the hemolysis has been completed. impella rp flex was returned for evaluation. pump was returned cut at catheter. upon visual inspection, no biomaterial, damage to impeller blades, or cannula kinks were noted. pump was not able to be hemolysis tested as the pump was returned cut. data logs were reviewed and a motor current spike a the beginning of the case with a simultaneous speed deviation and step down in flows was noted. placement signal also stepped up at time of motor current spike. lt logs of entirety of case show pump was never able to achieve proper flows per p-level from start-up of pump. additionally, later in case at time of suction alarms, motor current becomes variable with spiking in motor speed and pump flow. cvp trends negatively throughout case with a noisy placement signal. clinical details noted that patient was experiencing hemolysis indicated from labs with suction alarms triggering. pump position confirmed to be in proper position. volume given with no improvement. the root cause of the low pump flow was most likely due to biomaterial ingestion at the beginning of the case. the root cause of the hemolysis was most likely patient condition related.
18869 | 1220648-2024-18913 | 20353767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353767&pc=OZD | 04/20/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/20/2022 | 10/02/2024 |  |  | 07/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114098 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the pump had signal loss which prompted the team to explant at the end of the surgical procedure. no patient harm was reported. | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the data logs showed no evidence of optical sensor failure with no major placement signal alarms/events occurring. therefore, it was determined that no optical signal issue occurred during the case. this report is being filed as part of a retrospective review of historical records.
18870 | 1220648-2024-18912 | 20353912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353912&pc=OZD | 09/16/2024 | 11/25/2024 |  | 10/02/2024 |  |  | 09/16/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025536800 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and during the implant, the doctor was suturing the graft, and it began to bleed at the anastomosis site. graft was removed and resutured. impella pump was placed without issue. during closure, the graft continued to ooze resulting in the surgical incision being packed and left open with a pressure dressing. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report: g1-mdr reporting contact email.
18871 | 1220648-2024-18915 | 20353986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20353986&pc=OZD | 09/09/2024 | 11/14/2024 |  | 10/02/2024 |  |  | 09/09/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025512711 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Limb Pain | [12 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump delivered to support a 50-year-old female patient admitted in with a in st elevated myocardial infarction. the pump had supported for just under 21 days when it was explanted. the explant was followed by limb injury. the left arm had numbness and pain after explant. the limb later was treated by thrombectomy. the patient survived the explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation for the reported peripheral nerve injury has been completed. no product was returned for analysis. the root cause of the peripheral nerve injury was not determined due to insufficient clinical details.
18872 | 1220648-2024-18914 | 20354470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20354470&pc=OZD | 09/16/2024 | 11/18/2024 |  | 10/02/2024 |  |  | 09/16/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025512877 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and the pump twisted into the mitral valve, causing some premature ventricular contractions. staff attempted to reposition away from the mitral, but it was still going toward the mitral valve. the doctor decided to leave the position as is and monitor. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b1 should have been both adverse event and product problem on manufacturer device report 1220648-2024-18914.
18873 | 1220648-2024-18919 | 20354481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20354481&pc=OZD | 07/28/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 07/28/2022 | 10/02/2024 |  |  | 10/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2022148312 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. however, implantation was not possible due to aortic valve stenosis. implantation was aborted, and no further intervention was specified. there was no reported harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was patient anatomy/condition. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18874 | 1220648-2024-18918 | 20354673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20354673&pc=OZD | 04/20/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/20/2022 | 10/02/2024 |  |  | 12/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171809 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were suction alarms during. the pump was ultimately removed, and the access site was closed. impella was successfully weaned. no patient harm was reported. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the data logs confirm the sudden low flow. the root cause of the low pump flow was most likely ingested biomaterial due to the clinical account. this report is being filed as part of a retrospective review of historical records.
18875 | 1220648-2024-18923 | 20354764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20354764&pc=OZD | 04/20/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/20/2022 | 10/02/2024 |  |  | 01/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189138 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the device had suction alarms. echo showed good position in left ventricle. the device was explanted, and the decision was made not to replace the impella. there was no harm or injury to the patient. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18876 | 1220648-2024-18916 | 20354779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20354779&pc=OZD | 09/16/2024 | 11/18/2024 |  | 10/02/2024 |  |  | 09/16/2024 | 10/02/2024 |  |  |  |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025536807 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 fail to deliver for a patient admitted in and needing care escalation from the cp. the native anatomy prevented delivery. no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy.
18877 | 1220648-2024-18917 | 20354817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20354817&pc=OZD | 07/28/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 07/28/2022 | 10/02/2024 |  |  | 04/30/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217198 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. during insertion, the pigtail crossed the aortic valve and the wire was placed. however, severe aortic tortuosity prevented the impella pump from crossing the aortic valve, and there was no more available catheter to push the impella further. the physician elected to abort the procedure, and no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. per clinical account, the root cause of the delivery issues is patient anatomy/condition. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
18878 | 1220648-2024-18920 | 20354839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20354839&pc=OZD | 08/03/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 08/03/2022 | 10/02/2024 |  |  | 10/31/2022 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021010887 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, flows were low and suction alarms occurred. the impella cp device was explanted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was not received for evaluation. the root cause of the low pump flow was unable to be determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
18879 | 1220648-2024-18921 | 20354852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20354852&pc=OZD | 08/04/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 08/04/2022 | 10/02/2024 |  |  | 04/30/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224857 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella access site. the impella was explanted and replaced with a new impella cp to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of access site bleeding. the device was returned for evaluation. the root cause of the access site bleeding was a lack of vessel recoil onto the sheath. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18880 | 1220648-2024-18925 | 20354884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20354884&pc=OZD | 08/07/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 08/07/2022 | 10/02/2024 |  |  | 04/30/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224334 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the pump had multiple high motor current and impella stopped alarms. the pump was able to be restarted a few times before eventually the final pump stop occurred. the physician decided to explant the device, it was reported that the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the product was returned, and the pump stop was confirmed. the root cause of the pump stop was bearing wear. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18881 | 1220648-2024-18928 | 20354894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20354894&pc=OZD | 04/09/2022 | 10/09/2025 |  | 10/02/2024 |  |  | 04/09/2022 | 10/02/2024 |  |  | 12/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023175066 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was bleeding which was treated by holding heparin. additionally, the purge side arm was cracked during a purge cassette change, therefore, the pump was exchanged. | The us complainant had a cp support ongoing when the team broke the purge sidearm. the sidearm cracked and prompted the team to replace the pump. additionally, it was reported that there was bleeding and this was treated by holding heparin. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the purge sidearm was found to be broken into two pieces between the pressure reservoir and the yellow luer check valve. considering the nature of the damage, it was determined that the cause of the sidearm bond damage was the use of excessive force during the purge cassette exchange.
18882 | 1220648-2024-18926 | 20354911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20354911&pc=OZD | 08/08/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 08/08/2022 | 10/02/2024 |  |  | 04/30/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223280 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, due to lack of adequate purge there was an impella pump stop that was not able to be fixed. the patient continued on support until the device was successfully weaned. it was reported that the patient survived the reported. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the pump stop was most likely blood ingress due to a lack of purge. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18883 | 1220648-2024-18933 | 20355027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355027&pc=OZD | 08/12/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 08/12/2022 | 10/02/2024 |  |  | 08/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022126967 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. the sterile sleeve was not able to retract or move forward. several attempts were made to retract it but were unsuccessful. the impella was explanted and replaced with a new impella cp to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the delivery issues was not determined since product was not returned. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
18884 | 1220648-2024-18922 | 20355068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355068&pc=OZD | 08/05/2022 | 02/28/2025 |  | 10/02/2024 |  |  | 08/05/2022 | 10/02/2024 |  |  | 04/30/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225080 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. it was reported that the device failed to prime. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported priming issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that product was visually inspected, and no abnormalities were observed. the product was then connected to aic and then de-aired. after de-air, the pump was connected to purge cassette and fluid was observed exiting the impeller purge gap. the root cause of the purge issue was use related. this report is being filed as part of a retrospective review of historical records.
18885 | 1220648-2024-18924 | 20355081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355081&pc=OZD | 08/05/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 08/05/2022 | 10/02/2024 |  |  | 04/30/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218745 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. the impella cp was unsuccessfully delivered over the 0.018¿ abbott steel core multiple times and couldn't cross the valve. the impella was explanted and replaced with a new impella cp to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the product was returned, and the delivery issue was confirmed. the returned pump was visually inspected and no abnormalities were found. the abbott 0.018" guidewire that was utilized during this case is not approved for use with impella per ifu 0048-9007. therefore, the cause of the delivery issue was use related. instructions for use for the related event are as follows: impella cp with smartassist system section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
18886 | 1220648-2024-18932 | 20355103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355103&pc=OZD | 08/12/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 08/12/2022 | 10/02/2024 |  |  | 04/30/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213989 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. on day 9 of support, the physician observed a pump stop. the impella cp was explanted and replaced to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the device was not received for evaluation. the root cause of the pump stop was unable to be determined as no product was returned for analysis. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18887 | 1220648-2024-18930 | 20355127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355127&pc=OZD | 08/10/2022 | 02/21/2025 |  | 10/02/2024 |  |  | 08/10/2022 | 10/02/2024 |  |  | 04/30/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216618 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, it was reported that there was a leaking purge side arm. the impella was explanted and replaced with a new impella cp. the patient survived the procedure. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | The returned pump was visually inspected and a crack in the yellow check valve female luer was observed. a syringe with colored was connected to the yellow luer and the purge leak was reproduced. the cause of the purge leak was sidearm yellow luer damage on day 8 was due to bicarb.
18888 | 1220648-2024-18934 | 20355177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355177&pc=OZD | 08/12/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 08/25/2022 | 10/02/2024 |  |  | 03/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211267 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention attempted to being implanted with an impella cp device for mechanical circulatory support. during implantation, a heavily calcified iliac was noted that made the placement of impella cp difficult. the impella cp was removed, and the patient was implanted with another impella for patient support. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was determined to be diseased patient vessels. patient had peripheral artery disease/peripheral vascular disease calcified vessels. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
18889 | 1220648-2024-18940 | 20355196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355196&pc=OZD | 05/06/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 05/06/2023 | 10/02/2024 |  |  | 12/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318155 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in was implanted with an impella cp device for mechanical circulatory support. during support, the pump stopped. the patient tested positive for heparin-induced thrombocytopenia while on support and was receiving bivalirudin at a rate of (b)(4). sodium bicarbonate was included in the purge solution. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18890 | 1220648-2024-18936 | 20355345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355345&pc=OZD | 07/03/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 07/03/2022 | 10/02/2024 |  |  | 07/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108786 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a priming issue with the impella cp. the priming did not proceed due to an error in the purge fluid. although the staff replaced the purge set, it did not improve. there was no outflow of purge fluid from the tip of impella. the patient was reportedly stable. | [5 empty] | The investigation for the reported priming issue has been completed. the returned pump was visually inspected. no kink was observed in cannula or catheter. there were also no biomaterials or broken impella blades. pump and purge cassette were connected to aic for testing. purge testing reproduction was performed. the pump purged normally. no issue was found. the root cause was technique/use. this report is being filed as part of a retrospective review of historical records.
18891 | 1220648-2024-18927 | 20355383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355383&pc=OZD | 08/09/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 08/09/2022 | 10/02/2024 |  |  | 03/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200400 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella cp device for mechanical circulatory support. it was observed that the red lumen was slightly twisted leading to resistance being felt when backloading the guidewire. the impella was explanted and replaced with a new impella cp to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was determined to be a red lumen issue. this is most likely a use related issue. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
18892 | 1220648-2024-18929 | 20355396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355396&pc=OZD | 04/24/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/24/2022 | 10/02/2024 |  |  | 01/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189140 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the device had lower than expected flows. the pump was removed, and a large clot was observed in the cannula. a new impella was implanted with no harm done to the patient. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was most likely biomaterial. this report is being filed as part of a retrospective review of historical records.
18893 | 1220648-2024-18931 | 20355403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355403&pc=OZD | 08/10/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 08/10/2022 | 10/02/2024 |  |  | 02/29/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198171 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after the pump was inserted, it was unable to start. multiple attempts were made to restart the pump but were unsuccessful. the impella cp was explanted and replaced to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump did not start. the product was returned. the root cause of the pump did not start was determined to be use related as the red lumen was not completely removed during case start. this report is being filed as part of a retrospective review of historical records.
18894 | 1220648-2024-18939 | 20355419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355419&pc=OZD | 04/27/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/27/2022 | 10/02/2024 |  |  | 02/29/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197028 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the pump fell out of the ventricle and would not recross while attempting to reposition. a new impella was used without any harm to the patient. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the root cause of the positioning issues was determined to be use related as the pump was accidentally pulled out of the ventricle. this report is being filed as part of a retrospective review of historical records.
18895 | 1220648-2024-18935 | 20355436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355436&pc=OZD | 04/24/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/24/2022 | 10/02/2024 |  |  | 12/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023179443 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped shortly after device insertion. the pump was removed and replaced with no harm to the patient. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the data logs show multiple motor current spikes followed by a high motor current pump stop. this is indicative of biomaterial ingestion. no product was returned for analysis. the root cause of the pump stop was most likely ingested biomaterial. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18896 | 1220648-2024-18941 | 20355466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355466&pc=OZD | 05/21/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 05/21/2023 | 10/02/2024 |  |  | 11/30/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303583 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella suddenly stopped. the team had thoughts that the pump had ingested clot or another biomaterial. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18897 | 1220648-2024-18942 | 20355469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355469&pc=OZD | 06/19/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 06/19/2023 | 10/02/2024 |  |  | 12/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024302412 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to being implanted with an impella cp device for mechanical circulatory support. the complainant reported a failure to deliver the device due to patient's anatomy. as a result of this complication the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation into the delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the most likely cause is due to the patient¿s condition related to small vessel. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18898 | 1220648-2024-18943 | 20355492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355492&pc=OZD | 06/20/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 06/20/2023 | 10/02/2024 |  |  | 01/31/2025 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024332152 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the impella device failed to cross the aortic valve following balloon aortic valvuloplasty. additionally, there was a kink in the catheter above the motor housing. the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation into the difficulty inserting/removing introducer has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the most likely cause is due to the patient's aortic valve stenosis. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18899 | 1220648-2024-18945 | 20355510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355510&pc=OZD | 04/27/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/27/2022 | 10/02/2024 |  |  | 01/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192420 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp pump cannula kinked due to the patient's anatomy and a small left ventricle/aorta. the procedure was aborted and the pump was explanted. the patient was supported by the intra-aortic balloon pump (iabp). no patient harm was reported. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the root cause of the low pump flow was determined to be a cannula kink caused by patient anatomy. this report is being filed as part of a retrospective review of historical records.
18900 | 1220648-2024-18937 | 20355563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355563&pc=OZD | 04/26/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/26/2022 | 10/02/2024 |  |  | 01/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188397 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella kinked and prolapsed in the left ventricle. the kinked pump was pulled back to the iliac. the pump at that time was found to be stuck within the repo sheath. this prompted the pump to be explant and replaced. no patient harm was reported. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the root cause of the positioning issue was determined to be use related as the pump was improperly positioned causing a kink and in attempting to pull pack and reposition the pump was pulled too far. this report is being filed as part of a retrospective review of historical records.
18901 | 1220648-2024-18947 | 20355612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355612&pc=OZD | 07/05/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 07/05/2022 | 10/02/2024 |  |  | 11/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167843 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. it was reported that there was an impella cp pump stop. there was an abrupt motor current spike and pump stop, restart and stop attempted a few times, then finally successful. the pump was running and attempted a deceleration to extend pump life. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the issue was not determined since product was not returned. the impella cp was used beyond the design life. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18902 | 1220648-2024-18944 | 20355649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355649&pc=OZD | 06/16/2023 | 09/16/2025 |  | 10/02/2024 |  |  | 06/16/2023 | 10/02/2024 |  |  | 01/31/2025 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331100 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. prior to insertion, a low purge pressure alarm was noted. the impella displayed a purge failure, and the pump was discarded. no further information available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the priming issue was not determined since product was not returned and sufficient clinical information was not provided. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2024-18944. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-18944 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-18944 in accordance with updated procedures.
18903 | 1220648-2024-18952 | 20355739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355739&pc=OZD | 09/06/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 09/06/2022 | 10/02/2024 |  |  | 02/29/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201363 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the physician implanted the impella cp and felt that the cannula was kinked, removed it, and implanted another pump. successful change to a new pump. no harm done to the patient reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. no product was returned for investigation. the cause of the delivery issue was most likely patient condition related. the kinked cannula was likely caused by patient condition. this report is being filed as part of a retrospective review of historical records. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
18904 | 1220648-2024-18949 | 20355740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355740&pc=OZD | 09/06/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 09/06/2022 | 10/02/2024 |  |  | 05/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231276 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a pump stop during use of the device. attempts to restart were not successful. the pump was removed. there was no harm to the patient reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the data logs confirm the sudden high motor current pump stop on the 7th day of cp support. the root cause of the pump stop was unable to be determined as no product was returned for analysis. as noted in the impella ifu: impella cp with smart assist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18905 | 1220648-2024-18951 | 20355763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355763&pc=OZD | 04/28/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/28/2022 | 10/02/2024 |  |  | 11/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162675 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a spike in motor current followed by a pump stop. the pump was removed and replaced with no harm done to the patient. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the pump stop was use related since improper technique was used to remove the introducer causing pump malposition and a catheter kink. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18906 | 1220648-2024-18948 | 20355776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355776&pc=OZD | 04/27/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/28/2022 | 10/02/2024 |  |  | 03/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191405 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump suddenly stopped and the impella could no longer be started. no additional details were provided regarding the treatment or management of this condition. | [5 empty] | D6 explanted unknown. the investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs showed the pump stopped immediately upon the sudden occurrence of the "impella stopped - motor current high" alarm. prior to the pump stop, the motor current had low pulsatility and there the pump flow waveform was clipped. a large impeller purge gap was observed on the returned pump. the high motor current pump stop was reproduced when attempting to run the pump. therefore, it was determined the root cause of the pump stop was bearing wear. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18907 | 1220648-2024-18955 | 20355919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355919&pc=OZD | 06/19/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 06/19/2023 | 10/02/2024 |  |  | 01/31/2025 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024334045 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon insertion the pump failed to start. the procedure was completed using a new device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump did not start. the product was not returned. the root cause of the pump did not start was determined to be use related since they attempted to start outside the body. this report is being filed as part of a retrospective review of historical records.
18908 | 1220648-2024-18957 | 20355955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355955&pc=OZD | 06/22/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 06/22/2023 | 10/02/2024 |  |  | 08/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267595 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to being implanted with an impella cp device for mechanical circulatory support. the physician was unable to implant the impella due do the patients vessels being too narrow. the impella was explanted and the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of difficulty inserting/removing introducer. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the difficulty inserting/removing introducer was due to patient condition related due to patient having small vessel in left femoral per clinical. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18909 | 1220648-2024-18959 | 20355957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355957&pc=OZD | 09/10/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 09/10/2022 | 10/02/2024 |  |  | 03/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023208253 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the us complainant had placed a cp for support when there were low flows present with the pump. the pump was explanted and replaced without harm or injury from the interruption. | [5 empty] | The investigation for the reported low pump flow has been completed. the product was returned, and the low flow issue was confirmed. per the results of the clinical review and investigation: the logs were reviewed. an abnormally high motor current was observed when the pump was turned on. product was visually inspected, and biomaterial was observed around the impeller. this report is being filed as part of a retrospective review of historical records.
18910 | 1220648-2024-18961 | 20355958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355958&pc=OZD | 07/01/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 07/01/2023 | 10/02/2024 |  |  | 04/30/2025 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349638 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella was inserted into the left ventricle, and the device functioned optimally. however, when the physician attempted to remove the peel-away sheath, he inadvertently advanced the impella further into the left ventricle, resulting in prolapse and kinking of the device, causing it to fold upon itself. consequently, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of product damage (cannula kink). the device was not received for evaluation. the root cause of the reported issue could not be determined; however, the most likely root cause of the product damage (cannula kink) was due to use related since an improper technique was used to remove the peel away. this report is being filed as part of a retrospective review of historical records.
18911 | 1220648-2024-18962 | 20355965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20355965&pc=OZD | 07/04/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 07/05/2023 | 10/02/2024 |  |  | 02/28/2025 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322864 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a broken hemostasis valve was reported, resulting in bleeding from the valve in two consecutive cases during the use of the peel-away sheath. there was no additional treatment other than replacing to a repositioning sheath. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of access site bleeding. the device was not received for evaluation. the root cause of the access site bleeding was most likely damaged valve based on provided clinical information. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18912 | 1220648-2024-18960 | 20356008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356008&pc=OZD | 07/03/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 07/03/2023 | 10/02/2024 |  |  | 12/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024300728 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella was introduced, but there was vascular tortuosity and the device would not advance. the insertion was aborted and the patient survived the procedure. | [5 empty] | The investigation into the delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the most likely cause is due to the patient¿s condition related to tortuous blood vessels. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18913 | 1220648-2024-18954 | 20356014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356014&pc=OZD | 09/09/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 09/09/2022 | 10/02/2024 |  |  | 06/30/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235324 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a pump stopping when the physician cross clamped the device. unable to restart while cross clamped despite multiple attempts and the physician decided to explant the device. it was reported that the patient survived the explant and procedure. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the data logs confirm high motor current pump stops. the root cause of the pump stop was most likely use related as the clinical details specify that cross clamping the pump caused the pump stop. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18914 | 1220648-2024-18963 | 20356036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356036&pc=OZD | 09/11/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 09/11/2022 | 10/02/2024 |  |  | 03/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023199207 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after nine days of support, the impella cp's pump stopped. the medical staff tried to start pump several times without success and the impella cp was explanted. no additional information is available. | [5 empty] | The investigation for the reported delivery issue has been completed. data logs showed the pump stopped immediately upon the sudden occurrence of the "impella stopped - motor current high" alarm. all pump performance metrics were within specification prior to the pump stop. visual inspection of the returned pump revealed a large impeller purge gap. no other abnormalities were found. in conclusion, it was determined that the cause of the pump stop was bearing wear. the device was used beyond its intended design life of 8 days per design trace matrix 0046-9910. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18915 | 1220648-2024-18958 | 20356092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356092&pc=OZD | 07/03/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 07/04/2023 | 10/02/2024 |  |  | 03/31/2025 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024334311 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to being implanted with an impella cp device for mechanical circulatory support. the impella cp optical pump was unable to advance through a patient's existing intravascular calcification. the implant was subsequently aborted as a result of the noted issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient intravascular calcification was too strong to deliver. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18916 | 1220648-2024-18964 | 20356109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356109&pc=OZD | 03/02/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 03/02/2022 | 10/02/2024 |  |  | 12/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173602 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported during the impella cp being placed across the aortic valve, the cannula kinked and prolapsed on itself. it was then placed across the aortic valve successfully. however, once the pump was started, continuous suction occurred with low flows. the pump was then removed and a new impella cp was placed. there was no patient harm reported and this device was replaced. | [5 empty] | The investigation into the low pump flow issue has been completed. the impella pump was returned for investigation. a slight kink was found on the returned product. the root cause of the low flows is most likely due to a cannula kink. this report is being filed as part of a retrospective review of historical records.
18917 | 1220648-2024-18965 | 20356163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356163&pc=OZD | 02/28/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 03/01/2022 | 10/02/2024 |  |  | 08/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127129 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was not purging through the catheter and the physician decided not to use the device. this device was replaced. there was no patient harm reported and the replacement device purged appropriately. | [5 empty] | The investigation into the purge leak - pump issue has been completed. the impella pump and purge cassette were returned for investigation. the returned pump and cassette were able to purge without issue. the data logs were returned and showed the pump purge pressure within a normal range. the logs showed the pump going through all case start steps successfully. the root cause of the priming issue was "no problem found" as the pump had no issues priming when returned. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
18918 | 1220648-2024-18938 | 20356224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356224&pc=OZD | 07/03/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 07/03/2022 | 10/02/2024 |  |  | 03/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023208584 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump had a sudden reduced flow and stopped. the restart pump protocols/algorithm was followed without success. the decision was made to replace the cp with a new one. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the returned pump was visually inspected and biomaterial was observed in the outflow. the cause of the issue was biomaterial. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18919 | 1220648-2024-18967 | 20356226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356226&pc=OZD | 09/16/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/16/2021 | 10/02/2024 |  |  | 06/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022104240 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella cp device for mechanical circulatory support. during insertion of the guidewire into the impella pump pigtail catheter, the wire perforated the pigtail catheter. the decision was made to remove the original device and implant a new impella pump. the procedural outcome is the patient survived. | [5 empty] | The investigation for the reported hole in catheter issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that visual inspection of the returned pump revealed a small puncture hole in the curved end of the pigtail. no damage or abnormalities were found on the returned guidewire. the failure mode was able to be reproduced when inserting the guidewire into the pump with force. therefore, it was determined that the cause of the perforated hole in the pigtail was excessive force when inserting the guidewire. this report is being filed as part of a retrospective review of historical records.
18920 | 1220648-2024-18968 | 20356238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356238&pc=OZD | 07/30/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 07/30/2023 | 10/02/2024 |  |  | 02/28/2025 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333106 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that due to the patient¿s vasculature the impella cp was unable to be placed. reportedly after multiple attempts with different wires and catheters, the physician was unable to cross the aortic valve to place the impella. the impella placement was aborted. additionally, upon impella removal, the physician placed an intra-aortic balloon pump and chose to keep the impella introducer occlusive. there was no patient harm reported and this device was removed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. the investigation into the delivery issue has been completed. based on clinical description, the root cause of delivery issue was most likely due to patient condition. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18921 | 1220648-2024-18971 | 20356283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356283&pc=OZD | 09/16/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/16/2021 | 10/02/2024 |  |  | 05/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022088930 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella device rapidly increased and the pump shut off. it was stated that it was likely a thrombus was ingested. however, upon further examination, the medical professional confirmed no clot was present. the impella pump was removed, and a second impella device was implanted to manage this complication. the procedural outcome was the patient survived. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the field rt log was reviewed and an immediate high motor current pump stop occurred upon initial impella activation. the returned pump was visually inspected and a fracted red lumen was observed in the outflow cage. the root cause of the pump stop was determined to be a fractured red lumen. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18922 | 1220648-2024-18946 | 20356289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356289&pc=OZD | 07/04/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 07/04/2022 | 10/02/2024 |  |  | 03/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206097 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp pump stopped as the flow was saturated. the doctor intervened and placed impella in the right femoral artery without issue. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
18923 | 1220648-2024-18972 | 20356296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356296&pc=OZD | 09/16/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/17/2021 | 10/02/2024 |  |  | 06/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100275 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, various alarms were received followed by the pump stopping after the device was pulled back into the aorta. the impella device was explanted and a new impella device was placed. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device revealed a large purge gap indicative of bearing wear. the bearing wear occurred after six days of use, which is less than the intended design life of 8 days per durability test p2011-144. the root cause of the pump stop was determined to be bearing wear. also, high purge pressure alarms were noted sporadically throughout the logs. the data logs show normal purge pressure and purge flow followed by a sudden spike in pp up to 1100 mmhg. the purge pressure then recovers, and the purge flow increases slightly before stabilizing at around 15 ml/hr. this indicates a periodic kink in the purge line before the pss as the root cause of the high purge pressure. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18924 | 1220648-2024-18973 | 20356318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356318&pc=OZD | 07/11/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 07/11/2023 | 10/02/2024 |  |  | 01/31/2025 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024332205 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed a small hematoma at the impella access. additionally, the heparin-induced thrombocytopenia returned positive, prompting a switch from heparin to argatroban. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
18925 | 1220648-2024-18975 | 20356328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356328&pc=OZD | 09/17/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/17/2021 | 10/02/2024 |  |  | 06/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022104377 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the motor current increased and flow saturation was noted. a pump stop alarm was received. the impella was explanted and a new impella device was placed. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs showed that the pump stopped immediately upon the occurrence of the "impella stopped - motor current high" alarm. prior to the pump stop, the motor current was elevated and increased rapidly following a "suction" alarm. the motor current elevation corresponded with the ps waveform suddenly becoming ventricular with a brief ventricular position alarm, likely indicating when biomaterial was ingested. additionally, purge pressure rapidly increased with heparin levels noted to be at 0 throughout support. the flow rate was elevated at this time. the returned pump was able to successfully activate and its performance was within specification. visual inspection of the pump after its run revealed biomaterial around the impeller and in the purge gap. in conclusion, it was determined that the root cause of the pump stop was ingested biomaterial. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18926 | 1220648-2024-18950 | 20356333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356333&pc=OZD | 02/22/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 02/22/2023 | 10/02/2024 |  |  | 11/30/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290592 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, foot pulses were not palpable, but skin color was good, and the feet were warm. however, no additional details were provided regarding the treatment or management of this condition. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
18927 | 1220648-2024-18976 | 20356337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356337&pc=OZD | 09/17/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/17/2021 | 10/02/2024 |  |  | 06/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094419 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, it was observed that the left ventricular signal was negative. fluoroscopy revealed that the cannula was kinked. the physician stated he met resistance when pump was advancing over the aorta arch and that the patient had a calcified aorta. the original impella was explanted, and a new device was placed for support. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs showed recurrent "suction" alarms with accompanying low pump flow and decreasing motor current pulsatility. no other relevant abnormalities were noted. a minor cannula kink was observed near the motor housing. otherwise, no damage or abnormalities were found on the returned pump. after reviewing the clinical information, it was determined that the root cause of the cannula kink was most likely patient condition, specifically when advancing the pump across the patient's calcified aorta. this report is being filed as part of a retrospective review of historical records.
18928 | 1220648-2024-18977 | 20356354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356354&pc=OZD | 09/17/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/17/2021 | 10/02/2024 |  |  | 03/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022067206 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. during implantation, numerous efforts were made to advance the pump over the guidewire, however attempts were unsuccessful. it was stated that the process was aborted and the impella would be implanted at a later time. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issues issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was most likely the patient¿s condition. impella cp with smartassist for use during cardiogenic shock and high risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
18929 | 1220648-2024-18953 | 20356364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356364&pc=OZD | 04/29/2022 | 10/02/2024 |  | 10/02/2024 |  |  | 04/29/2022 | 10/02/2024 |  |  | 12/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023179011 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support for over 14 days, the pump stopped. the pump was unable to be restarted and was explanted. no patient harm was reported. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. the pump was used beyond the indication for use. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18930 | 1220648-2024-18979 | 20356377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356377&pc=OZD | 09/17/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/17/2021 | 10/02/2024 |  |  | 07/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022112035 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. while on support, the initial impella failed after insertion. efforts were made to get the device to start, but it would not. the patient became unstable, and the decision was made to explant the original impella pump, and implant second impella device. the procedural outcome is the patient survived surgery. | [5 empty] | The investigation for the reported pump did not start issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18931 | 1220648-2024-18980 | 20356385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356385&pc=OZD | 09/22/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/22/2021 | 10/02/2024 |  |  | 06/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022096808 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during preparation, the impella device failed to purge despite the medical staff identifying clear fluid at the outlet cage. it was stated that the process was aborted. there is no indication as to whether a second device was utilized, and the procedural outcome is unknown. | [5 empty] | The investigation for the reported priming issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that no data logs were sent. no damage or abnormalities were found on the returned pump. the pump purged with no abnormalities or purge alarms. therefore, the root cause of the pump failing to purge during case setup was use related. this report is being filed as part of a retrospective review of historical records.
18932 | 1220648-2024-18981 | 20356404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356404&pc=OZD | 07/20/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 07/20/2023 | 10/02/2024 |  |  | 10/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024291654 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported due to the patient¿s vasculature, an abdominal aneurysm, the impella cp was unable to be placed. there was no patient harm reported and this device was removed. | [5 empty] | The investigation into the delivery issue has been completed. the product was not returned for review. based on clinical description, the root cause of delivery issue was most likely due to patient condition. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18933 | 1220648-2024-18982 | 20356409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356409&pc=OZD | 09/19/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/19/2021 | 10/02/2024 |  |  | 06/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022101256 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, suction alarms were noted from the impella pump, and the device was removed to manage the complication. weaning was successful, and the device explanted. | [5 empty] | The investigation for the reported pump stop has been completed. the product was returned and the pump stop was confirmed. the root cause of the pump stop was due to improper technique in that the red lumen was not removed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
18934 | 1220648-2024-18956 | 20356410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356410&pc=OZD | 06/22/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 06/22/2023 | 10/02/2024 |  |  | 10/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024286802 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed emergently and experienced low flows rates and suction events. echocardiography indicated proper placement across the aortic valve. the physician noted that these issues were likely attributed to the presence of a clot. consequently, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was returned for evaluation and confirmed low pump flow. the root cause of the low pump flow was determined to be biomaterial. this report is being filed as part of a retrospective review of historical records.
18935 | 1220648-2024-18983 | 20356422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356422&pc=OZD | 09/19/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/19/2021 | 10/02/2024 |  |  | 02/28/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022051842 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support suction alarms were noted from the impella pump, and the device was removed to manage the complication. weaning was successful and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18936 | 1220648-2024-18984 | 20356436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356436&pc=OZD | 07/20/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 07/20/2023 | 10/02/2024 |  |  | 07/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022109782 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported due to the patient¿s vasculature, the impella was unable to be placed. there was no patient harm reported and this device was removed. | [5 empty] | The investigation into the delivery issue has been completed. the product was not returned for review. based on clinical description, the root cause of delivery issue was most likely due to patient condition. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18937 | 1220648-2024-18966 | 20356447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356447&pc=OZD | 09/15/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/15/2021 | 10/02/2024 |  |  | 07/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022109304 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a complainant with severe myocarditis and sinus tachycardia was implanted with an impella cp device for mechanical circulatory support. while on support, persistent high purge pressure alarms were received. the impella pump was removed, and a new pump was implanted to manage this complication. weaning was successful, the device was explanted, and the procedural outcome is the patient survived. | [5 empty] | The investigation for the reported high purge pressure issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
18938 | 1220648-2024-18985 | 20356454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356454&pc=OZD | 07/14/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 07/14/2023 | 10/02/2024 |  |  | 03/31/2025 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024344762 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella placement was aborted due to the patient¿s vasculature creating placement/delivery issues. there was no patient harm reported and this device was removed. | [5 empty] | The investigation into the delivery issue has been completed. the cause of the delivery issue was patient condition related to anatomy. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18939 | 1220648-2024-18969 | 20356467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356467&pc=OZD | 09/16/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/16/2021 | 10/02/2024 |  |  | 05/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022087028 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the pump flow only increased to one liter per minute. the original impella device was removed and reinserted. as the issue persisted, fluoroscopy identified that the impella cannula was kinked. the original impella pump was removed and replaced with a new impella device. there was no patient harm reported. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the kink in the cannula was confirmed. the root cause of the low pump flow was determined to be a cannula kink. the root cause of the cannula kink was most likely delivery issues due to patient condition or improper technique. this report is being filed as part of a retrospective review of historical records.
18940 | 1220648-2024-18970 | 20356473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356473&pc=OZD | 07/26/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 07/26/2023 | 10/02/2024 |  |  | 03/31/2025 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341746 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported due to the patient¿s vasculature the physician was unable to cross the aortic valve to place the impella. the physician attempted to use a straight guide wire and j-wire. however, the physician was still unable to cross the aortic valve after 15 minutes of attempting. the decision was made to abort the impella placement and proceed with the percutaneous coronary intervention of the occluded left anterior descending artery. there was no patient harm reported and this device was removed. | [5 empty] | The investigation into the delivery issue has been completed. the impella device was not received from the customer. per the given clinical condition, the root cause of the delivery issue was patient condition related to stenotic valve. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18941 | 1220648-2024-18987 | 20356475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356475&pc=OZD | 07/11/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 07/11/2023 | 10/02/2024 |  |  | 04/30/2025 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024339272 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 89 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a thrombus was confirmed near the inflow cage on echocardiography. therefore, impella cp was removed early as a precautionary measure. the patient survived the procedure. | [5 empty] | The investigation for the reported thrombus has been completed. the impella device was received from the customer. no thrombus was observed on the impeller, cannula and inflow cage. no other abnormalities were observed. data logs showed the motor current spike due to biomaterial interaction. the root cause of the thrombus issue was most likely patient condition related. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18942 | 1220648-2024-18974 | 20356481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356481&pc=OZD | 09/16/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/16/2021 | 10/02/2024 |  |  | 05/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022091173 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the purge system was unable to prime. the impella pump was explanted and replaced with a new impella device. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported priming issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the field rt log was reviewed and after insertion into the patient, a purge disc not detected alarm triggered. the returned purge cassette was connected to an aic, run through de-air, connected to the returned pump and purge flow was within specification. the cause of the priming issue was most likely use related since the purge disc was likely not inserted and the purge operated to specification on the returned product. this report is being filed as part of a retrospective review of historical records.
18943 | 1220648-2024-18988 | 20356488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356488&pc=OZD | 09/22/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/22/2021 | 10/02/2024 |  |  | 06/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022093188 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while the impella catheter was being inserted, the impella cannula became kinked at the aortic arch. the initial impella device was explanted and a new device was successfully placed. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18944 | 1220648-2024-18978 | 20356492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356492&pc=OZD | 07/25/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 07/01/2023 | 10/02/2024 |  |  | 03/31/2025 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024357296 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported due to the patient¿s vasculature, the physician was unable to cross the aortic valve. the aortic valve measured at 2.5 centimeters. the impella cp was unable to be placed. the percutaneous coronary intervention procedure was performed without impella support. there was no patient harm reported and this device was removed. | [5 empty] | The investigation has been completed. the impella device was discarded by the customer. based on clinical description, the root cause of the delivery issue was most likely due to patient condition. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18945 | 1220648-2024-18989 | 20356512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356512&pc=OZD | 09/23/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/23/2021 | 10/02/2024 |  |  | 07/31/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110554 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during implantation the device crossed the aortic valve but would not advance. when the pump was being removed, it was observed that the impella was filled with thrombus. the medical staff was not comfortable with flushing the device for reuse. a new impella pump was placed without issue. | [5 empty] | The root cause of the reported thrombus was ingested biomaterial and was most likely anticoagulation management since the patient was believed to be improperly anticoagulated during the first pump's insert. the patient's anticoagulation was adjusted, after which the second pump was inserted without issue. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).
18946 | 1220648-2024-18990 | 20356518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356518&pc=OZD | 07/11/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 07/11/2023 | 10/02/2024 |  |  | 02/28/2025 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024337414 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that upon insertion of the short sheath access was lost, at which point the team used a single perclose to achieve hemostasis. reportedly the physician then attempted access on the left side. however access was not able to be gained on the left side due to tortuosity and the depth of the artery. the team went back to the right side and placed the impella. however, immediate suction was noted and the decision was made to abort the procedure and remove the impella. there was no patient harm reported and this device was removed. | [5 empty] | The investigation into the delivery issue has been completed. no product was returned. as per clinical information, impella attempted access on the left side femoral and was unable to advance due to tortuosity and depth of the artery. the cause of the delivery issue was patient condition related due to patient having tortuosity over the left femoral artery. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18947 | 1220648-2024-18986 | 20356524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356524&pc=OZD | 09/20/2021 | 10/02/2024 |  | 10/02/2024 |  |  | 09/20/2021 | 10/02/2024 |  |  | 06/30/2023 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095069 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was placed on venoarterial extracorporeal membrane oxygenation (va ecmo) with an impella cp for decompression. while preparing the pump, the device repositioned. efforts were made to advance the impella, however the impella cannula was observed to be kinked, twisted. the pump was explanted and a new impella device was placed. | [5 empty] | The investigation for the reported product damage (cannula kink) has been completed. the product was returned and the cannula kink was confirmed. the root cause of the cannula kink was determined to be wireless re-access which is not permitted with the use of impella cp. the patient was successfully weaned from the impella and the device explanted.
18948 | 1220648-2024-18991 | 20356552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356552&pc=OZD | 07/11/2023 | 10/02/2024 |  | 10/02/2024 |  |  | 07/11/2023 | 10/02/2024 |  |  | 03/31/2025 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024345156 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp pump could not be fully advanced into the patient because of challenging anatomy. therefore, the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation into the delivery issues has been completed. the impella device was not returned for evaluation. root cause of the delivery issue was patient condition related because patient had a history of severe peripheral vascular disease/peripheral artery disease and bilateral iliac stents from several years ago. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18949 | 1220648-2024-18992 | 20356624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20356624&pc=OZD | 01/24/2023 | 09/26/2025 |  | 10/02/2024 |  |  | 01/24/2023 | 10/02/2024 |  |  | 10/31/2024 |  | 10/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023278411 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there was an optical sensor failure at the start of priming. the aic was swapped in an attempt to troubleshoot the issue, but the failure remained. the pump was subsequently explanted and the impella cp was replaced. no further information is available. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the data log shows a placement signal not reliable alarm. the optical sensor system error was reproduced on the console when the returned pump was plugged into an aic in the lab. no abnormalities were observed during the light test and the sensor was not broken. an abnormality was observed on the patient cable plug, specifically a piece of the plug was missing. the root cause of the optical signal issue was use related and related to damage to the patient cable plug causing an air gap condition in which the plug could not be fully seated in the receptacle. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d8/d9 device returned to manufacturer updated h3 device evaluated by manufacturer updated. corrections that were made from the initial manufacturer device report number 1220648-2024-18992 g1 reporting contact email and g1 manufacturing site name and address updated. | The impella device investigation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
18950 | 1220648-2024-18788 | 20343524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343524&pc=OZD | 09/01/2024 | 10/15/2024 |  | 10/01/2024 |  |  | 09/01/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505859 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support had an ischemic stroke. the patient had been critically ill and there had been a concern of impaired neurological status prior to the stroke. the patient eventually expired after the family withdrew care. medical safety review of the event noted that the use of the device cannot conclusively be associated with the outcome. patient's peri-procedural condition was critical. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
18951 | 1220648-2024-18789 | 20343551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343551&pc=PYX | 09/01/2024 | 02/26/2025 |  | 10/01/2024 |  |  | 09/01/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1324936 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 70-year-old female patient was on impella rp support when there was a controller failure on the automated impella controller (aic). the engineering team determined that the controller error was due to persistent placement signal issues. the aic was replaced. the patient had been implanted with the rp to help generate flow to the left ventricular assist device (lvad). the patient ultimately expired on support, despite being on support with lvad and triple vasopressors. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). patient's peri-procedural condition was critical. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) and aic - loss of opm data has been completed. according to the device and data analysis the root cause for the aic issues was a defective optical bench lamp. d2 common device name revised on manufacturer device report in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g4 pma/510(k)number was inadvertently entered incorrectly on manufacturer device report.
18952 | 1220648-2024-18790 | 20343635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343635&pc=OZD | 09/08/2024 | 11/01/2024 |  | 10/01/2024 |  |  | 09/08/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025518092 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and after the implant, the patient had ventricular fibrillation (vf). cardiopulmonary resuscitation was performed, 1 lidocaine pushed, and the impella position and performance was optimized. furthermore, it was reported that there was oozing at the access site. a fem stop was placed and sodium bicarb was kept in the purge. anticoagulation was not administered until after the oozing subsided. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the ventricular arrhythmia and access site bleed has been completed. the device was not returned for evaluation. clinical details were provided which were sufficient to determine root cause. the root cause of access site bleeding is determined to be patient movement because the patient was moving his legs. the root cause of the ventricular arrhythmia could not be determined without additional clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18953 | 1220648-2024-18791 | 20343650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343650&pc=OZD | 06/04/2024 | 10/01/2024 |  | 10/01/2024 |  |  | 09/06/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1365450 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a 34-year-old male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, there were placement signal not reliable alarms. the automated impella controller (aic) audio was disabled and the patient's support continued. on september 6, 2024, the investigation was completed and it was found via logs that a controller error alarm (yellow) was noted on the aic. the patient survived the explant. | [5 empty] | The investigation into the reported automated impella controller (aic) - controller error (yellow alarm) has been completed. the product was returned for investigation. during data analysis, the logs confirmed that a placemental singal not reliable alarm had occurred, followed by a controller error alarm. following the set of errors, the optical bench was intentionally reset, which resolved the placement signal issue. real time logs confirmed that immediately prior to the controller error alarm, the optical bench reported an oscillating contrast value far outside the normal range of contrast. during device analysis, the console was investigated and the issue was notable to be reproduced in the lab when running a known-good pump and purge system. the cause of the aic issue was a defective optical bench.
18954 | 1220648-2024-18792 | 20343677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343677&pc=OZD | 09/10/2024 | 12/03/2025 |  | 10/01/2024 |  |  | 09/10/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1138237 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that impella automated impella controller (aic) had a controller error alarm. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation has been completed. data review: the case associated with the reported complaint was started on (b)(6) 2024. immediately after the bootup console displayed controller error (loss of comm with keypad or between the kbd and i2c for kbd hw revision 1) ¿ alarm, event #24. this confirms the reported failure mode. device analysis: this console was tested after arriving in fs. upon power cycling the console 200 times, it was unable to reproduce the reported failure mode - controller error. found the front panel optical connector contaminated, unable to remove the debris upon cleaning (unrelated to the reported failure mode). conclusion: the root cause of the intermittent keyboard communication issue could not be determined due to the inability to reproduce. g1: corrected the manufacturer email address h3: added information regarding the investigation status. h6: added codes per the results of the investigation h10: added the results of the investigation
18955 | 1220648-2024-18794 | 20343704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343704&pc=OZD | 09/09/2024 | 11/21/2024 |  | 10/01/2024 |  |  | 09/09/2024 | 10/01/2024 |  |  | 11/30/2025 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024424640 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella 5.5 support. the impella 5.5 was primed and placed in the left ventricle. immediately after pumping started, the "purge pressure increase" alarm sounded, and the condition did not improve despite multiple attempts to bleed air from the purge line. the purge set was replaced, but the alarm did not go away. using impella connect, it was discovered that the "purge pressure increase" alarm had been present from the priming process before the pump was inserted into the body. the pump was replaced and there was no reported patient harm. | [5 empty] | The investigation for the high purge pressure has been completed. the pump was returned and evaluated. the high purge pressure reproduced in the lab while priming and running the pump for 15 minutes. the purge fluid built up above the sidearm reservoir, which caused the high purge pressure. no purge fluid exited from the sidearm to the purge gap. the sidearm was cut off and the purge cassette tubing was connected straight to the red handle without the sidearm. the high purge pressure resolved, which isolates the blockage to the sidearm. a bond failure from internal manufacturing causing a blockage in the sidearm is suspected. added- h6 impact code 4629 | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
18956 | 1220648-2024-18793 | 20343710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343710&pc=OZD | 09/12/2024 | 11/06/2024 |  | 10/01/2024 |  |  | 09/12/2024 | 10/01/2024 |  |  | 10/31/2025 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445544 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Disability | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp to support a 57-year-old male patient admitted in an st-elevation myocardial infarction (stemi). the pump was placed femorally but, the access site had blood loss that was treated by femstop placement and a purse string suture. when then transferred from ccl to the icu the patient had runs of ventricular tachycardia (vt) and to remedy the pump was repositioned and p-level was dropped. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report.
18957 | 1220648-2024-18795 | 20343718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343718&pc=OZD | 09/13/2024 | 11/07/2024 |  | 10/01/2024 |  |  | 09/13/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520888 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 49-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. the patient was on impella cp support and the after repositioning, with good position, the impella p-levels were increased and the flows failed to meet the appropriate range. there were suction alarms, and the staff attempted to reposition to improve flows to no avail. the pump was explanted, and it was reported that a clot was found to be ingested. | [5 empty] | The investigation for the low pump flow and thrombus has been completed. data logs were reviewed which confirmed the failure mode and clinical narrative. logs showed pump was started and run without issue for 95 hours, then motor current (mc) spiked followed low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. based on clinical description, the root cause of low flow was most likely biomaterial ingestion. added-h6 impact code 4628. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
18958 | 1220648-2024-18796 | 20343727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343727&pc=OZD | 09/13/2024 | 11/11/2024 |  | 10/01/2024 |  |  | 09/13/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532194 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A decision was made by the physician to remove the impella due to bleeding concerns and unable to stop aggrastat. the cp was removed in the critical care unit and the patient expired within three hours of admission to the unit. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant reported that during the impella cp placement they were unable to get the orogastric (og) tube in the patient. the patient was on aggrastat and started to bleed out of the nose and mouth. trauma surgeon in room to place rhino-rockets in nose for bleeding. in the cardiac care unit (ccu), the patient developed a large soft hematoma with bleeding. pressure was held and fem stop applied. a decision was made by the physician to remove the impella due to bleeding concerns and the inability to stop aggrastat. the cp was removed in the ccu. the patient¿s outcome is unknown. |  |  |  |  | The investigation of access site bleeding and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. based on provided clinical details, the root cause of the access site bleed was anticoagulation management. due to lack of clinical details and potential lack of relation between the complication and the pump, the root cause of the vascular damage could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-18796: b2 death and date of death added. b5 patient outcome and death rational statement added. e4 should have been left blank. g1 reporting contact fax number missing. h1 type of reportable event. h6 health effect - impact code 1802 added. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
18959 | 1220648-2024-18803 | 20343841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343841&pc=OZD | 09/02/2024 | 11/16/2024 |  | 10/01/2024 |  |  | 09/06/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025513202 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 54-year-old male patient with cardiomyopathy and other systemic comorbidities was implanted with an impella cp device for mechanical circulatory support. during impella cp support, the impella cp purge fluid was leaking. the day after the impella was placed, the purge fluid was slowly leaking from the pressure reservoir. the patient survived the explant. | [5 empty] | The investigation of purge leak ¿ pump has been completed. the impella cp was returned for evaluation. no visual abnormalities were noted. pump was purged for 10 minutes then run while purging for an additional 10 minutes. no leaks purge leaks from the pump were noted. pump was also purged using a syringe in attempt to exaggerate any leaking however no leaks were again noted. data logs were reviewed and no significant decreases in purge pressure with increased purge flow were noted throughout case. clinical details noted that a leak was coming from the light blue area of the purge reservoir. the root cause of the purge leak - pump was most likely use related as no leak was able to be reproduced in the lab. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
18960 | 1220648-2024-18800 | 20343847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343847&pc=OZD | 09/15/2024 | 11/07/2024 |  | 10/01/2024 |  |  | 09/15/2024 | 10/01/2024 |  |  | 09/30/2025 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444883 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had no pulses in the impella leg. additionally, the impella groin shot through the repositioning sheath showed dissection where the two perclose devices failed. the impella was removed, and vascular repair was successfully completed. the patient survived the event. | [5 empty] | The investigation for the limb ischemia and vascular damage has been completed. no product was returned for evaluation. clinical details mention damage to vasculature occurred after perclose device failure. the root cause of the vascular damage - peripheral vessel was most likely due to a use related issue. the root cause of the limb ischemia could not be determined without clinical details. corrections to the initial mfr report: d4 model number. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
18961 | 1220648-2024-18801 | 20343865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343865&pc=OZD | 09/15/2024 | 11/11/2024 |  | 10/01/2024 |  |  | 09/15/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508826 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during the impella 5.5 placement the md made multiple attempts using several different wires and catheters to cross valve, but it was unsuccessful. the 5.5 placement was aborted due to the patient anatomy. the patient¿s outcome is unknown. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the pump could not be implanted and was aborted due to the patient's anatomy. added- h5 labeled for single use is yes. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
18962 | 1220648-2024-18797 | 20343882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343882&pc=OZD | 09/04/2024 | 11/23/2024 |  | 10/01/2024 |  |  | 09/06/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534317 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an attempted impella cp implant and it was aborted due to the pump not being able to pass through the heart valve. the cannula kept kinking. the implant was aborted and a different pump was used. the was no reported patient harm. | [5 empty] | The investigation for the delivery issues has been completed. the pump was returned for investigation and the kinking was noted on the cannula. the root cause of delivery issue was determined to be use issue because they were able to insert the second pump without any issues. | A.5 is unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
18963 | 1220648-2024-18798 | 20343939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343939&pc=PYX | 09/14/2024 | 09/17/2025 |  | 10/01/2024 |  |  | 09/14/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025543823 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the physician attempted to peel away the 23fr sheath but did not have control of the choker on the sutures and the patient lost around 1l of blood. the physician was able to tamp down the choker and get hemostasis. three units of red blood cells were administered in the operation room. the impella flows dropped to 2.2lpm. in the intensive care unit, one unit of blood was given for the bleeding and hemolysis was noted. the impella placement signal was also not working properly. the physician decided to remove the impella. the patient¿s outcome is unknown. | [5 empty] | The investigation for the hemolysis, access site bleed, and low pump flow has been completed. the pump was not returned for evaluation. the data logs show a motor current spike outside a p-level change followed by a decrease in pump flows. there was an upward trend in placement signal with rising motor current trend and low flows during the case. the root cause of hemolysis was most-likely due to biomaterial. the root cause of the access site bleeding major was most likely due to a use-related issue. the root cause of the low pump flow was most likely due to biomaterial. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the root cause of hemolysis was most likely due to biomaterial. the root cause of the access site bleeding was most likely due to a use-related issue. the root cause of the low pump flow was most likely due to biomaterial. the root cause of the optical signal issue was not determined as product was not returned and limited information related to the failure was provided. all pertinent information available to abiomed has been submitted at this time.
18964 | 1220648-2024-18799 | 20343950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343950&pc=OZD | 09/14/2024 | 11/08/2024 |  | 10/01/2024 |  |  | 09/14/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532457 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that heparin was discontinued due to bleeding concerns for the patient on impella cp support. it was further reported that the patient had gi bleed. the patient survived the event. | [5 empty] | The investigation for the gastrointestinal bleed has been completed. the product was not returned for investigation. a data log review was not required for this case. according to the given clinical information, the patient was off anticoagulants due to a gi bleed on the 5th day of pump use. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not returned. the relation between gi bleed and impella device was unknown. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: d4-model number. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
18965 | 1220648-2024-18804 | 20343984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343984&pc=OZD | 09/06/2024 | 12/03/2025 |  | 10/01/2024 |  |  | 09/06/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1183444 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller had a controller error. this issue was discovered during use in the catheterization lab. however, the patient information is unknown. there was no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
18966 | 1220648-2024-18802 | 20343999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20343999&pc=OZD | 05/06/2024 | 11/11/2024 |  | 10/01/2024 |  |  | 09/10/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed complaints team was forwarded a publication entitled "preemptive impella 5.5 insertion to reduce operative risk in high-risk cardiac surgery: a case report" in which the impella 5.5 patient had atrial fibrillation on day 5 post op (on impella support) treated by amiodarone with successful conversion to normal sinus rhythm. the impella was explanted on day 7 post op and no other information provided. | [5 empty] | The investigation for the reported atrial arrhythmia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the atrial arrhythmia could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. literature citation: echieh, c. p., ryan, a., cherian, a., rohilla, y., wang, k., & kazui, t. (2024). preemptive impella 5.5 insertion to reduce operative risk in high-risk cardiac surgery: a case report. international journal of surgery case reports, 121, 109947. https://doi.org/10.1016/j.ijscr.2024.109947.
18967 | 1220648-2024-18805 | 20344043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20344043&pc=OZD | 03/06/2024 | 11/12/2024 |  | 10/01/2024 |  |  | 09/10/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | There was a literature review done and one of the publications was entitled "the sideclose technique: a novel method for achieving hemostasis with an indwelling impella cp". it was reported "persistent bleeding from the groin site, despite impella sheath manipulation and tightening of the perclose suture. the sideclose technique was performed, resulting in immediate hemostasis". | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue since hemostasis was obtained by using sideclose technique along with tightening the perclose. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ a.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. literature citation: korngold, e., & wollmuth, j. (2024). the sideclose technique: a novel method for achieving hemostasis with an indwelling impella cp. journal of the society for cardiovascular angiography & interventions, 3(7), 102141. https://doi.org/10.1016/j.jscai.2024.102141.
18968 | 1220648-2024-18806 | 20344258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20344258&pc=OZD | 09/06/2024 | 11/06/2024 |  | 10/01/2024 |  |  | 09/06/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142850 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during preparation, the automated impella controller shuts down at start up with an error message. the aic had to be restarted several times and still gives the same error message. there was no patient involvement. | [5 empty] | The investigation of automated impella controller - pump stop has been completed. when console was booted up, there was an immediate controller error (defective optical sensor lamp) alarm issued. the 5s was triggered by the bad lamp and had a voltage of 0.46 volt (v) which is the value for a disconnected or dead lamp. this failure mode reproduced later on in the month as well. the failure mode was not reproduced during bench top testing. 5s testing showed the light outputting >1v and the lamp was visibly outputting light. there was no evidence for why the failure mode repeatedly occurred in the field. the cause of the issue was most likely a defective lamp assembly. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
18969 | 1220648-2024-18808 | 20344269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20344269&pc=OZD | 09/06/2024 | 11/08/2024 |  | 10/01/2024 |  |  | 09/06/2024 | 10/01/2024 |  |  | 09/30/2025 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024399417 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 57-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. during impella support, the pump failed to deliver beyond the native iliacs despite all troubleshooting measures. the iliac had angioplasty but the pump delivery failed. the team instead placed an intra-aortic balloon pump (iabp). | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the impella pump implant was aborted due to patient's anatomy. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
18970 | 1220648-2024-18807 | 20344299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20344299&pc=OZD | 09/10/2024 | 11/04/2024 |  | 10/01/2024 |  |  | 09/10/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531510 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella cp implant however, once the impella was in the left ventricle it would not move freely. the doctor attempted to reposition the pump, but ended up pulling the pump out of the left ventricle. the doctor said the pump was not tracking well over the wire and elected to replace the pump. the was no reported patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. product was not returned for review. based on clinical description, pump was pulled back across valve inadvertently by physician, the root cause of positioning issue was most likely use related. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
18971 | 1220648-2024-18809 | 20344883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20344883&pc=PYX | 09/10/2024 | 11/21/2024 |  | 10/01/2024 |  |  | 09/11/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025462121 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Us complainant reported the patient was on impella rp flex support and the patient had oozing at the access site after the implant. the mattress sutures were tightened, and support continued. additionally, there was an abrupt decrease in impella flow with concurrent decrease in purge flow to >2ml/hr and purge pressure high alarm. purge pressure high with intermittent purge blocked alarms for the next 36hrs. tissue plasma activator was implemented in the purge. the patient survived the event. | [5 empty] | The investigation for the reported high purge pressure, access site bleeding - major and low pump flow has been completed. the product was returned and the reported observations were not confirmed. the cause of the high purge pressure issue was most likely biomaterial based on data log review and returned product investigation. the biomaterial interaction could lead to temporary low pump flow during a high purge pressure event. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the placement signal issue was dp sensor drift. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation of low pump flow, high purge pressure, and access site bleeding has been completed. the cause of the high purge pressure issue was most likely biomaterial based on data log review and returned product investigation. the biomaterial interaction could lead to temporary low pump flow during a high purge pressure event. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the placement signal issue was dp sensor drift. d9 date device returned to manufacturer added. g4 pma/510(k)number incorrect.
18972 | 1220648-2024-18810 | 20345044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20345044&pc=OZD | 09/11/2024 | 11/13/2024 |  | 10/01/2024 |  |  | 09/11/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525337 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 67-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. while on support the patient had poor leg perfusion. the decision was made to escalate to an impella 5.5 and perform a percutaneous coronary intervention (pci). | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
18973 | 1220648-2024-18812 | 20345718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20345718&pc=OZD | 09/03/2024 | 11/05/2024 |  | 10/01/2024 |  |  | 09/11/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025506814 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality (loqi) indicator impella registry indicating an allegation of left ventricle (lv) thrombus: ¿9/3/24 lv thrombus on echo¿ for a 64-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella 5.5 device for mechanical circulatory support. the patient outcome was not provided. additionally, the complaints team received a notification via abiomed¿s long-term outcome and quality (loqi) indicator impella registry indicating an allegation of ventricular tachycardia (vt): ¿cardioverted¿. the patient recovered with no sequelae. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of vt and thrombus could not be determined as the device was not returned nor sufficient clinical details. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
18974 | 1220648-2024-18811 | 20345768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20345768&pc=OZD | 09/11/2024 | 11/11/2024 |  | 10/01/2024 |  |  | 09/11/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527742 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 59-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a large pannus that was pressing on the impella repositioning sheath. a pool of blood was noted at the patient¿s side that appeared to be coming from repositioning sheath location. the gauze were saturated with blood and it appeared that proper angle match was not in place. the sheath sutures had become loose, so the decision was made to cut them and place new ones. once new sutures and new angle match dressing were in place the bleeding was resolved. surgeon recommended the patient receive one unit of blood. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was determined to be use issue because it was noted that proper angle match was not in place and sheath sutures has become loose. the bleeding was resolved by angle match and new sutures in place. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
18975 | 1220648-2024-18814 | 20346139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20346139&pc=OZD | 12/29/2022 | 10/01/2024 |  | 10/01/2024 |  |  | 01/10/2023 | 10/01/2024 |  |  | 07/31/2024 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247149 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a thrombus was observed on the impella cp device upon removal of the device. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of thrombus. no product was returned. a returned photo confirmed the thrombus on the pump. the cause of the thrombus was patient condition related. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18976 | 1220648-2024-18815 | 20346499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20346499&pc=OZD | 12/29/2022 | 10/01/2024 |  | 10/01/2024 |  |  | 12/29/2022 | 10/01/2024 |  |  | 08/31/2024 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270049 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to being implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp pump could not be fully advanced into the patient because of challenging anatomy. therefore, the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the impella device was discarded by the customer. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was patient condition related and related to the tortuous and calcified iliac artery. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18977 | 1220648-2024-18816 | 20346578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20346578&pc=OZD | 12/30/2022 | 09/16/2025 |  | 10/01/2024 |  |  | 12/30/2022 | 10/01/2024 |  |  | 07/31/2024 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253691 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, flows were low. the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs show suction alarms began to trigger within a minute of initial ramp-up and persisted throughout support; the flow waveform was diastolic throughout support, with an average value of <1 l/min. meanwhile, high purge pressure alarms triggered following a 10-minute rise in purge pressure. visual inspection of the returned pump revealed the presence of biomaterial on impeller and in the purge gap. the pump passed all the post sterile inspection checks. no other abnormalities were found. therefore, the cause of the low pump flow was ingested biomaterial. corrections to the initial mfr report: g1 mdr reporting contact email.
18978 | 1220648-2024-18813 | 20346648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20346648&pc=OZD | 12/29/2022 | 10/01/2024 |  | 10/01/2024 |  |  | 12/29/2022 | 10/01/2024 |  |  | 08/31/2024 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023266092 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to being implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp pump could not be fully advanced into the patient because of calcified vasculature. therefore, the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation into the reported delivery issues has been completed. the impella device was discrded by the customer. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was determined to be patient condition as the patient had severely calcified iliacs. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18979 | 1220648-2024-18818 | 20346677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20346677&pc=OZD | 09/12/2024 | 11/20/2024 |  | 10/01/2024 |  |  | 09/12/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1432077 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that the automated impella controller (aic) is not recognizing the purge cassette. no harm has been reported. | [5 empty] | The controller was received for evaluation. upon completion of the investigation, a supplemental mdr will be filed. | The investigation for the reported "aic - purge issue - other" issue has been completed. the product was returned and the "aic - purge issue - other" was confirmed. the purge flag was confirmed to be missing which would trigger the inability for the purge disc to be recognized. the cause of the issue was a missing purge flag.
18980 | 1220648-2024-18821 | 20347180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20347180&pc=OZD | 09/12/2024 | 11/13/2024 |  | 10/01/2024 |  |  | 09/12/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525122 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon starting the impella cp, flows of 0.1l noted with flattened motor current. the physician repositioned cp back and forth in the left ventricle under fluoroscopy with some pulsatility noted but no resolution with the poor flows. the impella cp was replaced. there was no harm to the patient. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of low pump flow has been completed. pump was returned for investigation and no cannula kinks or biomaterial was observed. the pump was primed and ran in a bucket of water to reproduce failure but the pump ran as per the specification. data logs were investigated and 2 suction alarms were observed corelating to p-level changes. clinical mentions that patient had tortuous illiac vessel but without any supporting evidence in the product investigation, this cannot be deemed as the reason for the failure. low pump flows were observed at the start of the case. the root cause of the low flow issue was most likely patient condition related.
18981 | 1220648-2024-18822 | 20347761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20347761&pc=OZD | 09/12/2024 | 11/11/2024 |  | 10/01/2024 |  |  | 09/12/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536668 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hypovolemia | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that hematoma was noted at impella cp access site post percutaneous coronary intervention procedure in holding area. the team provided manual pressure to site with no change. the patient became volume dependent and required pressor support. the patient was brought for a computed tomography scan and arterial bleed noted at the right femoral artery. the patient was transfused with multiple blood products. in the operating room, impella was weaned and explanted and vessel repaired. hemostasis achieved. the patient survived the event. | [5 empty] | The investigation for the vascular damage has been completed. the pump was not returned for investigation, and a data log review was not required for this case. according to the provided clinical details, the vascular surgery successfully resolved the access site bleeding issue. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
18982 | 1220648-2024-18823 | 20347944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20347944&pc=OZD | 09/12/2024 | 11/06/2024 |  | 10/01/2024 |  |  | 09/12/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527392 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had placed a cp to support a patient admitted in a non-st elevated myocardial infarction. the pump was placed but low flows and suction occurred. the pump remained on for support until eventual wean and explant. the patient survived the event. | [5 empty] | The investigation for the low pump flow has been completed. pump was not returned for investigation. data logs were investigated. there was a suction alarm followed by low flow where flow dropping to 0.2l/m. there was a p-level change observed following that but the flow never recovers. there was no motor current spike or any placement signal trend. the data logs are inconclusive and without the product the complaint cannot be further investigated. therefore, the root cause of the issue was not determined since product was not returned and data logs were inconclusive. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
18983 | 1220648-2024-18826 | 20348159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348159&pc=OZD | 08/31/2021 | 10/01/2024 |  | 10/01/2024 |  |  | 08/31/2021 | 10/01/2024 |  |  | 06/30/2023 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108231 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump stopped. the device was attempted to be restarted multiple times, which was unsuccessful. the device was then explanted, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was discarded by the customer, clinical information did not provide sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
18984 | 1220648-2024-18825 | 20348229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348229&pc=OZD | 08/31/2021 | 10/01/2024 |  | 10/01/2024 |  |  | 08/31/2021 | 10/01/2024 |  |  | 03/31/2023 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022067330 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the team observed lower than expected flows from the pump. the device was then removed and replaced with no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18985 | 1220648-2024-18827 | 20348250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348250&pc=OZD | 08/31/2021 | 10/01/2024 |  | 10/01/2024 |  |  | 09/14/2021 | 10/01/2024 |  |  | 05/31/2023 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089351 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis. there was no reported patient harm. | [5 empty] | The investigation for the reported hemolysis issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
18986 | 1220648-2024-18817 | 20348275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348275&pc=OZD | 12/31/2022 | 10/01/2024 |  | 10/01/2024 |  |  | 12/31/2022 | 10/01/2024 |  |  | 08/31/2024 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268839 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient attempted to get out of bed and the pump was pulled back into the aorta. the impella cp was explanted and replaced with a new impella cp. the patient survived the procedure. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the positioning issue was determined to be patient movement per the clinical details. this report is being filed as part of a retrospective review of historical records.
18987 | 1220648-2024-18819 | 20348295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348295&pc=OZD | 08/11/2024 | 11/05/2024 |  | 10/01/2024 |  |  | 09/12/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025513949 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had placed a 5.5 pump for the support of a patient who had a heart transplant at age of 2. patient admitted in with acute myocardial infarction / cardiogenic shock and heart failure. patient had a venous-arterial extra corporeal membrane oxygenation (va ecmo) running while decisions being made as the patient was not a ventricular assist device or transplant candidate. the 2nd day of support the patient developed a gi bleed that prompted purge change to sodium bicarbonate and removal of heparin. the pump remains on for support now over a month later. | [5 empty] | The investigation for the gastrointestinal bleed has been completed. pump was not returned for investigation. data log review was not required for this case. as per the clinical, gi bleed was reported during case run. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. the relation between gi bleed and impella device was unknown. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
18988 | 1220648-2024-18828 | 20348334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348334&pc=OZD | 09/01/2021 | 05/29/2025 |  | 10/01/2024 |  |  | 09/01/2021 | 10/01/2024 |  |  | 12/31/2022 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022024963 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, there were multiple pump stops. the device was removed and replaced, and there was no reported harm to the patient. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related. the patient refused to lay flat due to back pain. the patient was moving constantly and refused to lay still | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18989 | 1220648-2024-18830 | 20348352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348352&pc=OZD | 01/03/2023 | 09/25/2025 |  | 10/01/2024 |  |  | 01/04/2023 | 10/01/2024 |  |  | 11/30/2023 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161052 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. it was reported that the impella cp was placed via femoral artery cutdown. due to ecmo cannula in the femoral artery, the cp was placed higher, in the iliac. given the depth of the artery, upon closure of the skin, both the entire blue hub and the junction where the sterile sleeve connects to the repo sheath were inside the skin. the impella was explanted and replaced with a new impella cp. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation of the positioning issues has been completed. after reviewing the clinical information, it was determined that the cause of the positioning issues was use related. corrections to the initial mfr report: g1 mdr reporting contact email h6 investigation findings 114, investigation conclusion 19 added.
18990 | 1220648-2024-18820 | 20348457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348457&pc=OZD | 12/31/2022 | 10/01/2024 |  | 10/01/2024 |  |  | 02/13/2023 | 10/01/2024 |  |  | 10/31/2024 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275619 |  | 048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced hemolysis. action taken to resolve the issue was unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was discarded by the customer. therefore, an investigation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. this report is being filed as part of a retrospective review of historical records.
18991 | 1220648-2024-18836 | 20348511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348511&pc=OZD | 09/03/2021 | 10/01/2024 |  | 10/01/2024 |  |  | 09/03/2021 | 10/01/2024 |  |  | 07/31/2023 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110242 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, there was a kink located distal to the motor housing after starting pump. attempts to re-position unsuccessful, and new pump inserted. case successfully supported with the second pump. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller instructions for use for the related event are as follows: section: warnings, cautions, and precautions ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ impella cp with smartassist for use during cardiogenic shock and high risk pci and impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records.
18992 | 1220648-2024-18834 | 20348526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348526&pc=OZD | 01/04/2023 | 10/01/2024 |  | 10/01/2024 |  |  | 01/05/2023 | 10/01/2024 |  |  | 08/31/2024 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256213 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. a thrombus was developed in the right femoral artery coinciding with impella support. the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | The investigation for the reported thrombus has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the thrombus was due thrombus formation in the right femoral artery. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
18993 | 1220648-2024-18837 | 20348591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348591&pc=OZD | 04/09/2023 | 10/01/2024 |  | 10/01/2024 |  |  | 04/09/2023 | 10/01/2024 |  |  | 12/31/2024 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318013 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the sterile sleeve was stiff and folded over itself during impella delivery. as a result of the fold, the repositioning sheath was unable to be pulled back to allow for the peel away sheath to be removed fully from the body before breaking away. the impella cp was explanted and care was escalated to impella 5.5. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the device was not received for evaluation. the root cause of the stiff sterile sleeve was unable to be determined as no product was returned for review. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
18994 | 1220648-2024-18842 | 20348607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348607&pc=OZD | 04/10/2023 | 10/01/2024 |  | 10/01/2024 |  |  | 04/10/2023 | 10/01/2024 |  |  | 11/30/2024 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024308431 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported a failure to prime the device. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
18995 | 1220648-2024-18824 | 20348688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348688&pc=OZD | 09/12/2024 | 11/06/2024 |  | 10/01/2024 |  |  | 09/12/2024 | 10/01/2024 |  |  |  |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025485973 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, the patient developed hemolysis. additionally, aortic and left ventricle waveforms became separate and rapidly moved to top of screen with simultaneous loss of flow and pulsatile motor current. the decision was made to explant the pump. the patient survived the event. | [5 empty] | The investigation of low pump flow and hemolysis has been completed. the cause of the low pump flow issue was most likely patient condition related, based on returned product investigation and data log review. the cause of the hemolysis issue was most likely improper positioning of the device based on given clinical details, data log review and returned product investigation. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
18996 | 1220648-2024-18838 | 20348724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348724&pc=OZD | 09/03/2021 | 10/01/2024 |  | 10/01/2024 |  |  | 09/03/2021 | 10/01/2024 |  |  | 06/30/2023 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022103274 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the pump became kinked after making the bend around steep aortic arch and then maintained the kink as pump was inserted into left ventricle unable to get device unkinked resulting in flows <2lpm. the impella device was removed and replaced with a new impella. | [5 empty] | The investigation for the reported cannula kink issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. a submitted video of fluoroscopic imaging confirmed the reported cannula kink adjacent to the pump outflow. the root cause of the cannula kink was most likely patient condition. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller instructions for use for the related event are as follows: section: warnings, cautions, and precautions ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ impella cp with smartassist for use during cardiogenic shock and high risk pci and impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records.
18997 | 1220648-2024-18831 | 20348805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348805&pc=OZD | 09/02/2021 | 10/01/2024 |  | 10/01/2024 |  |  | 09/02/2021 | 10/01/2024 |  |  | 07/31/2023 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105331 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, flow saturation, and pump stop during use of the device was reported. the pump was removed and no new impella was inserted. there was no adverse impact to the patient. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
18998 | 1220648-2024-18832 | 20348806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348806&pc=OZD | 01/03/2023 | 03/27/2025 |  | 10/01/2024 |  |  | 01/03/2023 | 10/01/2024 |  |  | 08/31/2024 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270813 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, the pump was pulled back and triggered a suction alarm. multiple attempts were made to reposition the pump, however, resistance was met and a bend was seen in the cannula. the impella was explanted and replaced with a new impella cp. the patient survived the procedure. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that no data logs were returned and no logs are available on constellation. a visual inspection of the pump revealed a kink in the cannula. the cause of the positioning issue was use related and related to re-crossing of the valve wirelessly. this report is being filed as part of a retrospective review of historical records. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
18999 | 1220648-2024-18840 | 20348842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348842&pc=OZD | 04/09/2023 | 03/07/2025 |  | 10/01/2024 |  |  | 04/09/2023 | 10/01/2024 |  |  | 12/31/2024 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024318023 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the sterile sleeve was stiff during impella delivery. the patient continued on support until the device was successfully weaned. | [5 empty] | The returned product was investigated and no abnormalities were seen on the sterile sleeve. a clamp mark and deformation to the catheter was seen at the 38cm marking. resistance was felt at this area when the repositioning unit was slid up the catheter. the cause of the product damage was use related and related to clamping of the catheter causing deformation. this report is being filed as part of a retrospective review of historical records. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
19000 | 1220648-2024-18841 | 20348844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348844&pc=OZD | 04/10/2023 | 10/01/2024 |  | 10/01/2024 |  |  | 04/10/2023 | 10/01/2024 |  |  | 03/31/2024 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213262 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. upon initial insertion of the pump, the patient exhibited irregular electrocardiogram changes. during manipulation of the pump, the patient progressed to an extended complete heart block. the physician withdrew the pump into the aorta and ordered it to be turned off. the impella was then fully removed from the body. a transvenous pacemaker was quickly placed, and the decision was made to reinsert the same impella after it was washed with heparinized saline. the impella was replaced and functioned without incident for the duration of the procedure. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
19001 | 1220648-2024-18833 | 20348884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348884&pc=OZD | 01/04/2023 | 09/25/2025 |  | 10/01/2024 |  |  | 01/04/2023 | 10/01/2024 |  |  | 09/30/2024 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278535 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump reportedly stopped as a result of the red lumen breaking off in the pump inlet. the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation of detached inflow/outflow - great vessel has been completed. the real time (rt) log shows an impella stopped motor current high alarm when a pump start is attempted. the pump is then removed. the cause of the pump not starting was most likely a red lumen removal issue. corrections to the initial mfr report: g1 mdr reporting contact email h6 investigation findings 114, investigation conclusion 19, investigation type 4121 added.
19002 | 1220648-2024-18835 | 20348951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348951&pc=OZD | 08/14/2021 | 05/22/2025 |  | 10/01/2024 |  |  | 03/23/2022 | 10/01/2024 |  |  | 04/30/2023 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022075998 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, hemolysis was observed at impella placement to removal. weaning was successful, the device was explanted, and the procedural outcome was the patient survived explant. | [5 empty] | The cause of the hemolysis was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
19003 | 1220648-2024-18839 | 20348964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20348964&pc=OZD | 04/09/2023 | 10/01/2024 |  | 10/01/2024 |  |  | 04/10/2023 | 10/01/2024 |  |  | 07/31/2024 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269012 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to being implanted with an impella cp device for mechanical circulatory support. during attempted insertion, it was noted that the pump failed to advance into position. the implant was aborted, and it was discovered upon removal that the shaft was kinked. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
19004 | 1220648-2024-18844 | 20349057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20349057&pc=OZD | 09/03/2021 | 10/01/2024 |  | 10/01/2024 |  |  | 09/03/2021 | 10/01/2024 |  |  | 06/30/2023 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022107665 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the first impella cp was implanted had inadequate low flows with impella in good position and adequate volume status. this second impella cp was inserted which had better flows but still decreased. after troubleshooting position and giving patient fluid it was determined that the patient¿s anatomy may be the cause of decreased impella flows. ultimately the decision was made to continue with primary percutaneous coronary intervention and left main was successfully ballooned three times and stented x1. impella was successfully weaned and explanted in the cath lab. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. as this failure occurred on two separate devices on the same patient it is unlikely that this failure was a malfunction of the product. it is likely that the patient's anatomy hindered the pump¿s ability to achieve optimal flows. this report is being filed as part of a retrospective review of historical records.
19005 | 1220648-2024-18843 | 20349077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20349077&pc=OZD | 04/11/2023 | 10/01/2024 |  | 10/01/2024 |  |  | 04/11/2023 | 10/01/2024 |  |  | 01/31/2025 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312913 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 45 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was malposition in the left ventricle due to patient's challenging anatomy. the impella was explanted and replaced with a new impella cp to address the issue. the patient survived the procedure. | [5 empty] | The investigation for the reported positioning issues has been completed. the device was returned for evaluation. the root cause of the positioning issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
19006 | 1220648-2024-18829 | 20349178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20349178&pc=OZD | 01/03/2023 | 10/01/2024 |  | 10/01/2024 |  |  | 01/03/2023 | 10/01/2024 |  |  | 09/30/2024 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023279008 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, cannula became kinked, and flows were low. the impella cp device was explanted and replaced with a impella 2.5. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was returned for evaluation, and the low pump flow was confirmed. the root cause of the low pump flow was due to a kinked cannula because of improper use technique. this report is being filed as part of a retrospective review of historical records.
19007 | 1220648-2024-18845 | 20349409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20349409&pc=OZD | 09/03/2021 | 10/01/2024 |  | 10/01/2024 |  |  | 09/03/2021 | 10/01/2024 |  |  | 07/31/2023 |  | 10/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110465 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, after starting the pump, the impella appeared to have a kink, and flows were less than optimal. when attempting to fix the left main the impella was not providing enough support due to poor flows. the impella was also not allowing the guide to get into position. the impella cp was removed and there was no harm to the patient. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. the root cause of the low flows was determined to be a cannula kink as it was seen on fluoroscopy. the root cause of the cannula kink was unable to be determined as no logs or product was returned for this investigation. this report is being filed as part of a retrospective review of historical records.
19008 | 1220648-2024-18742 | 20332331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20332331&pc=PYX | 09/04/2024 | 11/21/2024 |  | 09/30/2024 |  |  | 09/05/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025495312 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient on impella rp support ongoing due to being admitted in with right ventricle failure. the pump was placed and after seven days, there was hemolysis observed. the team wants the impella rp investigated despite the assertion that there could be many causes of the hemolysis, per communication with field representative. the patient remains on impella support. | [5 empty] | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
19009 | 1220648-2024-18743 | 20332396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20332396&pc=OZD | 09/05/2024 | 11/21/2024 |  | 09/30/2024 |  |  | 09/05/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527088 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and while on support, the flow of the pump was reduced and did not match the p-level. echocardiography showed that the cannula appeared to be kinked. the pump was explanted and replaced. the patient survived the event. | [5 empty] | The investigation for the low pump flow has been completed. the pump was returned for evaluation and was visually inspected. no biomaterial observed. no cannula kink or broken blade found. mold observed around purge gup and some places in the pump. upon functional testing, the pump was connected to aic to produce the pump low flow issue. pump run on p-9 for 10 minutes. no low flow observed and flow was within required range. no abnormality observed. data log not returned for review.the cause of the low pump flow cannot be determined since no abnormality found and low pump flow cannot be reproduced and data log not available for review. added-h6 med dev prob code 2889 | A.5 is unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
19010 | 1220648-2024-18744 | 20332423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20332423&pc=OZD | 08/18/2024 | 11/18/2024 |  | 09/30/2024 |  |  | 09/07/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025509348 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. it was reported that the patient dropped out from the bed and because that, the purge filter broke. the pump was explanted the same day. it was noted that the patient did well and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the purge leak- pump and purge leak-cassette has been completed. the purge cassette was returned and evaluated. the purge cassette was de-aired and connected to a test pump. purge flow was ~18 ml/hr and purge pressure was ~350-450 mmhg. a window was cut into the cassette near the piston. a leak was observed from the piston seal. the root cause of the purge cassette leak was a piston failure. the part number for the purge cassette was updated from 2000163 to 0043-3011 because the returned product (associated with the complaint pump set) was a gen 1 cassette. the pump was not returned. the root cause of the pump purge leak was due to use-related sidearm filter damage. | The purge cassette was received, but the pump was not. upon completion of the investigation, a supplemental mdr will be filed.
19011 | 1220648-2024-18745 | 20332477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20332477&pc=OZD | 09/05/2024 | 11/21/2024 |  | 09/30/2024 |  |  | 09/05/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508658 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported that impella 5.5 insertion was aborted because the pump could not advance into the aorta. it did not pass the truncus. no patient harm has been reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of determined to be patient condition because the pump was unable to pass through due to the stenosis in the vessels. | A.5 is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
19012 | 1220648-2024-18746 | 20332601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20332601&pc=OZD | 09/05/2024 | 11/06/2024 |  | 09/30/2024 |  |  | 09/05/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494550 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed over the .018 wire for support of a patient admitted in with a st elevated myocardial infarction (stemi). the patient had delayed support for the stemi. the patient had left the hospital and then returned. the patient had mitral valve disease, coronary artery disease, and pulseless electrical activity and cardiopulmonary resuscitation during the cardiac catheterization lab support. the first impella cp was accidentally pulled from the left ventricle and the second impella cp (subject device) was kinked during delivery. the patient expired. | [5 empty] | The investigation has been completed. the failure mode was changed to delivery issues as this is what was alleged and the case was aborted. both the pump and guidewire were damaged as a result of this issue. the medical doctor felt resistance when trying to remove the wire and afterwards the wire and pump appeared kinked on fluoroscopy and had to be removed. the patient was obese. no other patient conditions were mentioned. the cause of the delivery issues was most likely use related as resistance was felt when removing the wire and excessive force likely was applied due to this. the wire and pump alleged to be kinked as a result but not returned to confirm with investigation. h.6 due to investigation results, code 3036 and 2889 should not have been coded on initial manufacturer device report number 1220648-2024-18746. codes 2682 and 829 were added due to investigation results. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
19013 | 1220648-2024-18748 | 20332760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20332760&pc=OZD | 09/07/2024 | 11/05/2024 |  | 09/30/2024 |  |  | 09/07/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526486 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support experienced a large brain bleed. left side flaccid, follows commands on right. heparin was turned off and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
19014 | 1220648-2024-18747 | 20332769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20332769&pc=OZD | 09/05/2024 | 11/08/2024 |  | 09/30/2024 |  |  | 09/05/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025510371 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a 14 french sheath placed via the femoral with single access for a high risk percutaneous coronary intervention. the sheath had a malfunction at the hub that lead to leaking. the team removed and replaced the leaking sheath. the patient survived the event. | [5 empty] | The investigation for the introducer damage has been completed. the product was not returned for investigation, and a data log review was not required for this case. according to the clinical report, the hub on the 14 french introducer was found to be split and a leak was reported from the valve. the cause of the access site bleeding issue was damaged introducer valve, based on given clinical details. introducer damage could not be verified without a returned product or provided clinical image. g.1 revised reporting contact fax number in accordance with updated procedures since initial manufacturer device report 1220648-2024-18747 was submitted. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
19015 | 1220648-2024-18749 | 20332896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20332896&pc=PYX | 09/08/2024 | 11/18/2024 |  | 09/30/2024 |  |  | 09/08/2024 | 09/30/2024 |  |  | 01/31/2025 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024299928 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella rp support. while on support, there was a sensor failure on the impella rp. the decision was made to remove the impella rp due to sensor failure. there was no reported patient harm. | [5 empty] | The investigation for the placement signal issue has been completed. no product was returned. log review showed that the placement signal and pump flow metrics were lost on the fourth day of support around the time of the "placement signal not reliable" alarms with no change in motor current at that time. the placement was gradually trending downward before going out of range, and the pump flow was trending upward before dropping to zero. all of these log observations indicate sensor drift. rp c3s do not have an optical sensor, so there is no optical signal review in this investigation. the root cause of the placement signal issue was most likely sensor drift. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
19016 | 1220648-2024-18750 | 20332981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20332981&pc=OZD | 09/05/2024 | 11/04/2024 |  | 09/30/2024 |  |  | 09/05/2024 | 09/30/2024 |  |  | 04/30/2025 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024339406 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Fluid/Blood Leak | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 65-year-old male patient was on impella 5.5 support and the pump had suction alarms. the patient was also on intra-aortic balloon pump (iabp). blood products were given. iabp was eventually discontinued due to suction alarms. furthermore, there was bleeding at the anastomosis site. graft dehisced from the aorta. entire impella and graft were removed to apply partial aortic clamp to perform aortic repair. moreover, the investigating engineer for the automated impella controller (aic) found that the purge issue that was originally thought to be the aic, is likely an issue from the pump. there were purge system open alarms. the staff replaced the purge cassette without resolution. the aic was replaced. the pump was successfully weaned and explanted after a day of support. | [5 empty] | The investigation of low pump flow and vascular damage - great vessel has been completed. log review showed noisy placement signal, motor current, and pump flow along with frequent fluctuations in p-level from p-3 to p-8. there were several suction alarms throughout the case. after 23 hours of support, the motor current steadied for about nine hours while at p-5. the root cause of the low pump flow was not determined since the product was not returned and insufficient clinical detail was provided. bleeding was found at the anastomosis site after the patient had increased chest tube output and was hypotensive. the root cause of the vascular damage was not determined since insufficient clinical detail was provided. the failure mode purge leak - pump was removed because the purge issue was resolved after switching automated impella controllers. the console investigation can be found under manufacturer device report 1220648-2024-18506. h.6 medical device problem code 1250 is to be omitted from manufacturer device report 1220648-2024-18750 due to the information above. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
19017 | 1220648-2024-18751 | 20333071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20333071&pc=OZD | 09/06/2024 | 12/03/2025 |  | 09/30/2024 |  |  | 09/06/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1125040 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when doing a routine automated impella controller (aic) check, the aic had a controller failure alarm. the purge system had stopped. they switched to a back-up controller. there was no patient involvement. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
19018 | 1220648-2024-18752 | 20333075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20333075&pc=OZD | 09/08/2024 | 11/08/2024 |  | 09/30/2024 |  |  | 09/08/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532884 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and the patient was profusely bleeding at the access site. the patient coded and the staff gave 3 units of blood products as well as epinephrine injection. the patient's pulse returned, but the bleeding continued. the impella was explanted and the vessel was repaired after explant. the patient outcome is unknown. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
19019 | 1220648-2024-18753 | 20333222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20333222&pc=OZD | 09/06/2024 | 11/11/2024 |  | 09/30/2024 |  |  | 09/06/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532230 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was on impella cp support and after the pump was started the purge fluid 9-10 milliliters (ml) gradually decreased to 4.9 ml where it stayed. no change in purge pressure (575-601), motor current or impella flow. no kinks in purge line. the impella cp was taken out of auto mode and turned down to p5 before wiring percutaneous coronary intervention (pci) vessel due to mean arterial pressure in the 140's. the purge fluid increased to 5.1 ml. pci was done and impella cp was weaned to p3 and purge fluid increased to 6-7 ml. by the time the cp was on p2, purge fluid 8 ml. the patient survived the event. | [5 empty] | The investigation of the purge leak ¿ pump has been completed. the field data logs were reviewed and initially purge flow drops from 10 milliliters per hour (ml/hr) to five ml/hr and was stable for the remainder of the run. no motor current spikes, purge alarms, or suction alarms were observed. the product was returned and evaluated. the returned purge cassette tubing was tied off so it was not used in reproduction testing. the pump was set up and ran for 10 minutes in the lab. no abnormalities were found. the failure mode was updated from "purge leak - pump" to "high purge pressure" because the purge flow decreased. there was no issue found during the case. the device functioned/operated to specification. h.6 due to investigation results, medical device problem code 1250 was reported incorrectly on the initial manufacturer device report number 1220648-2024-18753 and has been updated on this report. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
19020 | 1220648-2024-18754 | 20333229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20333229&pc=OZD | 09/08/2024 | 10/24/2025 |  | 09/30/2024 |  |  | 09/08/2024 | 09/30/2024 |  |  | 11/30/2025 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024429577 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | [5 empty] | The complainant reported the patient was on impella cp support and there were placement signal issues. the doctor was questioning the validity of the sensor on the device so the pump was replaced with an impella 5.5. the patients outcome is unknown. | Care was eventually withdrawn, and the patient expired. |  |  |  |  | The investigation of optical signal issue has been completed. the impella device was not returned for evaluation. the root cause of the optical signal issue cannot be determined as limited clinical details and insufficient product and data logs were provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned for evaluation. data logs were only available for up to day three of support and not at time of reported issue. clinical details noted that clinical team was concerned about the validity of the sensor as the placement signal reading was different than the patient's actual blood pressure readings. the root cause of the optical signal issue cannot be determined as limited clinical details and insufficient product and data logs were provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 adverse event updated. b2 updated to death. b5 updated to include death. h1 updated to death. h6 updated to include death and optical signal coding. b7 other relevant history, including preexisting medical conditions updated. g1 reporting contact email updated.
19021 | 1220648-2024-18755 | 20333346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20333346&pc=OZD | 09/08/2024 | 10/31/2024 |  | 09/30/2024 |  |  | 09/08/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533049 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp support initiated for a st elevation myocardial infarction and percutaneous coronary intervention. the team placed the cp but on the day of implant the patient was seen to have left ventricle clot. heparin was stopped and the pump was explanted the next day. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
19022 | 1220648-2024-18756 | 20333461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20333461&pc=OZD | 09/09/2024 | 11/21/2024 |  | 09/30/2024 |  |  | 09/09/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527122 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 64-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella implant, the sheath was not able to be inserted in the groin. the procedure was completed with a new sheath. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue most likely use issue since the procedure was completed successfully with the second sheath. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
19023 | 1220648-2024-18757 | 20333526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20333526&pc=OZD | 09/09/2024 | 02/20/2025 |  | 09/30/2024 |  |  | 09/09/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1112113 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Hemodynamic instability | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 73-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the impella automated impella controller (aic) showed a blank white screen and was constant alarming. when doing a controller exchanged the pump stopped and the patients pressure dropped rapidly and became unstable. the patient was medically managed until back on the new console. the patient was transferred to another facility via flight. the nurse was taking the console out of flight mode and then switching over to the new facility owned console when the problem occurred. the nurse stated that the console was not taken out of flight mode yet when the screen went solid white, and the pump stopped working. once the new console transfer was completed the patient rapidly became stable. | [5 empty] | The investigation for the console (aic) pump stop has been completed. data logs were reviewed which confirm that unexpected controller shutdown occurred. the console was returned for evaluation. the console screen went white with a constant screech upon reboot in the lab. attempts to restart the console in maintenance mode were futile with the console no longer booting up in bios. isolation of the epc by swapping with a well known good one resolved issue. the root cause for the unexpected controller shutdown was a defective epc. | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
19024 | 1220648-2024-18758 | 20333679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20333679&pc=OZD | 09/09/2024 | 11/25/2024 |  | 09/30/2024 |  |  | 09/09/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025497403 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 59-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. during impella support, there was blood leaking around the connection between the catheter plug and shaft. the bleeding was uncontrollable so the impella was removed. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed and product damage has been completed. data logs are not relevant to the investigation as the complication occurred immediately after insertion. returned product was investigated. the red handle was opened and it was full of blood. then, the catheter was inspected, and a puncture hole was found right around the 80 cm mark. based on the puncture in the catheter on returned product, the root cause of the product damage was determined to be a use issue. the root cause of the access site bleed could not be determined without additional information. corrections to the initial mfr report: d4-udi number.
19025 | 1220648-2024-18759 | 20333788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20333788&pc=OZD | 09/09/2024 | 10/31/2024 |  | 09/30/2024 |  |  | 09/09/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546084 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 35-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, there was significant kinking at the catheter between the junction of the motor and the outlet during insertion. initially pump flows were adequate but then slowly decreased only allowing the pump to flow at p-1 with flows of 0.3 liters. the decision was made to remove and replace the pump with another cp. | [5 empty] | The investigation for the low pump flow has been completed. product and data logs were not returned for review. based on clinical description, the root cause of low flow was most likely due to cannula kink. the cause of cannula kink was patient condition related. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
19026 | 1220648-2024-18760 | 20334099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20334099&pc=OZD | 09/09/2024 | 11/04/2024 |  | 09/30/2024 |  |  | 09/09/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025482105 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 62-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, the pump became mal-positioned. the pump was shallow and posterior. the physician declined to reposition, and the pump was removed. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
19027 | 1220648-2024-18761 | 20334253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20334253&pc=PYX | 09/10/2024 | 11/04/2024 |  | 09/30/2024 |  |  | 09/10/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025529719 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a rp flex placed for a patient who had arrested and been given cardiopulmonary resuscitation in the field. an st elevated myocardial infarction was confirmed in the cardiac catheterization lab and the team placed the rp for support. the pump was malpositioned deep into the pulmonary artery and subsequently stopped. the pump was able to be restarted after attempts. the patient was sent to the intensive care unit and the family withdrew care. the patient expired. | [5 empty] | The investigation for the pump stop has been completed. the pump was not returned for analysis. the data logs showed that there was a motor current spike followed by a decrease in pump flow and placement signal, and a slight increase in purge pressure. the pump was able to be restarted after multiple attempts but flows remained lower throughout the rest of the case. based on the data logs the root cause of the pump stop was most likely due to ingested biomaterial. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
19028 | 1220648-2024-18762 | 20334364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20334364&pc=OZD | 09/10/2024 | 11/18/2024 |  | 09/30/2024 |  |  | 09/10/2024 | 09/30/2024 |  |  | 11/30/2025 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024426555 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported that the patient on impella cp support experienced hemolysis. the impella was in a good position. hemolysis was resolved with impella cp removal. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined.
19029 | 1220648-2024-18763 | 20334406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20334406&pc=OZD | 01/09/2023 | 09/16/2025 |  | 09/30/2024 |  |  | 01/09/2023 | 09/30/2024 |  |  | 08/31/2024 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272381 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, poor flows were observed as there was an unfixable kink in the cannula. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported product damage issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a visual inspection of the returned product revealed a cannula kink near the outflow cage. no other damage or abnormalities were found. therefore, it was determined that the cause of the low pump flow was a kinked cannula due to the patient's small aorta and ventricle. corrections to the initial mfr report: g1 mdr reporting contact email.
19030 | 1220648-2024-18765 | 20334622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20334622&pc=OZD | 09/10/2024 | 10/28/2024 |  | 09/30/2024 |  |  | 09/10/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1324017 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the automated impella controller (aic) had a battery failure. the aic was turned on to prime the impella cp and a battery failure was noted. the battery switch was flipped on and off to no avail. the aic was replaced. the aic was tested after, and the battery alarm persisted. no patient harm has been reported. | [5 empty] | The investigation for the console (aic) battery failure red alarm has been completed. after reviewing the imc logs, the reported battery failure alarm issue was confirmed. there were multiple instances of battery failure alarms (transport connection blown/missing) (event set #360) observed from the reported complaint date. the console was returned for evaluation and the battery failure issue was able to be reproduced. results of running batttest utility on telnet revealed that cell 2 in battery 1 had a voltage that was significantly less than the rest of the cells in the battery. the reported battery failure issue was determined to be caused by internal cell imbalance (found in cell 2 of battery 1). added- d9 device return date, h6 impact code 4629. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
19031 | 1220648-2024-18764 | 20334654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20334654&pc=OZD | 01/09/2023 | 10/08/2025 |  | 09/30/2024 |  |  | 01/09/2023 | 09/30/2024 |  |  | 09/30/2024 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275035 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, flows were low and suction alarms occurred. the team changed the concentration of dextrose in the solution. additionally, the patient developed a temperature with positive blood cultures. the decision was made to explant the device. the patient survived the procedure. | The complainant reported that the patient was being placed on impella cp for hemodynamic support. it was reported that a patient was on impella cp support ongoing when purge alarms were present. the team had low purge pressure and as such they troubleshot by pc and aic replacement. neither remedied the issue. the team then adjusted the purge solution from d5w to d10w. low pump flow while on support. it was reported that the patient developed a temperature with positive blood cultures. sepsis while on support . |  |  |  |  | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist. section: table 3.2 impella catheter components . ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿. section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. this report is being filed as part of a retrospective review of historical records. | The lt log confirms purge pressure low alarms. since switching out the purge cassette and the automated impella controller (aic) did not resolve the issue, the issue can be isolated to the pump. increasing the concentration of dextrose increases the viscosity of the purge fluid. the cause of the priming issue was not determined though since product was not returned.
19032 | 1220648-2024-18766 | 20334734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20334734&pc=OZD | 01/09/2023 | 09/30/2024 |  | 09/30/2024 |  |  | 01/10/2023 | 09/30/2024 |  |  | 01/31/2024 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206264 |  | 1000115 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella stopped multiple times and was not able to be restarted. the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | A.5 ethnicity is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
19033 | 1220648-2024-18767 | 20334839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20334839&pc=OZD | 09/10/2024 | 11/05/2024 |  | 09/30/2024 |  |  | 09/10/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501353 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Hematuria | [10 empty] | Required Intervention | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION (ECMO). | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed along with extra corporeal membrane oxygenation (ecmo) for a patient admitted in acute myocardial infarction / cardiogenic shock. the pump was placed and on day of placement, there was ventricular tachycardia that required shocks. this shock along with underlying condition may have caused/contributed to signs of hemolysis. the urine color change did not prompt explant. the pump remains on for support however there has been continuous renal replacement therapy added to the support and reperfusion sheath for the ecmo limb. the patient survived the event. | [5 empty] | The investigation of vf/vt/af/at and hemolysis has been completed. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ . section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿.
19034 | 1220648-2024-18768 | 20334859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20334859&pc=OZD | 01/09/2023 | 09/30/2024 |  | 09/30/2024 |  |  | 01/09/2023 | 09/30/2024 |  |  | 09/30/2024 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273716 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella was accidentally pulled back into the aorta. the physician could not recross the valve. therefore, leading to the decision to explant the device and replace it with a new impella cp. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records.
19035 | 1220648-2024-18769 | 20334942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20334942&pc=OZD | 01/10/2023 | 03/27/2025 |  | 09/30/2024 |  |  | 01/10/2023 | 09/30/2024 |  |  | 06/30/2024 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247238 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. a kink was found in the device via an x-ray. repositioning was unable to resolve the issue. therefore, the impella cp was explanted and replaced with a new impella cp. the patient survived the procedure. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned product was visually inspected, and a cannula kink was observed. the cause of the delivery issue was most likely a tight aortic arch patient condition resulting in a cannula kink. this report is being filed as part of a retrospective review of historical records. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
19036 | 1220648-2024-18770 | 20335544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20335544&pc=PYX | 09/10/2024 | 11/12/2024 |  | 09/30/2024 |  |  | 09/10/2024 | 09/30/2024 |  |  | 09/30/2024 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023266250 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Arrhythmia | [12 empty] | Life Threatening | Required Intervention | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when placing the impella rp and upon advancing the rp catheter to the tricuspid valve using clockwise rotation, the cannula buckled above the motor housing and wouldn't cross the valve. a second attempt was made again the cannula buckled. on this attempt the temp pacer was dislodged and the 027 wire fell back into the right ventricle (rv). patient then had an episode of ventricular tachycardia (vt) requiring 1 shock and cpr for approximately 1 minute before return of spontaneous circulation (rosc). patient stabilized and the rp insertion was aborted. | [5 empty] | The investigation for the ventricular tachycardia (vt) and delivery issues has been completed. the impella was not returned for evaluation. clinical details were provided to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature as per the clinical description provided. the root cause of the vt could not be determined without additional clinical details. the instructions for use for the rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ added h6 impact code 4628. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
19037 | 1220648-2024-18771 | 20335902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20335902&pc=OZD | 01/10/2023 | 03/06/2025 |  | 09/30/2024 |  |  | 01/10/2023 | 09/30/2024 |  |  | 11/30/2024 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278291 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump stopped. the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | A.5 is unknown. the impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issued has been completed. the returned product was visually inspected after soaking in water and biomaterial was observed in the purge gap and blood was observed in the purge line. the cause of the pump stop was biomaterial in the purge gap leading to blood ingress. this report is being filed as part of a retrospective review of historical records.
19038 | 1220648-2024-18773 | 20336621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20336621&pc=OZD | 01/10/2023 | 09/30/2024 |  | 09/30/2024 |  |  | 01/10/2023 | 09/30/2024 |  |  | 09/30/2024 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281619 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was falling out of the ventricle while the patient was rolling from side-to-side within the hospital bed. the physician was unable to recross the valve. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records.
19039 | 1220648-2024-18774 | 20336734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20336734&pc=OZD | 01/11/2023 | 09/30/2024 |  | 09/30/2024 |  |  | 01/11/2023 | 09/30/2024 |  |  | 09/30/2024 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275458 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention attempted to being implanted with an impella cp device for mechanical circulatory support. the impella cp pump could not be fully advanced into the patient because of challenging anatomy. therefore, the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation into the delivery issues has been completed. the impella device was not returned for evaluation as it was dscarded by the customer. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the delivery issues was determined to be patient condition as the patient had tortuous vasculature. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
19040 | 1220648-2024-18772 | 20337313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20337313&pc=OZD | 09/10/2024 | 10/31/2024 |  | 09/30/2024 |  |  | 09/10/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545466 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 61-year-old female/male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. during impella support, the limb pulses were diminished. the ischemia resolved when the pump was explanted after 2 days of support. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿.
19041 | 1220648-2024-18775 | 20337807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20337807&pc=OZD | 07/24/2024 | 11/07/2024 |  | 09/30/2024 |  |  | 09/10/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | On 2024-09-10 abiomed japan forwarded the publication entitled: "intracardiac shunt reversal and early right ventricular failure after left ventricular assist device implantation" in which an impella 5.5 pump support at uconn had a "he had an episode of hemodynamically significant blood loss from gastrointestinal bleeding which improved with blood transfusions although endoscopic evaluation was unrevealing". | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. article citation: pillai, a., jedeon, z., hammond, j., gluck, j., & jaiswal, a. (2024). intracardiac shunt reversal and early right ventricular failure after left ventricular assist device implantation. cureus. https://doi.org/10.7759/cureus.65320 the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-18775. h10 article citation missing. publication attachment missing.
19042 | 1220648-2024-18778 | 20338805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20338805&pc=PYX | 09/10/2024 | 10/31/2024 |  | 09/30/2024 |  |  | 09/11/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025484858 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had patient admitted in for the placement of a left ventricular assist device. post operation, the patient needed right sided support. the rp flex was placed to support the dilated right ventricle and heart failure. during the support the team had purge flow alarms and purge system blocked alarms that prompted tissue plasminogen activator to be added to the purge, and, when this did not resolve the alarm, the pump was explanted. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. according to the clinical details, high purge pressure alarms were observed the day after the implant. the doctor added tissue plasminogen activator (tpa) to the purge line to address the issue. the cause of the high purge pressure issue was most likely biomaterial, based on data log review.
19043 | 1220648-2024-18779 | 20338983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20338983&pc=OZD | 08/31/2024 | 11/05/2024 |  | 09/30/2024 |  |  | 08/31/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025530464 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 57-year-old male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, there was saturated pump flow. the motor current was flat and purge flow decreased. the pump placement was confirmed. the decision was made to add tissue plasminogen activator to the purge and monitor. dextrose was infusing through the purge cassette. additionally, the pump was pointed towards the mitral valve. aggressive attempts to reposition were made to no avail. the pump was left mispositioned. care was ultimately withdrawn and the patient expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the positioning issue and saturated flows has been completed. data logs were reviewed and saturated flows were confirmed on the date they were reported. there was a motor current spike and speed deviation followed by a two minute rise in purge pressure, drop in purge flow, and saturated pump flows. the pump was running at p-5 and was achieving flows above 4.0 liers per minute for the next hour. these are all indications of biomaterial ingestion, that ultimately clogged the purge gap, wrapped around the impeller shaft, and caused saturated flows. though purge pressure spiked it never was higher than 1023 mmhg, so the high purge pressure alarm never triggered. the saturated flows and purge pressure issues both resolved, likely due to the decision to infuse tissue plasminogen activator through the purge. clinical details provided suggest that continuous renal replacement therapy lines were rinsed back at the time of the complication which further supports the likelihood of biomaterial ingestion. of note, there were also "impella position in aorta" alarms triggering. the product was not returned for investigation. the root cause of saturated pump flows was determined to most likely be biomaterial based on data logs and clinical details provided. at the same time as the saturated flow complication, it was reported that positioning alarms triggered. data logs confirmed that the "impella in aorta" alarm did go off during the period of saturated flows. that being said, clinical details state that a transthoracic echocardiogram was done and confirmed pump position was acceptable. there were no reports of patient movement, patient conditions, or any other abnormalities that could have caused the positioning issues. the product was not returned for investigation. due to lack of clinical detail, the root cause of the positioning issues could not be determined. h.6 code 4628 was inadvertently omitted from initial manufacturer device report number 1220648-2024-18779 and was added.
19044 | 1220648-2024-18777 | 20339135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20339135&pc=OZD | 01/13/2023 | 09/30/2024 |  | 09/30/2024 |  |  | 01/13/2023 | 09/30/2024 |  |  | 08/31/2024 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269931 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | Difficult to Remove | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. attempts were made to advance the impella cp to remove a kink were unsuccessful. the impella catheter was pulled back into the descending aorta and kink was removed. another attempt was made to advance the impella cp to the left ventricle, but the kink was visualized in the same spot around the aortic arch. the physician stated the kink was due to patient anatomy and not a catheter issue. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
19045 | 1220648-2024-18776 | 20339894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20339894&pc=OZD | 01/12/2023 | 09/30/2024 |  | 09/30/2024 |  |  | 01/12/2023 | 09/30/2024 |  |  | 08/31/2024 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260922 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the pump flow was assessed and found to be suboptimal in auto-mode. a subsequent fluoroscopy revealed a kink in the cannula proximal to the outflow. the impella cp device was explanted. the patient survived the procedure. | [5 empty] | A.5 is unknown. the impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records.
19046 | 1220648-2024-18780 | 20340107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20340107&pc=OZD | 09/11/2024 | 12/04/2025 |  | 09/30/2024 |  |  | 09/11/2024 | 09/30/2024 |  |  |  |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1159707 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump support ongoing when there were automated impella controller (aic) alarms that prompted the team to replace the aic. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
19047 | 1220648-2024-18781 | 20340129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20340129&pc=OZD | 08/30/2021 | 03/13/2025 |  | 09/30/2024 |  |  | 04/19/2022 | 09/30/2024 |  |  | 12/31/2022 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022023302 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 90 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported that the patient experienced limb ischemia during use of the device. no intervention or premature device removal was indicated. in addition, the patient experienced bleeding and hematoma requiring a blood transfusion. no intervention was specified, but it was reported that the patient survived. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ limb ischemia: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
19048 | 1220648-2024-18782 | 20340451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20340451&pc=OZD | 07/07/2021 | 09/30/2024 |  | 09/30/2024 |  |  | 08/31/2021 | 09/30/2024 |  |  | 05/31/2023 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083592 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed hemolysis. there is no further information regarding treatment or resolution, but the patient survived with no further harm reported. | [5 empty] | The investigation for the reported hemolysis has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: hemolysis: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
19049 | 1220648-2024-18783 | 20340565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20340565&pc=OZD | 03/05/2021 | 09/30/2024 |  | 09/30/2024 |  |  | 08/31/2021 | 09/30/2024 |  |  | 12/31/2022 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022028949 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing aortic valvuloplasty was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient developed hemolysis during impella support and was found to have a left ventricle filling issue. treatment details were not provided. infusion was provided for hemolysis; however, no blood products was administered. no further harm to the patient was reported. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on the clinical account, the root cause of the hemolysis was due to the patient's ventricle filling issue. as noted in the impella ifu: hemolysis: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ . section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. this report is being filed as part of a retrospective review of historical records.
19050 | 1220648-2024-18785 | 20340652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20340652&pc=OZD | 12/28/2022 | 09/30/2024 |  | 09/30/2024 |  |  | 12/28/2022 | 09/30/2024 |  |  | 09/30/2024 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275722 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was noted that the patient had a tortuous vasculature which caused the cannula to kink during delivery. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of product damage (cannula kink). the device was discarded by the customer. cannula kink happened during support. the cause of the issue was a kinked cannula due to patient condition. this report is being filed as part of a retrospective review of historical records.
19051 | 1220648-2024-18784 | 20340879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20340879&pc=OZD | 12/28/2022 | 09/30/2024 |  | 09/30/2024 |  |  | 12/28/2022 | 09/30/2024 |  |  | 08/31/2024 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023261255 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, a side arm leak was observed. the impella cp was explanted and replaced with a new impella cp. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of purge leak - pump. the device was not received for evaluation. the root cause of the purge leak - pump was unable to be determined because sufficient clinical information was not provided and the product was not returned. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
19052 | 1220648-2024-18786 | 20340968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20340968&pc=OZD | 12/28/2022 | 09/30/2024 |  | 09/30/2024 |  |  | 12/28/2022 | 09/30/2024 |  |  | 08/31/2024 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270095 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while moving the patient from the table to the bed, the staff inadvertently pulled out the impella cp. the impella cp was explanted and replaced with a new impella cp. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records.
19053 | 1220648-2024-18787 | 20340970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20340970&pc=OZD | 12/28/2022 | 09/30/2024 |  | 09/30/2024 |  |  | 12/28/2022 | 09/30/2024 |  |  | 09/30/2024 |  | 09/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273564 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp exhibited low flows, and the cannula was observed to be kinked due to difficult patient anatomy. despite repositioning, the impella cp within the aorta, the catheter remained kinked, with flow rates of 1.4. the impella cp was removed and the patient was implanted with an impella 2.5. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records.
19054 | 1220648-2024-18692 | 20329474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329474&pc=OZD | 08/05/2021 | 09/28/2024 |  | 09/29/2024 |  |  | 08/06/2021 | 09/29/2024 |  |  | 10/31/2022 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021012161 |  | 0048-0034-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 21 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was replaced after the purge pressure increased. the pressure of the purge system gradually increased, and the flow rate decreased. | [5 empty] | The investigation for the reported high purge pressure issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the field lt and rt logs were reviewed, and high purge pressure and low purge flow were confirmed throughout the case. visual inspection of the returned pump revealed no obvious abnormalities associated with the purge system. the pump was plugged into a console and purged normally for approximately 30 minutes. further inspection of the pump catheter revealed a kink at the 36cm mark. stripping the catheter revealed kinked wiring indicative of catheter kink. the root cause of the high purge pressure and low purge flow was determined to be a kinked purge line during support. this report is being filed as part of a retrospective review of historical records.
19055 | 1220648-2024-18693 | 20329479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329479&pc=OZD | 08/05/2021 | 09/28/2024 |  | 09/29/2024 |  |  | 08/05/2021 | 09/29/2024 |  |  | 03/31/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022076037 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, there was a low pump flows after initiating the pump. the pump was removed, and a clot was found entrained within the device's inlet. a new pump was placed. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the field rt log was reviewed and low pump flow with suction was confirmed upon activation of the pump. the returned pump was visually inspected, and no abnormalities were observed. the pump was run in a basin of water, the pump operated to specification, and the low pump flow was unable to be reproduced. the cause of the low pump flow was most likely ingested biomaterial due to high acts as reported by the clinical account. this report is being filed as part of a retrospective review of historical records.
19056 | 1220648-2024-18695 | 20329483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329483&pc=OZD | 08/08/2021 | 09/28/2024 |  | 09/29/2024 |  |  | 08/08/2021 | 09/29/2024 |  |  | 03/31/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022061812 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a pump stop. the patient was sent to the cath lab with plans to escalate to the impella 5.0 device with fem graft. no further information provided. | [5 empty] | D6 explant date unknown. the investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that data logs from the case were not sent for investigation. visual inspection of the returned pump revealed a significantly large impeller purge gap and significant amounts of biomaterial wrapped around the tip of the impeller. no other relevant external damage or abnormalities were found on the pump or the purge cassette. after reviewing the clinical information, it was concluded that the cause of the pump stop was bearing wear after 17 days of use, which is beyond indication per design trace matrix 0046-9910. the root cause of the pump stop was determined to be bearing wear after 17 days of use, which is beyond indication per design trace matrix 0046-9910. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
19057 | 1220648-2024-18694 | 20329490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329490&pc=OZD | 08/08/2021 | 09/28/2024 |  | 09/29/2024 |  |  | 08/08/2021 | 09/29/2024 |  |  | 06/30/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022093255 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Migration or Expulsion of Device | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp came out of the left ventricle while the physician was trying to readvance under fluoro in the cath lab, and the cannula flipped onto itself. the physician was able to pull the impella down to the distal aorta and allow the device to uncurl. the decision made to replace with a new pump with concern of potential kink in cannula. a second pump was placed without issue. | [5 empty] | The investigation for the reported product damage (cannula kink) issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the cannula kink was repositioning. this report is being filed as part of a retrospective review of historical records.
19058 | 1220648-2024-18696 | 20329564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329564&pc=OZD | 09/02/2024 | 11/21/2024 |  | 09/29/2024 |  |  | 09/03/2024 | 09/29/2024 |  |  |  |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025513832 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support due to shock vital signs caused by cardiovascular angiography and interventions, myocarditis, and worsening heart failure. while on support, p-6 was being managed, but the left ventricle waveform was not appearing and the flow rate was 0. suction alarm was sounding. the a-line waveform had pulse pressure and the motor waveform was flat. the catheter was removed. after removal, there was a blood clot-like substance attached to the catheter. the device was explanted due to pump flow issue. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for thrombus and low pump flow has been completed. returned complaint device was visually inspected. no biomaterial observed. no cannula kink or broken blade found. impella connected to aic to reproduce the low pump flow issue. pump run for more than 10 minute at p-9. no low pump flow observed and it was running within specification flow range. data logs confirm low flow at p7 with a sudden drop in mc and drop flow to less than 1l/min with slight down shift in ps that recovers. the flows slightly increase but still under 1 liter and then drop again to practically 0 with mc also dropping. there are suction alarms, ps low, and position in ventricle alarms. the flows never recover. the cause of the low pump flow most likely was improper positioning issue due to improper technique. the root cause of the thrombus could not be determined without any additional clinical details. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
19059 | 1220648-2024-18700 | 20329570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329570&pc=OZD | 09/03/2024 | 10/17/2024 |  | 09/29/2024 |  |  | 09/03/2024 | 09/29/2024 |  |  |  |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510678 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella 5.5 pump to support a 73-year-old male patient admitted in with cardiomyopathy and other systemic comorbidities. the pump was placed and after two plus weeks, the pump had a malfunction. the sterile sleeve became damaged and torn. the pump remained on for support despite this and was unable to be repositioned. the pump was explanted five days later and the patient survived the explant. | [5 empty] | The investigation for the product sleeve damage has been completed. no product was returned for analysis. based on the clinical details provided, the root cause of the product damage (sterile sleeve damage) is due to use. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings & cautions: cautions ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
19060 | 1220648-2024-18701 | 20329572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329572&pc=OZD | 09/04/2024 | 10/21/2024 |  | 09/29/2024 |  |  | 09/04/2024 | 09/29/2024 |  |  |  |  | 09/29/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1740558 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when the automated impella controller (aic) was turned on, purge cassette was primed and white impella cable was connected. once the cable was connected, the purge did not begin priming and the critical alarm appeared on the aic stating ¿controller failure¿, with instructions to ensure connection between the red impella plug and grey connectors. aic was exchanged and the pump was reconnected. no further issues with the new aic were noted. there was no patient involvement. | [5 empty] | The investigation for the console (aic) controller failure-red alarm has been completed. console logs from the reported day of event are consistent with complaint and show large number of error messages associated with eep/epm malfunction, culminated in controller failure alarm (#1029, due to epm firmware corruption). analysis of logs data from before the reported event revealed that sau_sens_eep_1 issues started while the console was in preventative maintenance but no issues were observed, since alarm trigger criteria had not been met. the console was returned for evaluation and the issue was reproduced - after seemingly correct boot-up, console logs were being flooded with sau_sens_eep_1 error messages. after extracting epm/eep firmware from its sub-circuitry on the impellatronic board and comparing to correct programming code, a memory corruption was found. after re-flashing epm firmware and restarting the console, the issues no longer reproduced. the cause of the aic issue was a firmware corruption. added- d9 device return date correction to the initial mfr report: h6 impact code 2645 changed to 4629 h8 usage of device changed to reuse. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
19061 | 1220648-2024-18702 | 20329582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329582&pc=OZD | 09/01/2024 | 10/25/2024 |  | 09/29/2024 |  |  | 09/01/2024 | 09/29/2024 |  |  |  |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530676 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, bleeding was noted at impella site and controlled with manual pressure and multiple dressing changes. forward tension sutures were added as it appeared the patient¿s abdomen was causing the sheath to shift. the pump continued to piston, but the doctor declined to reposition. two units of blood were given because of the retroperitoneal bleed. it was also reported that the patient had an unknown mass in left ventricle. when the impella was removed the thrombus was removed with an alpha vac. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. it was noted that adding a forward tension suture greatly improved the bleeding. the root cause of the access site bleeding was most likely use-related. the root cause of the thrombus could not be determined without additional clinical details. the root cause of the vascular damage was not determined since insufficient clinical details were provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
19062 | 1220648-2024-18697 | 20329588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329588&pc=OZD | 09/03/2024 | 02/25/2025 |  | 09/29/2024 |  |  | 09/03/2024 | 09/29/2024 |  |  |  |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1664272 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) had a controller failure. this issue was discovered during preparation. | [5 empty] | The investigation for the automated impella controller (aic) controller error - yellow alarm has been completed. data logs confirmed that the controller error (defective optical sensor lamp) - alarm was issued. the console was returned and was investigated in the lab. preventative maintenance procedure was performed and the console found to be within specification. the console was able to run a well known pump and purge. the root cause for the defective optical sensor lamp alarm could not be determined since the failure could not be reproduced. d.2 revised common device name since the initial manufacturer device report number 1220648-2024-18697 was submitted in accordance with updated procedures. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-18697 was submitted in accordance with updated procedures. | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
19063 | 1220648-2024-18704 | 20329589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329589&pc=OZD | 09/03/2024 | 10/30/2024 |  | 09/29/2024 |  |  | 09/03/2024 | 09/29/2024 |  |  |  |  | 09/29/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1382144 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a battery failure. the team attempted to leave the aic plugged in for extended amounts of time and to work through making changes to switch the battery and attempts were unsuccessful. | [5 empty] | The investigation for the automated impella controller (aic) - battery failure (red alarm) has been completed. data logs were reviewed and the reported battery failure issue was confirmed. there were numerous instances of transport connection blown/missing alarms observed in the logs from the reported occurrence date. it was noted that the issue occurred during the consoles first boot after preventive maintenance (pm) was performed by field service. the console was returned for evaluation. the reported battery failure alarm issue was unable to be reproduced. results of running the batttest utility on this console revealed that the health of the batteries were normal. the reported battery failure issue was determined to be use related. it is likely that the aic was booted up while it was connected to alternating current power and the battery switch was disengaged since the issue occurred upon first boot of the console after routine pm was performed on it. the root cause of this issue was likely the failure to engage the battery transport switch before booting up the aic. | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
19064 | 1220648-2024-18699 | 20329593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329593&pc=OZD | 09/03/2024 | 11/22/2024 |  | 09/29/2024 |  |  | 09/03/2024 | 09/29/2024 |  |  | 10/31/2025 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445167 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after implanting and starting on impella cp, the 14 french (fr) x 25 centimeter (cm) peel away sheath was noted to be leaking blood at the valve. the impella was pulled and the sheath was replaced with a 14 fr x 13 cm sheath. the impella was rinsed in sterile five percent dextrose water and reinserted. the new sheath was not leaking and the case continued successfully. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation into introducer damage ¿ mechanical has been completed. the device analysis showed there were no damages were observed.
19065 | 1220648-2024-18705 | 20329595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329595&pc=OZD | 09/02/2024 | 11/21/2024 |  | 09/29/2024 |  |  | 09/04/2024 | 09/29/2024 |  |  |  |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521442 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 17 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 17-year-old female patient was on impella cp support and during support, the pump was repositioned and there was no visible flow. the pump was explanted and replaced with a new impella cp. there was no reported patient harm. the patient survived the expalnt. | [5 empty] | The investigation for the low pump flow has been completed. product and data logs were not returned for review. the root cause of the low flow was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of the pump stop was unable to be determined as limited clinical details were provided and no product was returned for review. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi states the following: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿. | A.5 is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
19066 | 1220648-2024-18698 | 20329604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329604&pc=OZD | 09/03/2024 | 10/15/2025 |  | 09/29/2024 |  |  | 09/03/2024 | 09/29/2024 |  |  |  |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512716 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a 73-year-old male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, the impella cp sensor was reading in aorta with flat motor current but on fluoroscopy the impella was noted across the valve. the impella was repositioned and was noted under the papillary muscle. after percutaneous coronary intervention was completed, the impella cp was removed. | [5 empty] | The investigation for the positioning issue has been completed. no product was returned. logs showed "impella in aorta" alarms after pump was turned down to p-4. the pump was explanted about 30 minutes after the first positioning alarm. there was a large motor current spike about twenty minutes prior to first positioning alarm. there were no abnormalities with the 5s parameters. the pump was noted under the papillary muscle. the root cause of the positioning issues was most likely use-related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the root cause of the optical signal issue was most likely biomaterial. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
19067 | 1220648-2024-18703 | 20329606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329606&pc=OZD | 09/03/2024 | 11/07/2024 |  | 09/29/2024 |  |  | 09/03/2024 | 09/29/2024 |  |  |  |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456623 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support due to non-st-segment elevation myocardial infarction. while on support, there was decreased flows and position alarms. the pump was repositioned. additionally, the patient experienced significant ectopy including several sustained runs of ventricular tachycardia. the echocardiogram showed the impella against the lateral wall and multiple attempts were made to optimize positioning with the final decision to secure it with the left ventricle measuring at 2.9 centimeters free of ectopy on p-7 with 3.1 liters per minute. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation has been completed. the pump was not returned for investigation. data logs showed several suction and impella position in aorta/ventricle alarms at the start of the case, with noted low pump flow. data logs confirmed good pump flow later during the case run without any alarms. as per the clinical information, adjustments were made to the impella position, and pump flow improved. several sustained runs of ventricular tachycardia (vt) were experienced during the case run. the cause of the low pump flow issue was improper positioning of the device, as verified by clinical review. data logs also confirm the flow improvement following repositioning. the positioning issue failure mode was removed, as it was addressed as a cascading event under low pump flow due to improper technique. as the necessary information was not provided, the root cause of the vt was not determined.
19068 | 1220648-2024-18706 | 20329988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329988&pc=OZD | 08/10/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/10/2021 | 09/29/2024 |  |  | 05/31/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089167 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, there were low pump flows after initiating the pump. the cannula was observed to be kinked and the was kinked after crossing the valve. the kinked cannula was just distal to the motor housing. the pump was removed and replaced with an impella 2.5. | [5 empty] | The investigation for the reported low pump flow and product damage (cannula kink)issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a kink could be seen in the cannula. the pump was run without the kink and performed within specification. the root cause of the low flows was determined to be a kink in the cannula and the root cause of the cannula kink was determined to be a result of delivery issues through the patient¿s tortuous vasculature. this report is being filed as part of a retrospective review of historical records.
19069 | 1220648-2024-18707 | 20329989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329989&pc=OZD | 08/10/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/10/2021 | 09/29/2024 |  |  | 03/31/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022064817 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a purge pressure high alarm followed by pump failure and was unable to restart the pump. it was indicated that the impella was supposed to be explanted the following day and that the patient was stable. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
19070 | 1220648-2024-18708 | 20329993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329993&pc=OZD | 08/11/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/11/2021 | 09/29/2024 |  |  | 06/30/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022101574 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. shortly after implantation, it was noted that blood was rapidly dripping out of the red impella plug on the cp catheter. the pump was removed and replaced with a second cp. no patient harm or injury noted. | [5 empty] | The investigation for the reported the access site bleeding ¿ major and product damage (hole in catheter) issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the pump handle was filled with blood. a small cut in the catheter was noted at the 85cm mark. therefore, the root cause of the blood in the pump handle was the small cut in the catheter which resembled a puncture that occurred during use. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
19071 | 1220648-2024-18709 | 20329997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329997&pc=OZD | 08/16/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/16/2021 | 09/29/2024 |  |  | 01/31/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022039743 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the color tone of the impella is slightly worse than that of the other side. the impella was replaced with intra-aortic balloon pump (iabp) in consideration of coronary artery bypass grafting (cabg) implementation. | [5 empty] | The investigation for the reported limb ischemia issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. information indicated that the patient¿s iliac vessels were thin and there was potentially calcification. therefore, the cause of the reported limb ischemia was most likely related to the condition/size of the patient vessel. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
19072 | 1220648-2024-18710 | 20330003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330003&pc=OZD | 08/16/2021 | 02/27/2025 |  | 09/29/2024 |  |  | 08/16/2021 | 09/29/2024 |  |  | 06/30/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022101229 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump had low flows, and at the sheath peel away, clot burden was observed. the team replaced the pump and support proceeded without harm or injury from interruption. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump was visually inspected, and biomaterial was observed on the impeller. the cause of the low pump flow was ingested biomaterial due to low patient activated clotting time (acts). this report is being filed as part of a retrospective review of historical records. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
19073 | 1220648-2024-18711 | 20330004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330004&pc=OZD | 08/17/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/17/2021 | 09/29/2024 |  |  | 04/30/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079824 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during support, the pump stopped unexpectedly. the physician lowered the support level, restarted the pump, and replaced the control device, but it still did not function. the resolution of the issue is unknown, but no patient harm was reported. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine that the root cause of the pump stop was most likely due to bearing wear after 23 days of use, which is beyond the device's indicated period of use. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella; rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
19074 | 1220648-2024-18712 | 20330005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330005&pc=OZD | 08/18/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/18/2021 | 09/29/2024 |  |  | 06/30/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2022097299 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Female |  |  |  | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during placement, the pigtail/easy guide lumen broke outside of the patient. a new impella pump was used, and there was no harm to the patient. | [5 empty] | The device was returned, and the investigation has been completed. the red lumen was not returned, but per the clinical account, the root cause of the case start issue was a fractured red lumen prior to insertion. as noted in the impella ifu: cannula/pigtail detach great vessel/heart: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings, cautions, and precautions ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ impella cp with smartassist for use during cardiogenic shock and high risk pci and impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ impella rp smart assist system with automated impella controller section: warnings & cautions: warnings ¿handle with care. the impella rp with smartassist system catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings ¿handle with care. the impella rp flex with smartassist system catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records.
19075 | 1220648-2024-18713 | 20330006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330006&pc=OZD | 08/18/2021 | 10/09/2025 |  | 09/29/2024 |  |  | 08/18/2021 | 09/29/2024 |  |  | 03/31/2023 |  | 09/29/2024 |  | IMPELLA CP EU |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2022066355 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a kinked cannula prior to initiating the device. the pump was removed and replaced without and adverse impact. | The complainant reported a patient with unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after implant, the device developed a kinked cannula prior to initiating the device. the pump was removed and replaced without patient harm. |  |  |  |  | This report is being filed as part of a retrospective review of historical records.the investigation into the reported event is complete. a kink in the cannula was confirmed, the root cause of the low pump flow was due to a cannula kink. g1: updated manf. email address. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
19076 | 1220648-2024-18714 | 20330007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330007&pc=OZD | 01/04/2023 | 09/29/2024 |  | 09/29/2024 |  |  | 01/04/2023 | 09/29/2024 |  |  | 08/31/2024 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262325 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention attempted to being implanted with an impella cp device for mechanical circulatory support. the impella cp pump could not be fully advanced into the patient because of calcified vasculature. therefore, the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation into the delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. it was determined that the cause of the delivery issues was patient condition related, specifically the patient's calcified vessels. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
19077 | 1220648-2024-18715 | 20330009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330009&pc=OZD | 08/18/2021 | 02/28/2025 |  | 09/29/2024 |  |  | 08/18/2021 | 09/29/2024 |  |  | 06/30/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095347 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella unexpectedly alarmed "motor current high" and the pump stopped. the pump automatically attempted to restart, but the motor current spiked again. the device briefly was able to be started at p-2, but the motor current was very high, and it would stop again after just a few moments. patient rapidly was decompensating and the impella was withdrawn into the aorta and drips were rapidly titrated up to stabilize patient. the pump was replaced, and there was no harm to the patient. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device found to have a large purge gap indicating the bearing had failed. the cause of the pump stop was bearing wear. the pump stop occurred beyond the 8 day impella cp indication for use. this report is being filed as part of a retrospective review of historical records.
19078 | 1220648-2024-18716 | 20330010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330010&pc=OZD | 01/04/2023 | 09/29/2024 |  | 09/29/2024 |  |  | 01/04/2023 | 09/29/2024 |  |  | 04/30/2024 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219214 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to being implanted with an impella cp device for mechanical circulatory support. the pump failed to pass through the iliac during the initial delivery attempt. the patient was subsequently prepped for an axillary access. however, their condition improved and the implant was aborted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the impella device was not discarded by the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the impella was attempted to be implanted without success. the root cause of the delivery issues was patient condition because of the patient's unknown peripheral vascular disease. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
19079 | 1220648-2024-18717 | 20330012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330012&pc=OZD | 01/06/2023 | 09/29/2024 |  | 09/29/2024 |  |  | 01/06/2023 | 09/29/2024 |  |  | 09/30/2024 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275678 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump exhibited high purge pressure. to address this issue tissue plasminogen activator was started, but purge pressure had not resolved. additionally, the pump stopped and was unable to restart. the impella cp device was explanted. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
19080 | 1220648-2024-18718 | 20330025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330025&pc=OZD | 01/06/2023 | 09/29/2024 |  | 09/29/2024 |  |  | 01/06/2023 | 09/29/2024 |  |  | 08/31/2024 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263957 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to being implanted with an impella cp device for mechanical circulatory support. the complainant reported a failure to deliver the device due to patient's anatomy and potential damage to the pigtail. as a result of this complication the insertion was aborted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the pump and guidewire were returned and visually inspected. no problem or damaged spot was observed in the impella pump. no problem found at the guide wire (gw) ends. the gw passed through the pigtail and no obstruct or resistance observed. based on the clinical review, the root cause was patient condition due to calcifications. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
19081 | 1220648-2024-18719 | 20330026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330026&pc=OZD | 08/19/2021 | 03/27/2025 |  | 09/29/2024 |  |  | 08/20/2021 | 09/29/2024 |  |  | 06/30/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022103753 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attmpted to be implanted with an impella cp device for mechanical circulatory support. during insertion, the impella tip and outflow made it through the impella sheath but would not advance past the aorto-iliac bifurcation due to the patient's diseased vasculature. the device was removed, and the procedure was completed successfully without impella support. there was no harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the root cause of the delivery issues was determined to be patient condition as the patient reportedly has severe peripheral arterial disease (pad). this report is being filed as part of a retrospective review of historical records. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
19082 | 1220648-2024-18720 | 20330110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330110&pc=OZD | 08/19/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/20/2021 | 09/29/2024 |  |  | 05/31/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022087400 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. shortly after the pump was inserted and started, the screen showed a "placement signal unreliable" alarm, and no placement signal was visible on the console. the physician then replaced the impella pump and console, which worked successfully. there was no harm to the patient. | [5 empty] | The investigation for the reported placement signal issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
19083 | 1220648-2024-18722 | 20330117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330117&pc=OZD | 08/20/2021 | 09/24/2025 |  | 09/29/2024 |  |  | 08/20/2021 | 09/29/2024 |  |  | 10/31/2022 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022042364 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there was a regular increase of purge flow up to 12ml/h, and in the later course the flow settled at 6.9ml/h. the device was explanted and replaced, and there was no harm to the patient. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the high purge pressure has been completed. only the purge cassette (pc) was returned. no leaks were found with the pc, it was connected to a pump and operated within specification. based on the data log analysis, the purge flow / purge pressure (pf/pp )remains within specification, around 600 mmhg, for the duration of the case. no returned product images of the cassette were saved. only the log analysis. the cause of the issue was not determined since the issue was unable to be reproduced with the returned cassette and the pump was not returned. corrections to the initial mfr report: g1 mdr reporting contact email. g1 mfr site name and address updated. h6 type of investigation code 10 , investigation conclusion code 4315 added.
19084 | 1220648-2024-18723 | 20330123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330123&pc=OZD | 08/20/2021 | 09/29/2025 |  | 09/29/2024 |  |  | 08/20/2021 | 09/29/2024 |  |  | 06/30/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022102525 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. during delivery, the cannula and catheter kinked due to patient stenosis. this was treated via a balloon aortic valvuloplasty, and a second impella was placed without further difficulty. there was no reported harm to the patient. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation has been completed. no product was returned for investigation. after reviewing the clinical information, it was determined that the cause of the delivery issues was patient condition, specifically aortic valve stenosis. this report is being filed as part of a retrospective review of historical records. d1 brand name was updated as it was erroneously entered on the initial manufacturer device report number. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures.
19085 | 1220648-2024-18724 | 20330148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330148&pc=OZD | 08/22/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/22/2021 | 09/29/2024 |  |  | 04/30/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080427 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported resistance when attempting to remove the easy guide lumen, during preparation of the pump. the easy guide lumen was removed, but this resulted in damage to the 0.018" guide wire. a new pump was opened and used without further issue, and there was no harm to the patient. | [5 empty] | The device was returned, and the investigation for the reported difficulty inserting the pump through the introducer. the returned pump was visually inspected, and a fractured red lumen was confirmed. this was attributed to user error. as noted in the impella ifu: difficulty inserting/removing pump through introducer: impella 5.5 & cp impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during impella section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
19086 | 1220648-2024-18725 | 20330192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330192&pc=OZD | 08/23/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/23/2021 | 09/29/2024 |  |  | 05/31/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095047 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the placement signal (ps) was not reliable during 90% of the case. however, the procedure was completed successfully with no harm to the patient. | [5 empty] | The device was returned, and the investigation for the reported placement signal issue was completed. data logs showed that the "placement signal not reliable" alarm occurred shortly after pump activation and remained for the entirety of support. placement signal metrics were constant at 3000 mmhg, while optical sensor metrics were -30000 at the time. visual inspection of the returned pump revealed the presence of biomaterial covering the optical sensor. the biomaterial was cleaned off and then the pump was run. no placement signal alarms occurred. therefore, it was determined that the cause of the optical placement signal alarm issue was patient condition, specifically biomaterial found on the optical sensor. this report is being filed as part of a retrospective review of historical records.
19087 | 1220648-2024-18726 | 20330216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330216&pc=OZD | 08/23/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/23/2021 | 09/29/2024 |  |  | 02/28/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022047443 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. during the first series of ultrasound pulses, the placement signal appeared to spike straight up and then disappeared. the aortic (ao) and left ventricle (lv) placement signals showed ¿-/¿-. no additional information related to the resolution was provided. there was no harm to the patient. | [5 empty] | The device was returned, and the investigation for the reported placement signal issue has been completed. investigation: the device was returned for review and the optical system was evaluated. when attempting to observe the fringe pattern on the ccd, no signal was received. this indicates that the sensor had catastrophically failed and fractured. the cause of the placement signal failure was most likely an interaction with another device. this report is being filed as part of a retrospective review of historical records.
19088 | 1220648-2024-18727 | 20330248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330248&pc=OZD | 08/23/2021 | 03/13/2025 |  | 09/29/2024 |  |  | 02/08/2022 | 09/29/2024 |  |  | 01/31/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022057183 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced a cerebral hemorrhage. the impella was explanted the same day, though it is not specified whether this was related to the hemorrhage. the patient survived post-explant. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: cva/ischemic/hemorrhagic: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
19089 | 1220648-2024-18728 | 20330255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330255&pc=OZD | 08/24/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/24/2021 | 09/29/2024 |  |  | 02/28/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022055728 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. after explant, the physician suspected hemolysis and asked whether hemolysis could be caused by the catheter. there was no report of harm or medical intervention during this event. | [5 empty] | The device was returned, and the investigation for the reported hemolysis has been completed. investigation: analysis of the returned product revealed a large clot wrapped around the impeller, likely contributing to the suspected hemolysis. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: hemolysis: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
19090 | 1220648-2024-18729 | 20330286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330286&pc=OZD | 08/24/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/24/2021 | 09/29/2024 |  |  | 05/31/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022091204 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, there was leaking of d5 from around the red impella plug. the percutaneous intervention (pci) was able to be performed. no patient harm was reported. | [5 empty] | The device was returned, and the investigation for the reported purge leak has been completed. investigation: evaluation of the returned pump found no leaks in the purge system. the returned pump was purged without issue. the root cause of the low purge pressure was most likely due to use, as no leaks were found in the purge system and the pump purged without alarms. as noted in the impella ifu: purge leak - pump: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
19091 | 1220648-2024-18730 | 20330303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330303&pc=OZD | 08/24/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/24/2021 | 09/29/2024 |  |  | 03/31/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022069171 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon starting the pump, it produced a "placement signal unreliable" alarm and dashed out. the decision was made to remove and replace the pump, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported placement signal issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that no logs were provided for review. the returned pump was run in the lab and the 'placement signal unreliable' alarms were unable to be reproduced. electrical testing found no open lines. no broken optical fibers were noted when using laser to test. no debris was found on the optical ferrule of the patient cable. analysis of the pump via the optical bench found lamp to be 64.55%, light 2.70 v, cont. 15.23%, gain 1.34, and snr 3135.3. these 5s values indicated that there were no optical issues with the pump. this showed that root cause of the placement signal issue was most likely use related. this report is being filed as part of a retrospective review of historical records.
19092 | 1220648-2024-18731 | 20330320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330320&pc=OZD | 08/23/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/24/2021 | 09/29/2024 |  |  | 05/31/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083555 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient bled from the access site due to physical movement. the attending physician told the patient not to move the impella side, but since the patient had moved, it is probable that this was due to physical movement. an additional skin suture was placed to secure the access site, and there was no further patient harm reported. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding was determined to be patient movement as the physician stated that the bleed was a result of the patient moving. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
19093 | 1220648-2024-18732 | 20330368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330368&pc=OZD | 08/26/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/26/2021 | 09/29/2024 |  |  | 06/30/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022106405 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the guidewire fell out of the ventricle during delivery of the impella. the pump and wire were removed and a new pump and wire were delivered without further issue. there was no adverse impact to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
19094 | 1220648-2024-18733 | 20330387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330387&pc=OZD | 08/26/2021 | 09/30/2025 |  | 09/29/2024 |  |  | 08/26/2021 | 09/29/2024 |  |  | 06/30/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105417 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, after a minor repositioning, the pump flow dropped from 4l/min to 1.4l/min. the pump was replaced due to concern for clot, and no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation has been completed. the product was returned and biomaterial could be seen surrounding the impeller. the pump produced high motor current pump stops until the biomaterial was successfully ejected. once the biomaterial was removed, the pump's flow level was within specification. the root cause of the low pump flow was determined to be biomaterial ingestion as their was a sudden decrease in flow reported and biomaterial was found impeding the flow path. this report is being filed as part of a retrospective review of historical records. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures.
19095 | 1220648-2024-18734 | 20330398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330398&pc=OZD | 08/29/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 08/29/2021 | 09/29/2024 |  |  | 05/31/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022084113 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there was excessive suction and reported flow saturations. the physician decided to wean and remove the pump, and if patient didn¿t tolerate, he would reinsert a new cp. there was no reported harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
19096 | 1220648-2024-18735 | 20330412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330412&pc=OZD | 08/30/2021 | 03/27/2025 |  | 09/29/2024 |  |  | 08/30/2021 | 09/29/2024 |  |  | 06/30/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095608 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump began to show suction at p8. the physician repositioned the device multiple times and began to drop the p level, but even at p2 the device continued to show suction. the pump was then removed and replaced, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the pump passed all the post sterile inspection checks. the evaluation of the device noted that the field real time (rt) log was reviewed, and suction alarms and low pump flows were confirmed. the returned pump was visually inspected, and biomaterial was observed on the impeller. the pump was plugged into a console and ran at p-9 with low pump flows and suction alarms reproduced. upon removal of the biomaterial, the pump functioned normally with no alarms. the cause of the reported suction alarms was the observed ingested biomaterial. this report is being filed as part of a retrospective review of historical records. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
19097 | 1220648-2024-18736 | 20330420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330420&pc=OZD | 08/30/2021 | 09/29/2024 |  | 09/29/2024 |  |  | 01/06/2022 | 09/29/2024 |  |  | 01/31/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022039731 |  | 0048-0034-JP |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis for the duration the device was implanted. the patient's hemodynamics were stable, so the patient was weaned. | [5 empty] | The investigation for the reported hemolysis issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
19098 | 1220648-2024-18737 | 20330441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330441&pc=OZD | 01/06/2023 | 09/29/2024 |  | 09/29/2024 |  |  | 01/06/2023 | 09/29/2024 |  |  | 09/30/2024 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273866 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to being implanted with an impella cp device for mechanical circulatory support. a 14 french x !3 cenimeter sheath was inserted. the pump was attempted to advance but did not cross the common iliac due to the tortuosity of the artery coupled with calcium build up. a 13 centimeter sheath was exchanged for a 25 centimeter sheath over a stiff wire due to the possibility that it might cross the lesion and provide track for the catheter. the impella device was attempted to be advanced for a second time without success. the physician did not feel comfortable in proceeding. the case was aborted. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was due to the pump not being able to cross the common iliac due to the patient's tortious condition. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
19099 | 1220648-2024-18738 | 20330445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330445&pc=OZD | 01/07/2023 | 09/29/2024 |  | 09/29/2024 |  |  | 01/07/2023 | 09/29/2024 |  |  | 08/31/2023 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127881 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during setup, the pump was started outside of the body with the easy guide lumen still in place. an impella stopped alarm shortly followed. the impella cp was explanted and replaced to address the issue. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
19100 | 1220648-2024-18739 | 20330450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330450&pc=OZD | 01/07/2023 | 09/29/2024 |  | 09/29/2024 |  |  | 01/07/2023 | 09/29/2024 |  |  | 09/30/2024 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278881 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump was having continuous suction alarms. consequently, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was returned for evaluation. a large thrombus was found around the impeller. the root cause of the low pump flow was determined to be ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
19101 | 1220648-2024-18740 | 20330454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330454&pc=OZD | 01/08/2023 | 09/29/2024 |  | 09/29/2024 |  |  | 01/08/2023 | 09/29/2024 |  |  | 07/31/2024 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023258971 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, a pump stop occurred. normal troubleshooting attempts to restart the device were unsuccessful. the impella cp was explanted to address the issue. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
19102 | 1220648-2024-18741 | 20330460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20330460&pc=OZD | 01/09/2023 | 09/29/2024 |  | 09/29/2024 |  |  | 01/09/2023 | 09/29/2024 |  |  | 09/30/2024 |  | 09/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023278554 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, saturated flows were observed. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of saturated flow. the device was returned for evaluation. data logs showed flow saturation began to occur immediately following a sudden motor current spike (accompanied by a motor speed drop) about two days into support. the flow saturation persists even after the pump speed was turned down for the remainder of support, with max flows at 3 liter per minute at p-2. visual inspection revealed remnants of biomaterial wrapped around the impeller and larger amount wrapped at the base of the motor housing. no other abnormalities were found. in conclusion, it was determined that the cause of the saturated flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
19103 | 1220648-2024-18665 | 20328939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20328939&pc=OZD | 07/02/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/02/2021 | 09/28/2024 |  |  | 04/30/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2022081132 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had bleeding at the access site. a blood transfusion was required. the impella was explanted. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the reported access site bleeding was most likely due to lack of vessel recoil. additional information indicated that the patient¿s blood vessel was dilated. best practices were followed and bleeding resolved when impella was removed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
19104 | 1220648-2024-18666 | 20328943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20328943&pc=OZD | 09/14/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 09/14/2021 | 09/28/2024 |  |  | 05/30/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2022079789 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed ventricular tachycardia and hemolysis. however, no additional details were provided regarding the treatment or management of these conditions. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: arrhythmia section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ hemolysis impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
19105 | 1220648-2024-18667 | 20328999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20328999&pc=OZD | 07/02/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/02/2021 | 09/28/2024 |  |  | 03/31/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022067125 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the cannula kinked just past the motor while in position. pulling back and advancing did not correct the problem. suction developed and reduced flows. the impella was explanted and another pump was implanted. | [5 empty] | The investigation for the reported cannula kink has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the cannula kink was due to repositioning of the device. this report is being filed as part of a retrospective review of historical records.
19106 | 1220648-2024-18668 | 20329006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329006&pc=OZD | 07/03/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/03/2021 | 09/28/2024 |  |  | 12/31/2022 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022063716 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was needing an impella cp device for mechanical circulatory support. during implantation the impella was not able to complete priming. the initial device was removed and a new impella was implanted. | [5 empty] | The investigation for the reported priming issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was completed. the pump went through case start and priming without any issues. there was no problem with the device found. the root cause of the priming issue could not be determined. this report is being filed as part of a retrospective review of historical records.
19107 | 1220648-2024-18669 | 20329007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329007&pc=OZD | 07/03/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/04/2021 | 09/28/2024 |  |  | 05/31/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022086252 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. while on support, the purge pressure dropped to zero with purge flow at 30. the staff reported that there was a leak at the yellow luer connection. as the patient was hemodynamically stable, the impella was explanted at bedside. | [5 empty] | The investigation for the reported pump purge leak issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device confirmed a luer leak. the root cause of the purge leak was the sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records.
19108 | 1220648-2024-18670 | 20329008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329008&pc=OZD | 07/05/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/05/2021 | 09/28/2024 |  |  | 10/31/2022 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2021019881 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. upon explant, a clot was observed adhered to the pump. no additional details were provided regarding the treatment or management of this condition. | [5 empty] | The investigation for the reported thrombus issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, information provided was sufficient to determine a root cause. the root cause of the delivery issues was biomaterial which was observed on the pump at the time of pump removal and was most likely related to patient condition. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
19109 | 1220648-2024-18671 | 20329024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329024&pc=OZD | 07/05/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/05/2021 | 09/28/2024 |  |  | 04/30/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022075335 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock needed an impella cp device for mechanical circulatory support. it was reported that the white connector cable was plugged into three (3) different controllers and got "controller error" alarm. the original pump was never inserted into the patient and a new pump was opened and successfully implanted. | [5 empty] | The investigation for the reported pump not detected issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device confirmed insufficient adhesive. the root cause of the of the case start issue was an open fiber line due to insufficient adhesive. the pump sensor passed manufacturing testing per lot traveler review. this report is being filed as part of a retrospective review of historical records.
19110 | 1220648-2024-18672 | 20329032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329032&pc=OZD | 07/05/2021 | 05/22/2025 |  | 09/28/2024 |  |  | 11/16/2021 | 09/28/2024 |  |  | 03/31/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 2022061922 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient developed hemolysis. there was no information provided on how this was treated. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19111 | 1220648-2024-18673 | 20329048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329048&pc=OZD | 07/06/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/06/2021 | 09/28/2024 |  |  | 06/30/2022 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1486359 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 25 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, a pump stop occurred during use of the device. the pump was removed and replaced with no adverse impact to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. no clinical details provided were sufficient to determine a root cause. this report is being filed as part of a retrospective review of historical records.
19112 | 1220648-2024-18674 | 20329084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329084&pc=OZD | 07/06/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/06/2021 | 09/28/2024 |  |  | 02/28/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022048772 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. due to groin site bleeding the patient was taken back to the cardiac catheterization lab where the reposition sheath was wired. the initial impella was weaned and removed. a 14 fr oscor sheath was placed for hemostasis in preparation for the new pump insertion. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the access site bleeding was not determined since product was not returned. instructions for use for the related event are as follows: access site bleeding. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
19113 | 1220648-2024-18675 | 20329086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329086&pc=OZD | 07/07/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/07/2021 | 09/28/2024 |  |  | 05/31/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083422 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a pump stop after 39 hours of support. the team attempted to troubleshoot by console replacement but, the pump did not restart. no lasting harm or injury known. the team replaced the pump and support proceeded. | [5 empty] | The impella device was not received from the customer, therefore, evaluation of the device was not possible. the clinical team reported that the impella stopped 39 hours into support. the clinicians noted an increasing motor current and purge pressure prior to the pump stop. the pump was connected to another aic in an attempt to troubleshoot the issue, but this was unsuccessful. all products were discarded after use. the root cause of the pump stop was not determined. this report is being filed as part of a retrospective review of historical records.
19114 | 1220648-2024-18676 | 20329102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329102&pc=OZD | 07/07/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/07/2021 | 09/28/2024 |  |  | 04/30/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022070268 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a kink in the cannula just below the outlet that was noted on fluoroscopy. the catheter seemed to spring back in the left ventricle. impella was advanced and kink resolved for a few moments but there was a noticeable kink noted again with impella retreating back. the impella was explanted and replaced with another impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. the cause of the cannula kink was likely use-related because the second implant went without issue. this report is being filed as part of a retrospective review of historical records.
19115 | 1220648-2024-18677 | 20329103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329103&pc=OZD | 07/07/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/07/2021 | 09/28/2024 |  |  | 05/31/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083680 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there was access site was bleeding and a hematoma around the impella sheath. vascular surgery had been consulted and it was recommended to remove the impella sheath and replace it with a 18f gore sheath to control the bleeding. | [5 empty] | The impella device was not received from the customer, therefore, evaluation of the device was not possible. the cause of the access site bleeding was a result of the guide wire being unable to be placed in the arteriotomy as the patient's tissue tract was too deep. root cause is due to patient anatomy. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
19116 | 1220648-2024-18678 | 20329113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329113&pc=OZD | 07/08/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/08/2021 | 09/28/2024 |  |  | 03/31/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022068674 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp failed to start on 2 different controllers. there was an impella controller error. the pump was replaced and support proceeded with new pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. any optical placement signal issue. excludes fluidic or differential placement signal issues and positioning issues. without hemodynamic compromise. the cause of the case start issue was not determined. this report is being filed as part of a retrospective review of historical records.
19117 | 1220648-2024-18679 | 20329114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329114&pc=OZD | 07/09/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/09/2021 | 09/28/2024 |  |  | 04/30/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022075335 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump¿s placement signal was failing. the white connector cable was reseated and the issue resolved without any harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. controller error alert on the pump, lost it's placement signal for a brief amount of time. the primary rn checked all connections & "wiggled" the white connector cable into the controller & the placement signal came back & the alert disappeared. no harm to the patient. the placement signal issue was most likely use related since the alarm resolved after wiggling the patient cable connector. the product was not returned to confirm. this report is being filed as part of a retrospective review of historical records.
19118 | 1220648-2024-18680 | 20329123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329123&pc=OZD | 07/10/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/10/2021 | 09/28/2024 |  |  | 05/31/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022088927 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, patient experienced low pump flow. the team watched the flows drop from 3.5 to 2.8 which prompted pump replacement. | [5 empty] | The investigation into the restricted flow rate has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the reported issue could not be determined. the cause of the restricted flow rate was not determined since neither the product nor the data logs were returned. this report is being filed as part of a retrospective review of historical records.
19119 | 1220648-2024-18681 | 20329140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329140&pc=OZD | 07/20/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/20/2021 | 09/28/2024 |  |  | 11/30/2022 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022032735 |  | 0048-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the purge system was blocked, and the patient had hemolysis. however, no additional details were provided regarding the treatment or management of this condition. the patient was successfully weaned off of support. | [5 empty] | The investigation for the reported high purge pressure issue has been completed. the impella device was received from the customer. the likely root cause of the high purge pressure was ingested biomaterial occluding the purge gap. this report is being filed as part of a retrospective review of historical records.
19120 | 1220648-2024-18682 | 20329166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329166&pc=OZD | 07/22/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/22/2021 | 09/28/2024 |  |  | 06/30/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022093184 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Device Damaged by Another Device | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | SHOCKWAVE | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, pump was working fine and md proceeded with lazer, rotoblator and shockwave. immediately after first shockwave, placement and lv signal went off screen and numbers dashed out, re-adjusted with out success of brining signals back. upon view of fluoroscopy it was noted that the guide catheter for intervention was right up against the cannula. but when md moved guide catheter a little, placement signal still did not come back. procedure was completed without incident. | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device was received from the customer. the placement signal failure was most likely caused by damage to the sensor through an interaction with another device. this report is being filed as part of a retrospective review of historical records.
19121 | 1220648-2024-18683 | 20329231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329231&pc=OZD | 07/24/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/24/2021 | 09/28/2024 |  |  | 01/31/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022040859 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump had high motor current (mc) spike, stopped, restarted at p2, and there was no mention of patient harm or medical intervention. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer, nor was clinical information provided sufficient to determine the root cause. therefore, investigation of the device was not possible. a root cause for the pump stop could not be determined through this investigation as no product was returned for evaluation. this report is being filed as part of a retrospective review of historical records.
19122 | 1220648-2024-18684 | 20329257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329257&pc=OZD | 07/26/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/26/2021 | 09/28/2024 |  |  | 05/31/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022087788 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported the impella cp failed to pass through the introducer and to the left ventricle (lv) due to iliac artery disease. the pump was found to have kinked. the product was explanted and the 2nd pump successfully placed. no harm or injury was noted. additional information provided that the patient survived the procedure. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the delivery issues was due to patient condition, specifically the patient's diseased iliac artery. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows.
19123 | 1220648-2024-18685 | 20329304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329304&pc=OZD | 07/26/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/26/2021 | 09/28/2024 |  |  | 04/30/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022082022 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for hemodynamic support. it was reported that there was a defective impella catheter. the pump was disconnected and the catheter was replaced. arrhythmia-induced cardiomyopathy (aic) was restarted to resume successful primary percutaneous coronary intervention (ppci). no harm or injury was noted. | [5 empty] | The impella device has been returned for evaluation. clinical details determined the cause of the product damage was likely due to a use-related hole in the catheter, resulting in the blood to ingress through the catheter to the red handle. this report is being filed as part of a retrospective review of historical records.
19124 | 1220648-2024-18686 | 20329334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329334&pc=OZD | 07/26/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/26/2021 | 09/28/2024 |  |  | 03/31/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022062075 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | Patient-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp support for hemodynamic support. the pump was unable to cross the valve due to patient having an aortic aneurysm and an odd valve angle. multiple attempts were made but insertion was unsuccessful. the procedure was aborted. no injury or patient harm was noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
19125 | 1220648-2024-18687 | 20329376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329376&pc=OZD | 07/26/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/30/2021 | 09/28/2024 |  |  | 05/31/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2022082897 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for hemodynamic support. it was reported that the patient had heparin induced thrombocytopenia, and increased purge pressure 1050 mmhg and poor purge flow of 2.0 ml/h. the purge pressure remained high after the liquid was changed. the decision was made to remove the impella cp device. no injury or patient harm was noted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has been returned for evaluation. however, due to the condition the pump was returned in and the lack of a returned purge cassette, a root cause for the reported high purge pressure could not be determined through this investigation. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
19126 | 1220648-2024-18688 | 20329409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329409&pc=OZD | 07/31/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 07/31/2021 | 09/28/2024 |  |  | 03/31/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022076019 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the cp pump was having low purge pressures and high flows and higher than normal flows for p level selected. the device was exchanged for a new impella cp device. no patient harm or injury was reported. additional information provided that the patient was successfully weaned, and survived the procedure. | [5 empty] | The investigation for the reported high purge pressure issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the root cause of the low purge pressure was most likely due to bearing alignment, as no leaks were found in the purge system and the pump was naturally purging at pressures on the lower end of the specification as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
19127 | 1220648-2024-18689 | 20329416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329416&pc=OZD | 08/02/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 08/02/2021 | 09/28/2024 |  |  | 01/31/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022044294 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a pump stop. the motor control was high and there was no flow. the pump was explanted and the patient was switched to left ventricular assist device (lvad). no patient harm or injury were noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
19128 | 1220648-2024-18690 | 20329461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329461&pc=OZD | 08/02/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 08/02/2021 | 09/28/2024 |  |  | 06/30/2023 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022097056 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for hemodynamic support. while on support, during setup for transfer, it was noted that the patient developed a severe hematoma at the groin and scrotum. pressure was held for a period of time. the physician explanted the impella and placed the femstop. once explanted, the groin began to stabilize. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the reported access site bleeding was most likely related to anticoagulation management as the patient¿s latest act readings was 299 seconds. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
19129 | 1220648-2024-18691 | 20329465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20329465&pc=OZD | 08/05/2021 | 09/28/2024 |  | 09/28/2024 |  |  | 08/05/2021 | 09/28/2024 |  |  | 10/31/2022 |  | 09/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021013657 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  |  | White | Malfunction | No | Yes | Migration or Expulsion of Device | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump came out of position and was replaced. the position of the impella cp device was unclear and transthoracic echocardiogram (tte) was not feasible. | [5 empty] | The investigation for the reported delivery issues issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the product was returned and a kink was found in the cannula. the root cause of the cannula kink during repositioning was determined to be a result of attempted wireless reaccess. the root cause of the delivery issues was determined to be the impella device is not indicated for wireless re-access. impella cp with smartassist for use during cardiogenic shock and high risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
19130 | 1220648-2024-18629 | 20319121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20319121&pc=PYX | 06/05/2023 | 10/29/2024 |  | 09/27/2024 |  |  | 06/05/2023 | 09/27/2024 |  |  | 01/31/2025 |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024305335 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Sepsis | Multiple Organ Dysfunction Syndrome | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a rp flex-cp bipella support ongoing when the rp had low flows that prompted pump replacement. there was a presumption of the pump ingesting biomaterial after the support of four days. the 2nd pump was placed successfully. the patient survived the event. | A notification was received from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry). it was noted that the patient on impella support endured sepsis and worsening multisystem organ failure. relatedness to impella device is unknown. the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was returned and evaluated. the root cause of the low pump flow was determined to be ingested biomaterial (thrombus), likely due to the patient's activated clotting times being less than therapeutic with no systemic anticoagulation. | B2 exact date of death is unknown. impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2024-18629 was provided.
19131 | 1220648-2024-18631 | 20320554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20320554&pc=OZD | 09/03/2024 | 10/15/2024 |  | 09/27/2024 |  |  | 09/03/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508806 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump support initiated for a patient post coronary artery bypass graft. the impella 5.5 was inserted but, after 4 days there was multiple runs of torsade's and ventricular tachycardia that were treated with shock and medical therapy. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
19132 | 1220648-2024-18544 | 20320582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20320582&pc=PYX | 04/01/2021 | 05/22/2025 |  | 09/27/2024 |  |  | 04/01/2021 | 09/27/2024 |  |  | 08/31/2021 |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1452112 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella rp device for mechanical circulatory support. during implantation, and upon start up, the flows were lower than expected. the device was repositioned, yet the flow rates did not improve. the device was removed, reinserted, with no flow improvement. the device was removed and discarded by staff before onsite support arrived. survival outcome at explant noted the patient expired. medical safety review of the event noted that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19133 | 1220648-2024-18632 | 20320651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20320651&pc=OZD | 09/04/2024 | 10/30/2024 |  | 09/27/2024 |  |  | 09/04/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025512698 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support received one unit of blood for recurrent gastrointestinal (gi) bleed. scoped was done for gi bleed but the source of the bleed was not located. it was further reported that the gi bleed seemed to be resolved however the patient outcome is unknown. | [5 empty] | The investigation for the gastrointestinal (gi) bleed has been completed. pump was not returned for investigation. data log review was not require for this case. as per the clinical, gi bleed was reported during impella use. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. the relation between gi bleed and impella was unknown. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
19134 | 1220648-2024-18633 | 20320712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20320712&pc=OZD | 08/31/2024 | 10/24/2024 |  | 09/27/2024 |  |  | 08/31/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521384 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 59-year-old male patient in a cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. post impella cp insertion, the peel away was removed and reposition sheath was placed without any issue. in the intensive care unit, the patient became agitated and with movement access site bleeding was noted. manual pressure was held without success. the impella cp was removed and an impella 5.5 was placed. the patient survived the explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be patient movement because the patient became agitated and the movement caused the bleed around the repo sheath.
19135 | 1220648-2024-18634 | 20320714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20320714&pc=OZD | 02/15/2024 | 11/15/2024 |  | 09/27/2024 |  |  | 09/05/2024 | 09/27/2024 |  |  | 11/30/2025 |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025437521 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed team received a literature article involving a 77-year-old female patient on imeplla 5.5 support. the patient had significant calcification at the bifurcation with the common carotid artery. it was reported that attempts to remove the impella using standard techniques were unsuccessful due to its entrapment in the calcified area of the right subclavian artery. a balloon dilatation catheter was used, and with the impella positioned on the balloon, careful and slow deflation allowed the device to pass through the calcification and be successfully removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the removal issues has been completed. no product was returned for investigation. the clinical details mention the difficulty removing the pump was attributed to the calcification in the right subclavian artery. the root cause of the removal issues was most likely due to patient condition. added-h6 impact code 4641.
19136 | 1220648-2024-18635 | 20320869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20320869&pc=OZD | 09/05/2024 | 10/30/2024 |  | 09/27/2024 |  |  | 09/05/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492767 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had significant bleeding due to high dose of heparin. the following day the patient was doing better but with bleeding issue, the team opted to remove the impella. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
19137 | 1220648-2024-18636 | 20320894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20320894&pc=OZD | 11/30/2023 | 11/04/2024 |  | 09/27/2024 |  |  | 08/31/2024 | 09/27/2024 |  |  | 09/30/2025 |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403749 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured large pericardial effusion with a "possible" relationship to the impella cp device used: the patient was taken emergently back to the catheterization lab for a large pericardial effusion seen on the transthoracic echocardiogram. the patient returned to the cardiac intensive care unit with a pericardial drain in place. it was further reported that the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The investigation for the reported vascular damage has been completed. the pump was not returned for investigation. as per the given clinical details, pericardial effusion was reported during impella use. the cause of the vascular damage issue was not determined since the product was not returned and sufficient clinical information was not provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
19138 | 1220648-2024-18637 | 20320922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20320922&pc=PYX | 08/18/2021 | 09/27/2024 |  | 09/27/2024 |  |  | 08/18/2021 | 09/27/2024 |  |  | 03/31/2022 |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1478848 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support. it was reported that while on impella rp support, the patient experienced a little bit of oozing. a second mattress suture was placed and the sheath secured in placed. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19139 | 1220648-2024-18638 | 20320977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20320977&pc=OZD | 09/01/2024 | 11/20/2024 |  | 09/27/2024 |  |  | 09/01/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025509003 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while advancing the impella 5.5 through the graft into the aorta, the graft anastomosis started bleeding. the anastomosis was re-stitched, and bleeding was resolved. four unit of blood were administered. the patient eventually expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was determined to be an use issue because the insufficient anastomosis through graft into aorta which caused the bleeding was resolved by re-stitching the graft anastomosis. | A.5 is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
19140 | 1220648-2024-18639 | 20320978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20320978&pc=PYX | 11/11/2022 | 05/21/2025 |  | 09/27/2024 |  |  | 11/11/2022 | 09/27/2024 |  |  | 07/31/2024 |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023250993 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced significant bleeding from the chest tube (1000lm) in cath lab and the patient's blood pressure was low on arrival. albumin 500ml bolus was administered. the patient continued on support until the device was successfully weaned. the patient outcome was noted to be stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ impella rp with smartassist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19141 | 1220648-2024-18640 | 20321055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20321055&pc=OZD | 09/05/2024 | 11/14/2024 |  | 09/27/2024 |  |  | 09/05/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025506106 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an attempt was made to insert impella 5.5, but it could not be passed through the aorta because the aortic arch had previously been replaced. an artificial blood vessel under the left clavicle on the other side was attempted unsuccessfully. the case of the impella 5.5 was aborted and impella cp was placed. no patient harm was reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, impella could not be passed through the aorta as per the clinical description provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ion: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
19142 | 1220648-2024-18641 | 20321096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20321096&pc=OZD | 09/02/2024 | 11/07/2024 |  | 09/27/2024 |  |  | 09/02/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025492172 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Defective Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump delivered for the support of the patient with microvascular decompression/coronary artery disease in need of coronary artery bypass graft. the first impella sheath that was placed had bleeding near the rim of the hemostatic valve. it appeared to broken. therefore, a new one was placed. the impella was prepped and placed without an issue. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for the vascular damage and the introducer damage has been completed. no product was returned for evaluation. clinical details noted that the vessel dissection occurred prior to graft anastomosis to the artery. a repair was required. additionally, first introducer sheath used experienced bleeding at hemostatic valve to shaft connection and damage was noted. no excessive force was noted and best practices were followed. the root cause of the vascular damage peripheral vessel could not be determined as limited clinical details were provided. the failure mode access site bleeding is more applicable in this case as there was bleeding reported from the rim of the hemostatic valve. the root cause of the access site bleeding major is most likely due a damaged valve, however the specific location and cause of the damage is unknown. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
19143 | 1220648-2024-18642 | 20321147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20321147&pc=OZD | 09/02/2024 | 11/07/2024 |  | 09/27/2024 |  |  | 09/02/2024 | 09/27/2024 |  |  | 10/31/2025 |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407913 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 64-year-old male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella cp support, the patient developed bleeding at the impella access site. two units of platelets and one unit of packed red blood cells were administered. the patient survived the explant. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
19144 | 1220648-2024-18643 | 20321166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20321166&pc=OZD | 08/03/2024 | 10/31/2024 |  | 09/27/2024 |  |  | 09/06/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525140 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 50 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombocytopenia | [10 empty] | Disability | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding, the complainant had an impella cp support ongoing for a 50-year-old female patient admitted in (b)(6) 2024 for a patient with history of post partum cardiomyopathy and heart transplant rejection. patient was supported by the impella cp and ecmo and distal perfusion catheter. during the support the purge was changed from heparin to sodium bicarb due to bleeding from all access sites. the support was successful for just under 11 days. later in (b)(6) 2024 the patient had thrombocytopenia with "definite" relationship to the impella. | [5 empty] | The investigation for the thrombocytopenia, access site bleed, and limb ischemia has been completed. the impella device was not returned for evaluation. however, clinical details were provided sufficient to determine the root cause of the access site bleed. the root cause of access site bleeding is determined to be patient condition because the patient is bleeding from multiple locations including the access site. the root cause of the thrombocytopenia and limb ischemia could not be determined without additional clinical details. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
19145 | 1220648-2024-18645 | 20321324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20321324&pc=PYX | 09/18/2021 | 09/27/2024 |  | 09/27/2024 |  |  | 09/18/2021 | 09/27/2024 |  |  | 02/28/2022 |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022052399 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the complainant reported multiple attempts were made to advance the impella rp to the pulmonary artery. the device was turned on with no positivity in the motor current or placement signal. the device was explanted, and the procedural outcome was the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella rp smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records.
19146 | 1220648-2024-18644 | 20321522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20321522&pc=OZD | 09/06/2024 | 12/03/2025 |  | 09/27/2024 |  |  | 09/06/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1712868 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller was not reading the pressure. this issue was discovered in use, in the intensive care unit however patient information is unknown. there was no reported patent harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
19147 | 1220648-2024-18646 | 20321575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20321575&pc=PYX | 05/16/2024 | 10/01/2024 |  | 09/27/2024 |  |  | 08/06/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025443714 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "possible" relationship to the impella rp device used. the patient developed hemolysis and increased right ventricular assist device speed to 7500 rpm with flow 3.9 liters per minute due to worsening hemolysis. the patient recovered with no sequelae. | [5 empty] | The investigation for the reported hemolysis has been completed. no product was returned for investigation. data logs showed a brief onset of multiple "suction" alarms in the middle of support; otherwise, there were no abnormal relevant pump performance metrics observed. therefore, the cause of the hemolysis was not determined since the product and sufficient clinical information were not provided. | The investigation has been completed. no product was returned. per the clinical investigation, the cause of the hemolysis was not determined since the product and sufficient clinical information were not provided.
19148 | 1220648-2024-18647 | 20321821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20321821&pc=OZD | 08/28/2024 | 11/12/2024 |  | 09/27/2024 |  |  | 09/06/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527155 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was bleeding from the red impella plug. excessive force was not applied when placing impella cp. the patient outcome is unknown. | [5 empty] | The investigation for the product damage has been completed. the product was returned and a significant hole was found around the 75 cm mark. blood was expelling from the area. the root cause of the product damage (hole in catheter) is due to use that most likely occurred from a puncture. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." a2 patient age and a3 patient gender are unknown d6b explant date is unknown.
19149 | 1220648-2024-18648 | 20321882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20321882&pc=OZD | 09/06/2024 | 11/12/2024 |  | 09/27/2024 |  |  | 09/06/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532769 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  |  |  | Malfunction | No | Yes | Positioning Failure | Migration or Expulsion of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support due to cardiogenic shock. during transfer from bed to computed tomography scan, impella position lost displayed. the pump was subsequently explanted and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. based on clinical description, the root cause of positioning issue was most likely patient movement. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
19150 | 1220648-2024-18649 | 20322010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20322010&pc=OZD | 09/06/2024 | 11/06/2024 |  | 09/27/2024 |  |  | 09/06/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523264 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Sense | Adverse Event Without Identified Device or Use Problem | [5 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support due to st elevated myocardial infarction (stemi). while on support, impella alarmed "impella in ventricle", however placement signal appears aortic. placement signal alarms were silenced. additionally, the patient had a gastrointestinal bleed and received 2 units of red blood cells. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported gastrointestinal bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the gastrointestinal bleed could not be determined. added- h6 med dev prob code 2993 corrections to the initial mfr report: b1-product problem selected in error. b5-the statement "while on support, impella alarmed "impella in ventricle", however placement signal appears aortic. placement signal alarms were silenced." is not considered a reportable event. h6 med dev prob code 1559 is not relevant to reportable event.
19151 | 1220648-2024-18650 | 20322228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20322228&pc=OZD | 09/06/2024 | 10/31/2024 |  | 09/27/2024 |  |  | 09/06/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532156 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainants reported a 48-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was further reported that after impella removal, the patient was taken to a ct scan due to the distal pulses that were weaker on doppler than they were before. ct result showed some clot formation distal to previously placed impella. vascular surgeon used fogerty catheter and expelled clot. the patient remained in stable condition with positive distal pulses. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined.
19152 | 1220648-2024-18651 | 20322415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20322415&pc=PYX | 09/26/2022 | 05/20/2025 |  | 09/27/2024 |  |  | 09/26/2022 | 09/27/2024 |  |  | 09/30/2023 |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023193952 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support status post cardiothoracic surgery. it was reported during impella rp insertion the physician experienced difficulty place the cannula. it was reported the cannula would bow during insertion and the wire placement was lost. multiple wires and techniques were used, and the patient's graft was trimmed down to finally insert the impella. it was additionally reported that the patient was bleeding from chest tubes, and was transfused with 12 units of packed red blood cells, 10 units of fresh frozen plasma, 4 units of platelets, and 5 units of cryoprecipitate. the patient's belly was noted to be very distended and hard, and continued to worsen as well as blood gas results indicating the patient was declining. the physician came in and opened the patient's chest around 4am to ensure it was not tamponade and found about 800 cc of blood in the chest due to chest tube that had clotted off. due to abdomen swelling, the physician opened the peritoneum at the tip of the sternal incision and placed an abdominal suction which output 700 of ascites. sump drain was placed and chest was covered with loban. the patient's family withdrew care, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support status post cardiothoracic surgery. it was reported during impella rp insertion the physician experienced difficulty place the cannula. it was reported the cannula would bow during insertion and the wire placement was lost. multiple wires and techniques were used, and the patient's graft was trimmed down to finally insert the impella. it was additionally reported that the patient was bleeding from chest tubes, and was transfused with 12 units of packed red blood cells, 10 units of fresh frozen plasma, 4 units of platelets, and 5 units of cryoprecipitate. the patient's belly was noted to be very distended and hard, and continued to worsen as well as blood gas results indicating the patient was declining. the physician came in and opened the patient's chest around 4am to ensure it was not tamponade and found about 800 cc of blood in the chest due to chest tube that had clotted off. due to abdomen swelling, the physician opened the peritoneum at the tip of the sternal incision and placed an abdominal suction which output 700 of ascites. sump drain was placed and chest was covered with loban. the patient's family withdrew care, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿. ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿. potential adverse events (united states) . arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19153 | 1220648-2024-18653 | 20322425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20322425&pc=PYX | 09/22/2021 | 09/27/2024 |  | 09/27/2024 |  |  | 09/22/2021 | 09/27/2024 |  |  | 08/31/2022 |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 1496359 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced access site bleeding, the resolution for this issue is unknown. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19154 | 1220648-2024-18652 | 20322426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20322426&pc=PYX | 07/09/2022 | 09/27/2024 |  | 09/27/2024 |  |  | 07/09/2022 | 09/27/2024 |  |  | 03/31/2024 |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2023204612 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic was implanted with an impella rp device for mechanical circulatory support. while on support the there was ooze at the right femoral vein site after the patient had kicked the right leg. the complication was successfully managed by placing a new forward tension suture and manual compression for 15 minutes. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions: "make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound." "be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open." "make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound." this report is being filed as part of a retrospective review of historical records.
19155 | 1220648-2024-18654 | 20322529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20322529&pc=OZD | 09/05/2024 | 05/27/2025 |  | 09/27/2024 |  |  | 09/06/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516536 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup | Followup | Followup | [6 empty] | Additional information was received via medwatch report (b)(4). the patient was in imaging. during the return transfer to the patient's bed, the impella catheter broke causing the machine to stop. the registered nurse at the bedside heard a ¿pop¿ and identified the problem immediately and intervened with vasopressors, as needed. the patient had a slight drop in blood pressure, which was corrected with the vasopressors. no other change in condition was noted and the patient tolerated well. | The complainant reported a patient on support with a impella 5.5 pump admitted in acute myocardial infarction and cardiogenic shock. it was reported that while the patient was being transferred back to intensive care unit bed the staff inadvertently pulled on the impella 5.5 where the catheter was fully disconnected from red plug. only the fiber optics held the catheter to red plug and the catheter failed to work. no patient harm was report and another pump was used to support the patient. |  |  |  |  | The investigation for the product bond failure has been completed. no product or images of damaged were returned for evaluation. clinical details noted that when patient was being transferred back to icu after ct scan, pump was ripped at red handle with excessive force applied. data logs from field were reviewed and real time (rt) logs from time of product damage shows sudden pump stop with a motor current spike to 30000. placement signal also goes out of range at time of damage. the root cause of the product damage (bond failure) and subsequent pump stop was most likely an electrical open at the patient cable to red handle due to a use issue as excessive force was applied when damage occurred. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter. | B.5 additional information was received. e.4 revised to yes as med watch repot was received. g.2 added med watch report number. med watch report attached to this report. b.3 date of event was revised as it was previously entered incorrectly on the initial manufacturer device report number. b.7 added information that was inadvertently omitted from the initial manufacturer device report number. hypetension was reported incorrectly on the initial manufacturer device report number. e.1 fax number was added as it was inadvertently omitted from the initial manufacturer device report number. g.1 reporting contact email was revised since the initial manufacturer device report number was submitted in accordance with updated procedures. g.3 date of awareness on the initial manufacturer device report number should of reflected 5-sep-2024. h.6 code was added as it was inadvertently omitted from follow up 1 manufacturer device report number. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | Related medwatch report (b)(4) populated in section h10.
19156 | 1220648-2024-18655 | 20322708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20322708&pc=OZD | 09/06/2024 | 10/31/2024 |  | 09/27/2024 |  |  | 09/06/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025531963 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 63-year-old male patient with caridomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, the patient had bleeding and vascular damage. the patient brought to cvor for emergent replacement of impella. the previous 5.5 was damaged, needed to exchange through same graft. device was previously placed a week prior, so decision was made to expel or bleed back through graft so no clot would leave in the graft. the goal was not to have clot ingested by new impella 5.5 that was going to be exchanged in same graft. the pocket exposed and graft secured and impella explanted. impella 5.5 was advanced with no issues. bleed of graft 200-300cc of blood, clamped and introducer placed, the plan was for 1 unit of rbc to be given in intensive care unit (icu). | [5 empty] | The investigation for the access site bleed and the vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and vascular damage could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
19157 | 1220648-2024-18656 | 20322906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20322906&pc=PYX | 05/15/2021 | 09/27/2024 |  | 09/27/2024 |  |  | 05/15/2021 | 09/27/2024 |  |  | 10/31/2021 |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 1460452 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | Failure to Sense | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that the patient was in poor condition prior to implanting the impella rp and cp. the patient's right coronary artery was dissected, the patient was intubated, lost pulse, cardiopulmonary resuscitation was performed, and circulation returned. after repair of the dissection, impella support was initiated. delivery issues were encountered at time of rp implant, however the impella rp was successfully implanted and supported the patient. the impella rp device slowly ramped up to p-9 and flows of 3.1. the motor current was stable and had stable flows. post procedure the patient was sent to the intensive care unit in critical condition for further care. the next day post bipella implant, the medical staff noted that the patient continued to be in very poor state of health. it was also stated that the impella rp was supporting the patient and noted to be in "great" position. second day post support, it was noted the patient eventually met her demise (prognosis was "very" poor). use of the device cannot conclusively be associated with the outcome (patient's demise). patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: & rp smart assist system with automated impella controller section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records.
19158 | 1220648-2024-18657 | 20322952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20322952&pc=OZD | 09/06/2024 | 10/31/2024 |  | 09/27/2024 |  |  | 09/06/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530788 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 77-year-old male patient had an attempted impella cp implant and it was aborted due to patient anatomy. the patient had multivessel disease. the introducer was said to be used on the patient. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
19159 | 1220648-2024-18659 | 20323304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20323304&pc=OZD | 09/07/2024 | 11/07/2024 |  | 09/27/2024 |  |  | 09/07/2024 | 09/27/2024 |  |  | 11/30/2025 |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426041 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 31-year-old male patient with caridomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. while on impella support, heparin was withheld pending heparin induced thrombocytopenia (hit) panel. the panel was negative and heparin was resumed. additionally, high purge pressure was observed. cassette was changed. tissue plasminogen activator (tpa) was started. d5 with bicarb was added to purge. purge pressures eventually normalized. | [5 empty] | The investigation for the high purge pressure has been completed. the pump was not returned for inspection. the data logs confirm high purge pressure and purge system blocked alarms. the high purge pressure was sustained for about 24 hours before it returned to normal levels aligning with tissue plasminogen (tpa) administration in patient updates. there were positioning related alarms during the case. the root cause of the high purge pressure was most likely biomaterial. h6 impact code 4644. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
19160 | 1220648-2024-18664 | 20324709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20324709&pc=OZD | 09/10/2024 | 02/25/2025 |  | 09/27/2024 |  |  | 09/10/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025543731 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information received on 09/23/2024 via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding the patient that endured hemolysis with a "definite" relationship to the impella device used. the patient recovered with no sequelae. | Us complainant reported the patient was on impella cp support due to st elevated myocardial infarction (stemi). while on support, suction alarms were noted overnight. the decision was made to explant cp. the patient survived the event. |  |  |  |  | Added-b1 selected adverse event. b2-selected life threatening. b5-additional narrative. h1-type of reportable event-changed to serious injury. h6-clinical code 1886. | The investigation of low pump flow and hemolysis has been completed. the pump was retuned and visually inspected. there was no cannula kink, no broken blade and no biomaterial observed. pump connected to controller to reproduce the low pump flow issue. pump flow was running within specification limit. no abnormality found. hemolysis test conducted. no hemolysis issue observed. no abnormality and issue found in the pump and hemolysis test. data log reviewed: suction alarms occurred a few times. no positioning issue observed. no motor current spike observed outside of p-level changes. flow looks correlated with p-levels changes. the cause of the low pump flow most likely was patient condition related. pump low flow cannot be reproduced and no abnormality found in the pump. patient had small left ventricle (lv). the cause of the hemolysis issue most likely was patient condition related as the patient had small lv. instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ g1 reporting contact email revised on manufacturer device report 1220648-2024-18664 in accordance with updated procedures. g2 report source; study was missing from manufacturer device report 1220648-2024-18664 h6 revised to reflect additional health effect - impact code for hemolysis manufacturer device report 1220648-2024-18664-1. h10 revised to include instructions for use for hemolysis which was missing from manufacturer device report 1220648-2024-18664-1. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
19161 | 1220648-2024-18660 | 20324748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20324748&pc=PYX | 09/07/2024 | 11/11/2024 |  | 09/27/2024 |  |  | 09/07/2024 | 09/27/2024 |  |  | 08/31/2025 |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024399874 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella rp being delivered to the patient admitted with multiple cardiac and systemic comorbidities. it was reported that there were multiple attempts to deliver the rp past the pulmonic valve without success due to patient anatomy. the pump was replaced by a new rp flex pump and no harm to the patient was reported. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the impella pump implant was aborted due to patient's anatomy.
19162 | 1220648-2024-18662 | 20325396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20325396&pc=OZD | 09/07/2024 | 10/31/2024 |  | 09/27/2024 |  |  | 09/07/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505128 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 55-year-old male patient with caridomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, the impella 5.5 had sodium bicarb in the purge but after 43 days there was elevated purge pressure (pp) alarms. the high pp was troubleshot with tpa discussions and eventual pump explant/replacement. | [5 empty] | The investigation for the high purge pressure has been completed. no product was returned. log review showed a gradual increase in purge pressure beginning on (b)(6) with quicker increase from ~500 to ~900 mmhg in 3 hours on (b)(6). "purge pressure high" alarms on (b)(6) (last day of support) after continued increase. several suction alarms occurred during the increase in purge pressure. there were no spikes in motor current before the purge pressure increase. there were also several "impella position unknown" and "impella position in aorta" alarms throughout the case. the icu team gave tissue plasminogen activator (tpa), which decreased the purge pressure. the root cause of the high purge pressure was most likely biomaterial. added- h6 impact code 4644. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
19163 | 1220648-2024-18661 | 20325906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20325906&pc=OZD | 09/07/2024 | 10/25/2024 |  | 09/27/2024 |  |  | 09/07/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526623 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a bipella support ongoing with an impella cp and impella rp for a 68-year-old female patient admitted in an acute myocardial infarction / cardiogenic shock (ami/cgs) when after a week the patient experienced ventricular tachycardia (v)t that prompted treatment. the patient was shocked and medically managed. the patient was being weaned at the time but, after the vt the cp pump remained on for support. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
19164 | 1220648-2024-18663 | 20326405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20326405&pc=OZD | 09/06/2024 | 11/19/2024 |  | 09/27/2024 |  |  | 09/09/2024 | 09/27/2024 |  |  |  |  | 09/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025537372 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp had an impella defective alarm during preparation. the impella was exchanged. there was no harm to the patient. | [5 empty] | The investigation for the pump not detected issue has been completed. the pump was returned for investigation. no physical abnormalities were reported on the returned pump. the pump still had a red lumen in place. a controller failure alarm appeared after the pump was plugged into the console. the pump could not start. all 5s parameters were within specification limits. the pump passed all electrical testing on the motor cable lines and the communication lines. the cause of the pump not detected issue was eeprom corruption. device return date, h6 impact code 4629. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
19165 | 1220648-2024-18590 | 20308991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20308991&pc=OZD | 09/04/2024 | 10/23/2024 |  | 09/26/2024 |  |  | 09/04/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486956 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a patient admitted in with covid myocarditis in need of impella 5.5 support. the team attempted to deliver the impella 5.5 pump but the delivery failed via the axillary artery. the case was aborted and no reported patient harm. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, physician unable to advance impella 5.5 through axillary artery as per the clinical description provided.
19166 | 1220648-2024-18589 | 20308994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20308994&pc=OZD | 08/28/2024 | 10/18/2024 |  | 09/26/2024 |  |  | 08/28/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521372 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the patient had hemolysis. hemolysis remained after repositioning the pump. it was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. data logs were reviewed which showed that the pump generally ran without issue during support. there was intermittent low flow near the end of the case that triggered auto suction response and suction alarms. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
19167 | 1220648-2024-18591 | 20309002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309002&pc=OZD | 09/04/2024 | 10/28/2024 |  | 09/26/2024 |  |  | 09/04/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474296 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that impella 5.5 insertion was aborted due to left axillary artery size. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of access site bleeding is determined to be patient condition because the pump insertion was aborted due to l axillary artery size. corrections to the initial mfr report: d4-udi number. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
19168 | 1220648-2024-18592 | 20309003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309003&pc=OZD | 09/05/2024 | 11/15/2024 |  | 09/26/2024 |  |  | 09/05/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508656 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 implant. the impella was not able to pass the axillary artery into the aorta. the implant was aborted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since pump was not able to pass axillary artery into aorta. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
19169 | 1220648-2024-18594 | 20309012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309012&pc=OZD | 09/05/2024 | 10/30/2024 |  | 09/26/2024 |  |  | 09/05/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528945 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a bipella support with an impella cp and impella rp for a 49-year-old male patient admitted in an acute myocardial infarction / cardiogenic shock (ami/cgs) event with history of other comorbidities such as cerebrovascular accident (cva), chronic kidney disease (ckd), diabetes mellitus and iv drug use. the pumps were placed however there was an access site hematoma and on multiple days the patient received blood products. the groin was closed by the use of a manta device to achieve hemostasis. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
19170 | 1220648-2024-18593 | 20309015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309015&pc=OZD | 09/05/2024 | 10/25/2024 |  | 09/26/2024 |  |  | 09/05/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532063 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp to support a 48-year-old male patient admitted in with ami/cgs and history of cardiac disease. a coronary artery bypass grafting (cabg) surgery was not possible; therefore, the patient had percutaneous coronary intervention (pci). the cp accessed limb developed signs of ischemia and reperfusion cannula was placed to the limb. the perfusion became occluded and after 2 days of support the cp was explanted. patient survived. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report.
19171 | 1220648-2024-18595 | 20309016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309016&pc=OZD | 09/05/2024 | 11/15/2024 |  | 09/26/2024 |  |  | 09/05/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The abiomed team in japan located a publication that notes an impella patient with a gastrointestinal bleeding (gi). the gi bleed was in a patient in japan in which the bleed occurred after coronary artery bypass grafting (cabg). the bleed resolved with the explant of the impella. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the gastrointestinal bleed has been completed. no product was returned for analysis. limited case clinical details were provided. the root cause of the vascular damage peripheral vessel was not determined as limited clinical information was provided. unknown relation to impella since the source of the bleed was gi related. the instructions for use for the impella states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
19172 | 1220648-2024-18596 | 20309019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309019&pc=OZD | 09/06/2024 | 10/31/2024 |  | 09/26/2024 |  |  | 09/06/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025468647 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Patient-Device Incompatibility | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 71-year-old male patient with caridomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. during support, heparin was stopped due to hemolysis. one unit of blood was given. continuous renal replacement therapy (crrt) was started. additionally, there was an impella in aorta alarm noted. multiple attempts were made to reposition but were unsuccessful. noted high purge pressure alarm followed by purge pressure blocked alarm. given that purge pressure was blocked, the surgeon and interventional cardiologist (ic) elected to remove the pump. | [5 empty] | The investigation of high purge pressure, hemolysis, and positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the hemolysis was not determined since insufficient clinical details were provided and no product was returned. the root cause of the high purge pressure was most likely a kinked purge line. the root cause of the positioning issues was most likely use-related. e4 should have been left blank on manufacturer device report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿.
19173 | 1220648-2024-18597 | 20309022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309022&pc=PYX | 09/08/2024 | 12/03/2024 |  | 09/26/2024 |  |  | 09/08/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025529158 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | Followup | [6 empty] | Us complainant reported the patient was on impella rp flex support. while on support, high purge pressure alarmed. bicarbonate purge was swapped for heparin purge. high purge pressure alarmed again, and tissue plasminogen activator (tpa) was placed in purge. tpa ran for about 6 hours without alarming again and purge flow increased. after tpa purge finished, heparin was placed back in purge solution. | It was additionally reported that when the physician was assessing the patient, the physician pulled the pulmonary artery catheter due to inaccurate pressures. shortly after pulling the catheter, the impella pressures decreased, so the physician increased the p level to 9 to keep the flows about 1.5 l/min. shortly after, a clot was ingested by the rp flex, and flows lowered to 2.2 on p-9. as a result of the lower than expected flows, the impella rp flex was explanted. the patient was successfully weaned and explanted. |  |  |  |  | Corrections to the mfr mdr report #1220648-2024-18597-02: g3-should have been 22-nov-2024. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. the pump was returned for investigation and a clot was found wrapped around the impeller. there were no other physical abnormalities with the pump. the data logs confirm high purge pressure alarms. there were minor motor current and placement signal spikes throughout the case. there was a gradual increase in purge pressure until alarm was triggered. the high purge pressure was sustained for about 15 hours before it returned to normal levels aligning with clinical mention of tissue plasminogen activator (tpa) addition and resolution. there was a motor current spike followed by low pump flows and an upward shift in placement signal. this aligns with the timeline of the pa catheter being pulled as mentioned in the clinical details. the root cause of high purge pressure was determined to be biomaterial. the root cause of low pump flow was determined to be biomaterial. | Additional information added to b5 and h6. corrected b1 and h1: serious injury.
19174 | 1220648-2024-18598 | 20309026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309026&pc=OZD | 09/09/2024 | 10/07/2025 |  | 09/26/2024 |  |  | 09/09/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516549 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Infusion or Flow Problem | [5 empty] | Arrhythmia | Hemolysis | Thrombocytopenia | Thrombosis/Thrombus | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [7 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had placed a 5.5 pump for support of a patient admitted in cardiogenic shock with care escalated from an intra-aortic balloon pump. the pump had saturated flows on day 4 but, the pump remained on for support while awaiting team decisions. the patient survived the event after an eventual wean and explant. | The complainant also reported that, weeks later, the long-term outcome and quality indicator (loqi) database indicated the pump had experienced optical signal issues, which caused no harm or injury. subsequently, the loqi also reported an instance of hemolysis with a 'probable' relationship to the 5.5 pump. the patient¿s laboratory results were indicative of hemolysis, and eight units of red blood cells were administered. additionally, the loqi database reported acute liver injury, thrombocytopenia, cardiogenic shock, and thrombus. |  |  |  |  | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the root cause of the saturated flow was most likely biomaterial as evidenced by the saturated flows immediately following motor current spike observed in the logs. the data logs show stable 5s parameters with periods of lower signal-to-noise ratio (snr) readings. returned product had 5s parameters within the expected range with passed laser continuity test. the cause of the optical signal issue was not determined as returned product functioned to specification and limited clinical information was provided on suspected optical issue. the root cause of the hemolysis was damage to the impeller, but the cause of the scratches on the blades is unknown due to insufficient clinical details. the cause of the vascular damage - peripheral vessel was not determined as limited information was provided. thrombocytopenia was added as a failure mode because it was later reported through loqi. there were no clinical details related to the thrombocytopenia. the cause of the thrombocytopenia was not determined due to insufficient clinical details. the root cause of the thrombus was unable to be determined due to insufficient clinical details. the root cause of the issue was unable to be determined due to insufficient clinical details. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 adverse event updated. b2 required intervention to prevent permanent impairment/damage updated. b5 updated to include placement signal and loqi issue descriptions. h1 updated to serious injury. h3 updated. h6 updated to include loqi failure modes added. corrections that were made from the initial manufacturer device report number 1220648-2024-18598. d10 concomitant medical products and therapy dates updated. e4 should have been left blank on the initial report. g1 reporting contact email and reporting contact fax number updated.
19175 | 1220648-2024-18600 | 20309036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309036&pc=OZD | 09/10/2024 | 11/04/2024 |  | 09/26/2024 |  |  | 09/10/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520162 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a cp pump delivered for a patient with cardiogenic shock and cardiomyopathy. after 66 hours of support the pump came out of position in the left ventricle and was unable to be re-advanced. the cp was explanted and new cp pump was placed. the 2nd pump supports while patient continues to be evaluated for a left ventricular assist device. the patient survived the event. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue most likely was use related issue due to improper technique as the pump came out of position in the lv and was unable to be re-advanced. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
19176 | 20309046 | 20309046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309046&pc=OZD | 08/29/2024 | 09/05/2024 | 09/05/2024 | 09/26/2024 | 09/05/2024 | 09/26/2024 |  | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP 14 FR |  |  | OZD |  |  | 0048-003 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-003 |  | Yes |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 64 YR |  | Female | 98 KG |  |  | Malfunction | No | Yes | [7 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | During site care and dressing changes to left groin site, the impella motor spontaneously stopped, triggering stop alarms. multiple attempts to restart impella to p2 failed.
19177 | 1220648-2024-18601 | 20309365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309365&pc=OZD | 08/28/2024 | 10/15/2024 |  | 09/26/2024 |  |  | 08/28/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526243 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that impella cp was not inserted due to inadequate vessel size. decision was made to insert impella via femoral access and escalate if patient progressed. once introducer sheath was prepped, patient went asystole. patient expired. use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, femoral angiogram revealed poor vessel size unsuitable for impella insertion as per the clinical description provided.
19178 | 1220648-2024-18602 | 20309398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309398&pc=OZD | 08/28/2024 | 10/18/2024 |  | 09/26/2024 |  |  | 08/28/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523304 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and they had a hematoma at the access site. a blood transfusion was performed. it was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of hematoma has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 code 1888 was reported incorrectly on manufacturer device report 1220648-2024-18602. new code was added to health effect - clinical code.
19179 | 1220648-2024-18604 | 20309526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309526&pc=PYX | 03/09/2022 | 09/26/2024 |  | 09/26/2024 |  |  | 03/09/2022 | 09/26/2024 |  |  | 05/31/2023 |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022103171 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported patient status post heart transplant left the operating room supported by venoarterial extracorporeal membrane oxygenation (va ecmo) and an intra-aortic balloon pump. va ecmo was decannulated and patient continued to decline. the patient went into right ventricular failure, and the decision was made to place an impella rp device for mechanical circulatory support. during support it was noted that the patient was having bleeding and oozing around peel-away sheath that the team wanted to address in operating room after explant. the patient was transfused with two units of blood cells due due to the blood loss. the impella device was not explanted due to the bleeding. the patient was successfully explanted and the impella access site was surgically closed. the patient continued on va ecmo support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19180 | 1220648-2024-18605 | 20309535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309535&pc=OZD | 08/29/2024 | 10/31/2024 |  | 09/26/2024 |  |  | 08/29/2024 | 09/26/2024 |  |  | 07/31/2025 |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442319 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support due to st elevated myocardial infarction (stemi). while on support, it was noted that the pump was unable to go up to p9 due to suction alarms. it was also reported that there was oozing from the right internal jugular and two units of packed red blood cells were given. additionally, the patient developed a gastrointestinal (gi) bleed. however, the patient remained on support until family decided to remove care. the patient expired after care was withdrawn. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the low pump flow and vascular damage - peripheral vessel has been completed. the pump was not returned for investigation. data logs were reviewed which showed suction alarms while the pump was set to p9. the suction alarms were resolved by lowering the p level. some discrepancies in the motor current and placement signal trends were reported while running on a steady p level. additionally, the pump experienced suction alarms again while set to p9 on the following day. pump flow was adequate during the entire case run. according to the clinical findings, blood products were given during the time of the suction event. additionally, the patient developed a gastrointestinal (gi) bleed during impella use. the cause of the vascular damage issue was not determined since the product was not returned and sufficient clinical information was not provided. the relation between gi bleed and impella was unknown. the cause of the low pump flow issue was most likely patient condition related, based on data log review and provided clinical details.
19181 | 1220648-2024-18603 | 20309536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309536&pc=OZD | 08/23/2024 | 10/28/2024 |  | 09/26/2024 |  |  | 08/29/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516553 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding 50-year-old male patient that endured hemolysis with a "definite" relationship to the impella 5.5 device used: ¿8/23/24: pfh 160, 1-unit packed red blood cells (prbc) infused. (on 8/20/24 pfh 30) 8/27/24: pfh 30¿ the patient recovered with no sequelae. it was further reported that the patient endured renal failure with a "possible" relationship to the impella device used: ¿8/23/24: cr increased from 1.2 (8/20/24) to 3.6, nephrology consulted, crrt initiated @1957.¿ it was reported that the patient/event has not recovered/not resolved. patient expired on support. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The investigation for the hemolysis and renal failure has been completed. data logs were reviewed which showed pump was in improper position (aorta) corresponded with the time hemolysis occurred per clinical description that triggered alarms impella position in aorta. the intermittent suction alarms were also triggered during the case. no other issues were found on the logs. product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. the root cause of renal failure was unable to be determined as limited clinical details were provided and no product was returned for review. added- h6 impact code 4641.
19182 | 1220648-2024-18606 | 20309693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309693&pc=OZD | 08/29/2024 | 10/17/2024 |  | 09/26/2024 |  |  | 08/29/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527442 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and while on support, the patient was in critical condition. the patient had extracorporeal membrane oxygenation (ecmo) placed, and the flows reduced on the impella. the impella was unable to keep up with ecmo flows, and the impella was reduced to p1. fluids were given to the patient. overnight, flows on impella read 0.0 with no alarms. the decision was made to withdraw care due to stone heart diagnosis and the patient expired after care was withdrawn. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. the cause of the low pump flow issue most likely was patient condition related.
19183 | 1220648-2024-18607 | 20309871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309871&pc=PYX | 08/30/2024 | 10/07/2024 |  | 09/26/2024 |  |  | 08/30/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025518061 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella rp flex to support a patient admitted in for aortic valve replacement and abdominal aortic aneurysm repair. the support was ongoing when on day two plus, the pump was explanted for signs of hemolysis and kinking. the pump was removed and replaced with a new impella rp flex. the procedural outcome was the patient survived the surgery. | [5 empty] | A.5 race is unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | H.6 code 4627 was reported incorrectly on the initial manufacturer device report number 1220648-2024-18607. a new code was added to health effect - impact code.
19184 | 1220648-2024-18609 | 20309955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309955&pc=OZD | 08/30/2024 | 10/30/2024 |  | 09/26/2024 |  |  | 08/30/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025504282 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 ingested a clot and the physician decided to remove the impella. it was further reported that care was withdrawn the next day and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
19185 | 1220648-2024-18608 | 20309962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20309962&pc=OZD | 08/30/2024 | 09/26/2024 |  | 09/26/2024 |  |  | 08/30/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251884 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) was turned on and a yellow controller error alarm occurred. the aic was replaced. | [5 empty] | A.5 is unknown. the automated impella controller (aic) was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
19186 | 1220648-2024-18613 | 20310247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20310247&pc=OZD | 08/30/2024 | 11/04/2024 |  | 09/26/2024 |  |  | 08/30/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461218 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump delivered for the support of a patient for two days for the support during coronary artery bypass grafting. the pump was weaned and explanted after the two days. at the time of explant, there was resistance and the pump was separated. the pump broke at the cannula. all pieces were able to be removed. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation for the product damage has been completed. the returned device confirmed the cannula detachment from the detached inflow cage and had deforming on the cannula that confirmed grasping of the device. the root cause of the product damage (detached inflow/outflow - peripheral vessel) was determined to be use-related. h.6 code 4627 was inadvertently omitted from the initial manufacturer device report number 1220648-2024-18613 and was added. | A.5 is unknown. the impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
19187 | 1220648-2024-18611 | 20310307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20310307&pc=OZD | 08/30/2024 | 10/31/2024 |  | 09/26/2024 |  |  | 08/30/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504425 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 65-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. following the planned explant of an impella cp pump, the perclose failed during site closure and hemostasis could not be achieved. a surgical closure of the site was subsequently required involving a right groin cut-down and surgical closure of the artery using a 5-0 prolene. the patient survived the explant. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related, the physician was struggling with the first perclose when the sutured broke and with the second perclose the physician felt did not capture as no hemostasis appeared to be achieved without manual hold as per the clinical description provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
19188 | 1220648-2024-18612 | 20310346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20310346&pc=OZD | 09/04/2024 | 11/06/2024 |  | 09/26/2024 |  |  | 09/04/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025520117 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Migration or Expulsion of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and while on support, the patient walked and the pump became out of position. the impella was in the aorta and the staff was unable to reposition. the impella was explanted and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the positioning issues has been completed. . the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue most likely was patient movement. complaint device not returned for analysis. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
19189 | 1220648-2024-18615 | 20310501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20310501&pc=OZD | 08/26/2024 | 11/19/2024 |  | 09/26/2024 |  |  | 08/28/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025511529 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp stopped while on patient support. the pump stopped due to motor deterioration caused by exceeding duration of use and was not able to restart. it was further reported that the patient died while on impella support. abiomed's medical safety officer review noted that there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | B.2 the exact date of death is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. no product was returned for analysis. the data logs confirmed "impella stopped (rpm deviation)" alarm and confirmed that the pump was used longer than the indication for use for impella cp pumps which is 8 days. the pump was used for about 13 days when the pump stop occurred. the root cause of the pump stop was not determined since no product was returned and insufficient information related to failure was provided d.6b date of explant should of been left blank on initial manufacturer device report number 1220648-2024-18615 as the exact date is unknown. g.1 revised reporting contact fax number in accordance with updated procedures since initial manufacturer device report number 1220648-2024-18615 was submitted.
19190 | 1220648-2024-18614 | 20310537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20310537&pc=OZD | 09/04/2024 | 10/15/2024 |  | 09/26/2024 |  |  | 09/04/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025458125 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 61-year-old male patient with caridomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, the patient had ventricular tachycardia (vt). the patient was defibrillated and the pump was repositioned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined.
19191 | 1220648-2024-18610 | 20310645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20310645&pc=OZD | 08/30/2024 | 10/31/2024 |  | 09/26/2024 |  |  | 08/30/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025490887 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hematoma | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support a hematoma was noted on the 79-year-old male patient at the percutaneous coronary intervention (pci) access site and the impella access site was oozing with a hematoma. pressure dressing was applied and 4 units of blood products were given. additionally, high purge pressure alarms occurred. tissue plasma activator was initiated. support continued. the patient was on impella support for 147.47 hours before the patient expired. the relationship between the impella and cause of death is unknown. medical safety review of the event noted there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the high purge pressure and the access site bleed has been completed. no product was returned for evaluation. data logs from time of alarm shows a gradual rise in purge pressure over a period of approximately 30 minutes before alarms were triggered. purge pressure remained high for approximately 20 hours before purge pressure resolves. clinical details noted that high purge pressure/purge system blocked alarms were resolved with tissue plasminogen activator (tpa) after approximately 19 hours of use. the root cause of the high purge pressure was most likely due to biomaterial buildup as it was able to be resolved with tpa. clinical details noted that oozing from impella access site was present and at pci site a hematoma was present. patient had received 4 units of prbcs and bleeding was resolved with pressure dressings. the patient had been on tpa protocol for approximately 20 hours before oozing was noted. the root cause of the access site bleeding - major was most likely patient condition related as bleeding/oozing was present from multiple sites. added- h6 impact code 4644.
19192 | 1220648-2024-18616 | 20311422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20311422&pc=OZD | 08/30/2024 | 11/05/2024 |  | 09/26/2024 |  |  | 08/30/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521703 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp via the 14 french sheath and the 7 french companion sheath to support a 77-year-old male patient admitted in for a high risk percutaneous coronary intervention (pci). the pump was placed but the sheath had a malfunction. there was a leak that prompted the team to apply a hemostat to the orange ring. the team made the choice not to explant but to replace the sheath. the pci was completed and then explant occurred after two plus hours of support. the patient survived the explant. | [5 empty] | The investigation for the introducer damage has been completed. the product was not returned and logs are not applicable. clinical details stated that there was blood leaking from beneath the orange ring along the white wing of the peel away. this was managed with a hemostat. therefore, the root cause of the introducer damage and subsequent access site bleeding was not determined since the product was not returned. b.7 added information that was inadvertently omitted from initial manufacturer device report number 1220648-2024-18616. h.6 code 4641 was inadvertently omitted from initial manufacturer device report number 1220648-2024-18616 and was added. | A.5 race is unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
19193 | 1220648-2024-18619 | 20311654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20311654&pc=OZD | 09/05/2024 | 11/20/2024 |  | 09/26/2024 |  |  | 09/05/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025464977 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Syncope/Fainting | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 80-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient was on impella cp support and after the explant, the patient had a small access site bleed. the staff applied pressure and a femostop and the bleeding subsided. additionally, when the femostop was placed, the patient had a vasovagal reaction. staff gave volume and shock storage. the patient was stable after. | [5 empty] | The investigation for the reported vasovagal has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vasovagal could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
19194 | 1220648-2024-18617 | 20311721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20311721&pc=OZD | 08/30/2024 | 10/22/2024 |  | 09/26/2024 |  |  | 08/30/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025502132 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was implanted on a 46-year-old male patient for support during percutaneous coronary intervention (pci). after pci was completed, the impella cp was weaned and removed. as the impella was weaned it was noted that the 7 french sheath came out of the right radial. it is also noted that there are no 8 french angio-seal in the facility. the decision was made to use perclose twice. however, the physician was unable to push the knots down due to the patient's large body habitus. the physician re-sheathed with 8 french to discuss options. protamine was given at this time. the physician pulled the sheath and planned to hold manual pressure on the groin site. the patient became hypotensive and extremely agitated and aggressive. a knee immobilizer was placed as a precaution to the right leg. the patient survived the explant. | [5 empty] | The investigation for the removal issues has been completed. clinical details state that the 7fr sheath accidentally came out of the right radial and the patient was additionally moving throughout the case when pressure was held for the access site bleeding issue. the root cause of the removal issue is most likely due to patient movement based on the provided clinical details. added-h6 impact code 4644. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
19195 | 1220648-2024-18618 | 20311805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20311805&pc=OZD | 08/30/2024 | 11/04/2024 |  | 09/26/2024 |  |  | 08/30/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532030 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the 39-year-old male patient developed hemolysis. the doctor attempted to reposition the impella. by this time, the patient was successfully weaned off all pressors and the doctor decided to explant the impella. the patient was noted to be stable. the patient survived the explant. | [5 empty] | The investigation of hemolysis has been completed. data loga showed the pump ran without issue during support. there were pump position in aorta alarms triggered in early and the end of the case but resolved quickly. the product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
19196 | 1220648-2024-18622 | 20311968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20311968&pc=PYX | 09/06/2024 | 11/11/2024 |  | 09/26/2024 |  |  | 09/06/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025484900 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Migration or Expulsion of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella rp flex placed to support a patient admitted in acute myocardial infarction / cardiogenic shock (ami/cgs) along with an impella cp. the bipella support was ongoing for 1 day when the rp flex came out of position and was replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of positioning issue was most likely use related.
19197 | 1220648-2024-18620 | 20312654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20312654&pc=OZD | 08/30/2024 | 10/01/2025 |  | 09/26/2024 |  |  | 08/30/2024 | 09/26/2024 |  |  | 05/31/2025 |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354073 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Ischemia | [11 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a 41-year-old male patient experienced multiple complications during impella cp support. it was documented that the implanted pump triggered intermittent impella position in aorta alarms following successful repositioning. support was able to continue despite the noted issue. it was then reported that the patient developed hematuria with hemolyzed labs and their right leg became ischemic. the conditions went unresolved until care was withdrawn after only one day of support. the patient expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | B1 product problem updated. b2 updated to death. h1 updated to death. h6 updated to include placement signal coding and death. corrections that were made from the initial manufacturer device report number 1220648-2024-18620. d10 concomitant medical products and therapy dates updated. | The investigation for the hemolysis and limb ischemia has been completed. clinical details noted the patient was experiencing hematuria and signs of hemolysis and was given volume in attempt to resolve. the pump was noted to be mispositioned deep in the ventricle and was able to be repositioned. it was not noted if hemolysis resolved after pump repositioning. additionally, clinical details noted the patient was in multisystem organ failure. data logs from the field were reviewed and variable placement signal was continuous during the case. no motor current spikes were preset outside of p-level changes to indicate possible biomaterial ingestion. the root cause of the hemolysis could not be definitively determined as the patient's condition and improper positioning of the pump could have been contributing factors. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
19198 | 1220648-2024-18623 | 20312825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20312825&pc=OZD | 09/08/2024 | 11/04/2024 |  | 09/26/2024 |  |  | 09/09/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528909 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient suffered cerebral vascular accident (cva). clots were visualized in extra corporeal membrane oxygenation circuit. the patient survived the event. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: warnings, cautions & precautions: "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers."
19199 | 1220648-2024-18621 | 20313056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20313056&pc=OZD | 07/22/2024 | 09/26/2024 |  | 09/26/2024 |  |  | 09/06/2024 | 09/26/2024 |  |  | 11/30/2025 |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025505210 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant in canada had a cp support ongoing for a patient admitted in with a st elevated myocardial infarction. while in the intensive care unit, the team observed a hematoma. the hematoma did not prompt any intervention. later the team further reported anemia. the anemia was noted to have a causal relationship to the impella procedure. the severity was noted to be moderate. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
19200 | 1220648-2024-18624 | 20313891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20313891&pc=OZD | 08/31/2024 | 10/29/2024 |  | 09/26/2024 |  |  | 08/31/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1648389 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Company representation | [5 empty] | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after several self-reboots of the automated impella controller (aic), a self-system check fail error was displayed on the screen. | [5 empty] | The investigation for the console (aic) self check error has been completed. console logs from the reported day of event are consistent with complaint and show persistent failure of pbm2 communication, leading to self-reboots and system self check errors. the console was returned for evaluation. upon inspection of the console, loose metal fasteners (from the fan bracket) were found inside the console housing, which shorted the circuitry of pbm. although fan assemblies are subject to periodic replacement every 5 years, console was manufactured only 2 years ago, and there were no prior complaints or repairs. it is most likely that the fasteners were loose since console¿s manufacturing on 2022-03-18. | The automated impella controller (aic) was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
19201 | 1220648-2024-18625 | 20314347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20314347&pc=OZD | 09/09/2024 | 11/12/2024 |  | 09/26/2024 |  |  | 09/09/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025487468 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, the patient had hematuria and hemolysis. undetermined if hemolysis is related to the pump itself but is a suspected factor. the patient remains on impella support. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date added h6 code 2558 was reported incorrectly on manufacturer device report. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
19202 | 1220648-2024-18626 | 20314364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20314364&pc=PYX | 09/09/2024 | 09/17/2025 |  | 09/26/2024 |  |  | 09/09/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025442089 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The us complainant had an rp flex placed for the support of a patient with right heart failure and cardiomyopathy. the team found an access site bleed that prompted holding off on heparin and further treating with a mattress suture. a knee immobilizer was present. the patient survived the event. | It was additionally noted that during transport, as patient was being lifted on the bed in unit, the impella catheter was caught under the patient as she was rolled to her side and caused the impella catheter to be compressed briefly. | The user facility reported a patient with cardiomyopathy and right heart failure was implanted with an impella rp flex for mechanical circulatory support and three days after start of support, bleeding from the access site was noted which prompted discontinuation of heparin and further treatment with a mattress suture. a knee immobilizer was present. it was reported the patient remained stable. additional information was received and noted two days later after the access site bleeding adverse event, the impella rp flex catheter and fiber optic cable were inadvertently compressed during move of the patient. during transport, as patient was being lifted on the bed in unit, the impella rp flex catheter became caught under the patient when rolled to the side and this caused the impella catheter to be briefly compressed. the automated impella controller display screen no longer showed the pulmonary artery and central venous pressure (cvp) placement signals or impella flows; only the motor current, purge flow and purge pressure were displayed. an x-ray and echocardiogram were completed to confirm correct placement and determined that impella rp flex pump was correctly positioned and continued to function. the cvp continued to improve. it was reported the patient was unharmed following the event despite the fiber optic cable not functioning as intended. |  |  |  | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | Additional information has been added to b2, b5, g3, g6, h6. | Additional information was received, and updates were made to sections b1, b5, and h6 accordingly. the evaluation and analysis of the returned impella rp flex device have not been completed. upon investigation closure, a supplemental mdr will be filed. correction was made to the following section(s): sections d1 (brand name) and d4 (model number) modified to reflect impella rp flex. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the pump was returned and evaluated. there was some biomaterial and blood ingress found on the dp sensor. there were no abnormalities with the electrical resistances. the pump was tested in the flow loop and sensor drift reproduced. there were no other abnormalities with the returned product. log review showed a downward drift in placement signal before it went out of range and triggered an alarm. pump flow went to zero at this time while motor current and p-level stayed stable. the failure mode "optical signal issue" was updated to "placement signal issue" because the issue involved the dp sensor. the root cause of the placement signal issue was most likely sensor drift since it was reproduced in the lab. the root cause of the access site bleeding was most likely use-related since the bleeding resolved by adding a mattress suture. the root cause of the product damage issue was most likely a catheter kink due to patient movement during support. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures from the initial manufacturer device report number 1220648-2024-18626. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-18626.
19203 | 1220648-2024-18627 | 20314822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20314822&pc=OZD | 09/10/2024 | 11/12/2024 |  | 09/26/2024 |  |  | 09/10/2024 | 09/26/2024 |  |  |  |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532219 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, hemolysis and limb ischemia developed. patient had hematuria and there was no flow per dropper from the popliteal down, on both legs. hemolysis due to malpositioned pump. pump could not be repositioned due to boot shaped left ventricle (lv). pump came across the aortic valve during repositioning and was unable to re-advance. decision was made to explant the pump. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of hemolysis and limb ischemia has been completed. the logs were returned and there were no motor current spikes, trend in placement signal, suction alarms or positioning alarms seen. the device was returned and there were no abnormalities found. the clinical confirmed that the pump was mal-positioned due to the patients boot shaped lv. the root cause of the hemolysis is most likely due to improper positioning. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. d9: date device returned to manufacturer added. h6: code 2558 was reported incorrectly on manufacturer device report 1220648-2024-18627.
19204 | 1220648-2024-18628 | 20315805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20315805&pc=PYX | 07/17/2023 | 11/22/2024 |  | 09/26/2024 |  |  | 07/17/2023 | 09/26/2024 |  |  | 12/31/2024 |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024308919 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | Failure to Advance | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Upon further review, it has been clarified that this issue is a positioning issue that occurred shortly after initial placement of the impella device. the pump migrated into the right ventricle following delivery which led to the previously reported advancement / positioning issues. | The complainant reported that the pump cannula kinked during attempted delivery for impella support. it was documented that the issue occurred while trying to advance the pump across the pulmonic valve. after multiple unsuccessful attempts, the decision was made to utilize another pump for implant. no patient harm was reported due to the issue. |  |  |  |  | A correction has been made to b5 and h6 clarifying that the noted issue was a positioning failure that occurred following implant and not a delivery failure. the additional code of positioning failure has been added to h6 to represent this correction as well as component code 3036 to represent the kinked cannula that was previously reported. should additional relevant information become available, a supplemental report will be submitted. | The investigation of the deliver issue is completed. product was not returned for review. the root cause of positioning issue was due to attempt to recross the valve wirelessly based on the clinical account. the impella device is not indicated for wireless re-access. this report is being filed as part of a retrospective review of historical records.
19205 | 1220648-2024-18630 | 20317089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20317089&pc=PYX | 10/02/2022 | 09/26/2024 |  | 09/26/2024 |  |  | 10/02/2022 | 09/26/2024 |  |  | 03/31/2023 |  | 09/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022068396 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella rp device for mechanical circulatory support. while on support, the patient experienced extensive bleeding noted near the impella access site with the reposition being pushed out due to high venous pressure. the physician was able to place 3 mattress sutures around the reposition sheath as well as sutured the reposition sheath to the patient's skin. site was angle matched and dressing was completed by the scrub nurse. later on, another physician came to the unit and re-did the sutures around the impella repositioning sheath due to extensive oozing. oozing was markedly decreased after the physician new sutures. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19206 | 1220648-2024-18477 | 20299397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20299397&pc=PYX | 03/09/2022 | 05/21/2025 |  | 09/25/2024 |  |  | 03/09/2022 | 09/25/2024 |  |  | 11/30/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022151985 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support. while on support, the patient was taken back to the operating room due to increased bleeding from the sternotomy and chest tubes. the chest was re-opened and washed out. at the end of procedure, it was decided by the surgeon to leave the chest open. the impella rp appeared to be in a good position. the purge system was changed out due to purge pressure low alarm. purge dextrose was increased to ten percent and no heparin was in the purge due to the patient¿s bleeding. after the procedure, the patient began to decompensate. care was withdrawn and the patient expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19207 | 1220648-2024-18478 | 20299450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20299450&pc=PYX | 03/09/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 03/09/2022 | 09/25/2024 |  |  | 01/31/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023167825 |  | 004334 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support. the impella rp was placed in right femoral vein and to decompress the right vein, the surgeon placed a mattress suture around the site. the sheath was swapped and mattress suture was tampered down. as there was still significant bleeding from the site, a figure eight around site was made as well. this helped and manual pressure was discontinued after ten minutes with good hemostasis. additional information received reported the patient expired while on support. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19208 | 1220648-2024-18479 | 20299489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20299489&pc=PYX | 03/10/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 03/10/2022 | 09/25/2024 |  |  | 04/30/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022098522 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella rp pump for mechanical circulatory support. following implant, bleeding was noted at the insertion site. manual pressure was held and two units of packed red blood cells were transfused to treat the condition. support was continued with the implanted pump following the successful intervention. the patient eventually passed away on support. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19209 | 1220648-2024-18480 | 20299530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20299530&pc=PYX | 01/20/2022 | 09/26/2025 |  | 09/25/2024 |  |  | 01/20/2022 | 09/25/2024 |  |  | 01/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022031806 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic was implanted with an impella rp device for mechanical circulatory support. while on support it was reported that the patient had significant venous bleeding despite three mattress sutures. the vascular surgeon was called and placed several sutures to control the bleeding. additional information received reported the patient expired while on support. medical safety review of the event noted that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. patient's peri-procedural condition was critical. | [5 empty] | A.5 race is unknown. the impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp with smart assist system with automated impella controller section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19210 | 1220648-2024-18481 | 20299567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20299567&pc=PYX | 01/26/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 01/26/2022 | 09/25/2024 |  |  | 11/30/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022156775 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention with an intra-aortic balloon pump (iabp) needing an impella rp for mechanical circulatory support. a decision was made to exchange the iabp for an impella cp/rp bipella support due to a worsening global heart failure not responding to current therapy. while on support, there was some persistent oozing present at the venous site after the peel away sheath was removed. a mattress suture was placed and manual pressure was held. however, oozing persisted. in addition, it was noted that they are awaiting partial thromboplastin time to come down. ultimately, care was withdrawn and the patient expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19211 | 1220648-2024-18482 | 20299598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20299598&pc=PYX | 01/30/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 01/30/2022 | 09/25/2024 |  |  | 06/30/2022 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP | [7 empty] | 1485848 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic was implanted with an impella rp device for mechanical circulatory support. while on support three mattress sutures were placed to reduce femoral oozing. additional information received reported the patient expired while on support. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp with smart assist system with automated impella controller section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19212 | 1220648-2024-18483 | 20299668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20299668&pc=PYX | 02/24/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 02/24/2022 | 09/25/2024 |  |  | 02/28/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022052376 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on impella rp support, the patient's leg was cool to touch indicating limb ischemia. the decision was made to place an external femoral-femoral bypass. post bypass, distal flow was confirmed with doppler. additional information noted that care was withdrawn and the patient expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: contraindications (united states). ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ this report is being filed as part of a retrospective review of historical records.
19213 | 1220648-2024-18484 | 20299791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20299791&pc=PYX | 07/26/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 07/26/2022 | 09/25/2024 |  |  | 09/30/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023199476 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. the initial attempt at insertion with a 23 french peel away sheath kinked inside the vessel and the physician was unable to advance the flow tip catheter into the sheath. a new impella rp introducer sheath kit was opened, and the sheath exchanged. placement was successful. the outlet was seen on fluoroscopy four centimeters above the catheter. the wire was removed and impella rp started at p2 and slowly ramped up to p7. a mattress suture was placed. two days later, the decision was made to withdraw care. bipella support was discontinued and the patient expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella rp smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella rp smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records.
19214 | 1220648-2024-18409 | 20299818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20299818&pc=OZD | 08/29/2024 | 10/31/2024 |  | 09/25/2024 |  |  | 08/29/2024 | 09/25/2024 |  |  | 09/30/2025 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445683 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 62-year-old patient developed some oozing from their access site following impella cp removal. heparin was subsequently held to assist with managing the condition. support remained ongoing with an impella 5.5 pump. the patient survived the explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was anticoagulation management due to high patient activated clotting time values. heparin was stopped to achieve hemostasis as per the clinical description. h.6 code 4629 was inadvertently omitted from the initial manufacturer device report number 1220648-2024-18409 and was added.
19215 | 1220648-2024-18485 | 20300153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20300153&pc=PYX | 09/01/2024 | 10/29/2024 |  | 09/25/2024 |  |  | 09/01/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025476538 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 68-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella rp device for mechanical circulatory support. during impella support, the md decided to remove the art line that was next to the impella rp flex. during removal, it was noticed that the suture and yellow suture loop was cut or split allowing the rp sheath to be pulled back resulting in rp catheter to be completely removed from the heart. the patient was taken to the cath lab and rp attempted to be advanced back in place, but it was unsuccessful. the patient began to decompensate, and decision was made to replace rp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of positioning issue was most likely use related. added- h6 impact code 4629.
19216 | 1220648-2024-18486 | 20300181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20300181&pc=OZD | 09/02/2024 | 10/23/2024 |  | 09/25/2024 |  |  | 09/02/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025550600 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support due to chronic decompensated heart failure. while on support, hemorrhage was seen in the right axilla starting near the level of the impella insertion site. two units of blood were given and heparin was paused. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report 1220648-2024-18486.
19217 | 1220648-2024-18488 | 20300191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20300191&pc=OZD | 09/05/2024 | 10/30/2024 |  | 09/25/2024 |  |  | 09/05/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527402 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support due to shortness of breath and non-st elevated myocardial infarction (nstemi). while on support, loss of pulses distally in right leg. the pump was explanted and distal pulses returned. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
19218 | 1220648-2024-18489 | 20300197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20300197&pc=OZD | 09/07/2024 | 10/22/2024 |  | 09/25/2024 |  |  | 09/07/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025501349 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and they had a gastrointestinal bleed (gib). the impella was ran with heparin in the purge prior to the gib and they had systemic heparin. the patient survived the event. | [5 empty] | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ corrections to the initial mfr report: h6 clinical code 1888 changed to 4476. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
19219 | 1220648-2024-18487 | 20300198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20300198&pc=OZD | 09/03/2024 | 02/21/2025 |  | 09/25/2024 |  |  | 09/03/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025477856 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient on impella cp ssupport. it was reported that there was "distorsion" .018 guidewire around the impella cp. there was no noted harm or injury and the patient was escalated to the impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the product damage has been completed. pump was not returned for investigation. data log review was not required for this case. as per the clinical, details, guidewire distortion around the impella cp. the cause of the guidewire damage issue was not determined since product was not returned and sufficient clinical information was not provided.
19220 | 1220648-2024-18491 | 20300205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20300205&pc=OZD | 09/09/2024 | 11/05/2024 |  | 09/25/2024 |  |  | 09/09/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2025537372 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant in the eu had a cp being placed to support a patient when the the pump failed to be detected by the console. the aic (automated impella controller) was not replaced, but the pump was. a new pump was placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the pump not detected has been completed. there was no product returned or data logs as the console serial number is unknown. the root cause of the pump not detected cannot be determined. corrections to the initial mfr report:
19221 | 1220648-2024-18490 | 20300206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20300206&pc=OZD | 09/08/2024 | 11/04/2024 |  | 09/25/2024 |  |  | 09/08/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528927 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an attempted impella 5.5 implant and they could not successfully deliver the pump. the pump would not be implanted due to the patient's anatomy, in which the physician believes could be related to calcification near the innominate artery. the implant was aborted. | [5 empty] | The investigation for the reported delivery issues has been completed. the returned product was visually inspected and no abnormalities were found. the root cause of the delivery issue was most likely patient condition related; the physician believes could be related to calcification near the innominate artery as per the clinical description provided. added- d9 device return date corrections to the initial mfr report: d4-model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
19222 | 1220648-2024-18492 | 20300414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20300414&pc=OZD | 09/02/2024 | 11/01/2024 |  | 09/25/2024 |  |  | 09/02/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025550549 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support developed a hematoma at the impella access site. one unit of red blood cells was given. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
19223 | 20300601 | 20300601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20300601&pc=OZD | 02/25/2024 | 03/21/2024 | 03/21/2024 | 09/25/2024 | 03/21/2024 | 09/25/2024 |  | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 1000100 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 1000100 |  | Device was returned to manufacturer |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 69 YR |  | Male | 90 KG |  |  | Injury | Yes | Yes | [7 empty] | Extravasation | Failure of Implant | [11 empty] | Required Intervention | Hospitalization | [23 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y |  |  |  | Patient received an impella 5.5. after the patient was transferred to the cardiac icu, it was discovered that the inserted impella 5.5 was leaking from the red connector (the junction between the impella implanted device and the power cord). the patient was taken back to the cardiovascular operating room for the removal and placement of a new impella 5.5 device. per the surgeon, "immediately postop in the cicu there was noted to be blood extravasating from the proximal portion of the catheter which is highly abnormal. consultation with manufacture and representative indicated pump disruption which could result in catastrophic failure. for this reason was decided to move the patient immediately back to the operating room for removal of the defective catheter and placement of a new impella 5.5. findings at surgery: patient remained surprisingly stable for the brief period in which there was no support. the new pump was placed in good anatomic position and resultant hemodynamics were as expected with flows at 4.5 on p9 the pump was sent with the manufacture representative for analysis." manufacturer response for temporary non-roller type left heart support blood pump, impella (per site reporter).
19224 | 1220648-2024-18493 | 20300622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20300622&pc=OZD | 09/02/2024 | 11/05/2024 |  | 09/25/2024 |  |  | 09/02/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532276 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 47-year-old male patient with caridomyopathy and other systemic comorbidities was implanted with an impella cp device for mechanical circulatory support. prior to the patient leaving the operating room, a small amount of blood was noted at the red impella plug site. after the patient arrived to the unit, the amount of blood increased and cloth was placed around the red plug to absorb the blood. patient was given 1 unit of packed red blood cells, 1 unit of platelets (has been thrombocytopenic since admission) followed by fresh frozen plasma (ffp) an hour later. after ct scan, the bleeding from the red plug appeared to have stopped. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root caused of the access site bleeding was most likely patient condition since the bleeding resolved on its own without any treatment or intervention.
19225 | 1220648-2024-18494 | 20300972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20300972&pc=OZD | 09/03/2024 | 11/06/2024 |  | 09/25/2024 |  |  | 09/03/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508036 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had placed a 5.5 pump to support a patient with history of postpartum cardiomyopathy in need of transplant. the 5.5 was placed as the patient awaited transplant. after over one month of support the pump cut down site had a bleed that prompted two units of blood. the pump also had high purge pressures remedied by adding tissue plasminogen activator to the purge and automated impella controller swapping. the pump remained on for support for seven more days, and the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of access site bleeding and high purge pressure has been completed. the impella device was not returned for evaluation. based on data logs and clinical details, the root cause of the high purge pressure was determined to most likely be biomaterial. the root cause of the access site bleed could not be determined due to lack of clinical detail and returned product.
19226 | 1220648-2024-18495 | 20301066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20301066&pc=PYX | 01/14/2022 | 05/21/2025 |  | 09/25/2024 |  |  | 01/14/2022 | 09/25/2024 |  |  | 01/31/2022 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1468332 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella rp device for mechanical circulatory support. while on support, the patient was taken to unit due to surgical bleeding. an echocardiogram was utilized to confirm impella rp placement, and the pump speed level was increased to fill the left ventricle. the patient¿s chest was opened at bedside to rule out constriction, and the left ventricle was reported to be functioning as normal. the device was explanted, care was withdrawn. survival outcome at explant noted the patient expired. medical safety review of the event noted that the use of the device cannot conclusively be associated with the outcome. the patient¿s peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19227 | 1220648-2024-18496 | 20301170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20301170&pc=OZD | 09/03/2024 | 10/15/2024 |  | 09/25/2024 |  |  | 09/03/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519238 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and they lost pulses in their leg. the impella was explanted. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
19228 | 1220648-2024-18497 | 20301286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20301286&pc=PYX | 04/27/2021 | 09/27/2024 |  | 09/25/2024 |  |  | 04/27/2021 | 09/25/2024 |  |  | 02/28/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022043149 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella rp device for mechanical circulatory support. while on support, the patient was bleeding and not doing well. the rp was later explanted due to low flows. at explant the rp had biomaterial present and observed at the cannula and they did not replace the pump. | [5 empty] | This report is being filed as part of a retrospective review of historical records. h10: remediation statement was not included in manufacturer device report 1220648-2024-18497. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp flex with smartassist instructions for use & clinical reference manual instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19229 | 1220648-2024-18498 | 20301319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20301319&pc=OZD | 08/17/2024 | 10/07/2024 |  | 09/25/2024 |  |  | 08/30/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025490847 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured sepsis with a "possible" relationship to the impella 5.5 device used: ¿white blood celss (wbc) fluctuated and purulent drainage out of impella site on (b)(6) 2024. blood culture from august 19 was negative. superficial wound culture was obtained from impella site on (b)(6) 2024 which showed staph epidermidis in culture (moderate wbc on gram stain). ct chest and upper extremity on (b)(6) 2024 showed subcutaneous air locules. vanco was given, and infection disease (id) following. wound vac was placed on impella site on (b)(6) 2024¿ it was reported that the patient/event has not recovered and issue is not resolved. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | G.2 study selection was inadvertently omitted from the initial manufacturer device report number 1220648-2024-18498 and was added.
19230 | 1220648-2024-18499 | 20301378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20301378&pc=OZD | 08/26/2024 | 11/19/2024 |  | 09/25/2024 |  |  | 09/04/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025493947 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 has a catheter purge side arm and purge yellow connector leak. reconnecting the connector did not improve the leak, so a three-way stopcock was connected to improve the situation. there was no reported harm to the patient. | [5 empty] | The investigation for the pump-purge leak has been completed. the pump was evaluated and there were no physical abnormalities. the pump was run and the leak did not reproduce in the lab. the data logs showed a 'purge line click-on not detected' alarm and a purge system open alarm. the root cause of the purge leak - pump was most likely due to a use-related issue. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
19231 | 1220648-2024-18500 | 20301446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20301446&pc=OZD | 08/03/2024 | 10/07/2024 |  | 09/25/2024 |  |  | 08/26/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525284 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that was on impella cp support. during support, it was noted that the patient experienced a pulmonary hemorrhage that required suctioning but the hemorrhage persisted. the patient expired when care was withdrawn by the family. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | G.2 study selection was inadvertently omitted from the initial manufacturer device report number 1220648-2024-18500 and was added.
19232 | 1220648-2024-18501 | 20301544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20301544&pc=OZD | 09/05/2024 | 11/04/2024 |  | 09/25/2024 |  |  | 09/05/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1606415 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the the automated impella controller (aic) displayed yellow ¿controller error¿ and lost aortic (ao) placement signal during impella 5.5 support. this error informed the user to switch to backup controller. patient tolerated switch to backup controller. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the console (aic) controller alarm (yellow alarm) has been completed. console logs from the reported day of event are consistent with complaint and show pump was disconnected from the console twice during case start, and after third connection operating correctly for less than an hour. at this point data from optical bench was no longer available, and a controller error alarm (#1043, opm communication stopped) was triggered. the logs also reported processsamplingdatafromopticalbench: sentstopacquisitioncmd ack, indicating ossiopsens receive thread gets a data sampling packet with an unexpected length. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are lastly reported as +16384 values and then no more samples were recorded. the absence of snr samples impacted the console¿s ability to recognize the pump disconnection and required the user to perform a hard shutdown. pump was then transferred to ic1027 where no issues with placement signal occurred. the console was returned for evaluation and was tested on an engineering lab bench. the failure mode was not reproduced while running an optical pump. there was no issue with the console hardware. the root cause of the controller error was due to an aic software issue.
19233 | 1220648-2024-18502 | 20301579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20301579&pc=PYX | 02/22/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 02/22/2022 | 09/25/2024 |  |  | 12/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023161419 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella rp for mechanical circulatory support. while on support, the purge was switched from d5 with 50 units of heparin to sodium bicarb due to drain to control bleeding. the patient received 13 packed red blood cells, 6 plasma, 2 platelets, and 2 cryo. in addition, there was oozing from the hemostatic valve as it appeared to be sunken in. the oozing from the oscor sheath and bleeding from chest were managed by the blood products, gauze and pressure dressing. the patient made a turn for the worse with the kidneys shutting down. the family withdrew care and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19234 | 1220648-2024-18503 | 20301628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20301628&pc=PYX | 03/24/2022 | 05/19/2025 |  | 09/25/2024 |  |  | 03/24/2022 | 09/25/2024 |  |  | 03/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 2022066545 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support after experiencing right heart failure following a right coronary artery dissection during a coronary intervention. it was reported while on impella rp support, the patient experienced bleeding from the impella access site. manual compression was held and the patient was given blood as a result of the blood loss. it was also reported the patient experienced gastrointestinal bleeding while on impella cp support. the patient was unstable and in overall very critical condition with an ejection fraction <5%. it was noted the patient died of multiple organ failure while on impella cp support. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19235 | 1220648-2024-18504 | 20301646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20301646&pc=PYX | 04/26/2021 | 09/26/2025 |  | 09/25/2024 |  |  | 04/26/2021 | 09/25/2024 |  |  | 06/30/2022 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1485848 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic needing an impella rp device for mechanical circulatory support. while on support the patient was experiencing some bleeding. the purge was cut down to address the issue. additional information received reported the patient expired while on support. medical safety review of the event noted there is not enough evidence to exclude impella as an associated factor in the patient's outcome. patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19236 | 1220648-2024-18505 | 20301736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20301736&pc=OZD | 08/19/2024 | 10/07/2024 |  | 09/25/2024 |  |  | 09/01/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525967 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "probable" relationship to the impella cp device used. the patient was also noted to have down trending hemoglobin, decreasing platelets with elevated hemolysis markers in the setting of having the impella at p9. the impella remained at high settings, and the patient began to have significant hemolysis and vasopressors were fluctuating therefore the impella was removed. it was further reported that the patient endured occlusion of left superficial femoral artery with a "probable" relationship to the impella device used. after the impella was removed the patient was found to have occlusion of the left superficial femoral artery. a catheter was placed, and the patient received 24 hours of tissue plasminogen activator. a vascular surgery was consulted as the patient was not having any distal pedal pulses on the left leg but was deemed to high risk. the patient was retaken to the catheterization lab to evaluate and remove the femoral sheath. the patient recovered with sequelae. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | G.2 study selection was inadvertently omitted from the initial manufacturer device report number 1220648-2024-18505.
19237 | 1220648-2024-18506 | 20301776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20301776&pc=OZD | 09/05/2024 | 11/04/2024 |  | 09/25/2024 |  |  | 09/05/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1162563 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the impella 5.5 set up, there were purge system open alarms during set up. the purge cassette was replaced to no avail. the automated impella controller (aic) and the pump moved to another aic and alarm resolved. there was no harm reported. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the console (aic) purge issues has been completed. console logs from the reported day of event are partially consistent with complaint and show sudden drop of purge pressure during case start, resulting in purge pressure low alarm (#431) and purge system open (#407). one of the alarms (#431) resolved on the same console, while the other (#407) remained unresolved. however, the real time log data shows that before pump was unplugged, purge pressure was increasing to the point that triggering criteria for the alarms would have been removed. the console was returned for evaluation and the issue was not reproduced. testing did not reveal any anomalies, and the unit¿s purge system operated within specification. the purge pressure drop and resulting purge alarms are captured under mfr report #1220648-2024-18750. added- h6 impact code 4629.
19238 | 1220648-2024-18462 | 20301968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20301968&pc=PYX | 07/25/2021 | 05/27/2025 |  | 09/25/2024 |  |  | 07/25/2021 | 09/25/2024 |  |  | 10/31/2021 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 1460452 |  | 004334 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. the patient¿s pedal pulses were not able to be dopplered prior to impella insertion. overnight the patient¿s right leg became cold and cyanotic. the resolution for this issue is unknown. additionally during support, the device exhibited high purge pressure. the resolution for this issue is unknown. survival outcome at explant noted the patient expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. | [5 empty] | A.5 ethnicity is unknown. the impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ impella rp smart assist system with automated impella controller section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (eg, infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19239 | 1220648-2024-18507 | 20302069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20302069&pc=PYX | 12/04/2021 | 09/25/2024 |  | 09/25/2024 |  |  | 12/04/2021 | 09/25/2024 |  |  | 02/28/2022 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1469832 |  | 004334 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [9 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported while on impella rp support, the patient began to bleed profusely from all access sites. the patient also went into ventricular tachycardia and the patient's blood pressure dropped. a large hematoma was observed at the impella access site. the patient then began to code for over 30 minutes and cardiopulmonary resuscitation (cpr) was initiated. the patient's hemoglobin was reported to be 7.2 and rapid infusion protocol was initiated. additional mattress sutures were attempted to be placed during cpr to achieve hemostasis, but the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia. this report is being filed as part of a retrospective review of historical records.
19240 | 1220648-2024-18508 | 20302179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20302179&pc=PYX | 01/19/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 01/19/2022 | 09/25/2024 |  |  | 10/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022143574 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella rp device for mechanical circulatory support. while on support, additional purse-string sutures were placed on exposed vein at cut down site to control bleeding. cryo was also ordered to control the bleeding. the patient was relatively stable before experiencing a sudden systolic cardiac arrest. the patient was a do-not-resuscitate and was not resuscitated. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ potential adverse events (united states). arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia. this report is being filed as part of a retrospective review of historical records.
19241 | 1220648-2024-18509 | 20302225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20302225&pc=PYX | 01/30/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 01/30/2022 | 09/25/2024 |  |  | 11/30/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022147052 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  | Black Or African American | Death | Yes | No | Increase in Pressure | [6 empty] | Renal Failure | Thrombosis/Thrombus | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support. the patient had low ejection fraction but due to the fresh valve it was decided for an intra-aortic balloon pump would be the best option. upon evaluating the patient bedside, a clot was found. the patient had a dialysis catheter placed and had to be put on continuous dialysis. the flows had dropped to 1.9 with a purge pressure in the upper 1100s. the patient started to brady down, however, purge pressure shot back up. the patient was then made a do-not-resuscitate and expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19242 | 1220648-2024-18513 | 20302430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20302430&pc=OZD | 09/01/2024 | 10/15/2024 |  | 09/25/2024 |  |  | 09/01/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534711 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 64-year-old female patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the impella cp implant, the patient had access site bleeding. the patient was bleeding from the initial access attempt to deep circumflex iliac artery. the 14 french sheath was left in place to remain occlusive in the artery to occlude and tamponade the deep circumflex iliac artery. as soon as 14 french sheath was removed to attempt to perclose the vessel, the patient became profoundly hypotensive while pressure was being applied. ultimately, the 14 french sheaths flushed reinserted and the doctor chose to flush the impella catheter in a bowl of saline and reinsert the impella. the impella was advanced over the 0.018¿ wire into the left ventricle and slow ramped up. during this time it was noted that hemoglobin dropped from 10 to 7. an emergency blood transfusion was started. the patient was stabilized and support was continued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
19243 | 1220648-2024-18514 | 20302449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20302449&pc=PYX | 07/17/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 07/17/2022 | 09/25/2024 |  |  | 01/31/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023173476 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella rp device for mechanical circulatory support. while on support, the patient had significant bleeding from ecmo sites. systemic heparin was off, and purge solution was switched to sodium bicarbonate. patient was transfused 1 unit packed red blood cells (prbc), 4 fresh frozen plasma (ffp), 2 cryoprecipitate antihemophilic factor, (cryo) and multiple platelets. additional information noted that patient expired the next day. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
19244 | 1220648-2024-18515 | 20302497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20302497&pc=PYX | 07/27/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 07/27/2022 | 09/25/2024 |  |  | 08/31/2022 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1498341 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | Hemodynamic instability | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock¿and with low cardiac output syndrome was attempted to be implanted with an impella rp for mechanical circulatory support. it was reported that after the patient was weaned from extracorporeal membrane oxygenation (ecmo) and was attempted to be placed on impella rp. the physician experienced great difficulty accessing and placing a wire into the patient's vessels due to the patient's anatomy. the physician attempted multiple times using different wires, and was able to get a amplatz super stiff wire to cross with great effort. the physician then attempted to advance the impella rp over wire, but again experienced great difficulty and reportedly used a lot of force. the impella catheter buckled, so the impella rp was removed and the physician placed a gore dryseal to reattempt placement, but the patient became hemodynamically unstable. the impella rp was aborted, and the patient was placed back on ecmo. it was reported after the patient was placed back on to ecmo, there were no flows and all cardiovascular function ceased. further efforts of patient care was terminated and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ this report is being filed as part of a retrospective review of historical records.
19245 | 1220648-2024-18517 | 20302686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20302686&pc=OZD | 08/30/2024 | 12/03/2025 |  | 09/25/2024 |  |  | 08/30/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1678477 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller had a broken optical cover. the biomedical engineer noted that the port had been snapped off and pushed in. the device was removed from service as a result of the noted damage. there was no patient harm. | [5 empty] | The automated impella controller was received for evaluation. upon investigation closure, a supplemental manufacturer device report will be filed.
19246 | 1220648-2024-18510 | 20302760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20302760&pc=PYX | 04/15/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 04/15/2022 | 09/25/2024 |  |  | 07/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022125440 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella rp device for mechanical circulatory support. while on support, it was noted that access site had oozing. hemostasis was achieved by putting another mattress suture. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
19247 | 1220648-2024-18511 | 20302800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20302800&pc=PYX | 05/29/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 05/29/2022 | 09/25/2024 |  |  | 12/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2023160836 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock infarction was implanted with an impella rp device for mechanical circulatory support. it was reported while on impella rp support, the patient experienced bleeding from the intra-aortic balloon pump site and intermittent oozing from the impella access site. the patient received multiple units of packed red blood cells and platelets as a result of the blood loss. it was also reported that the patient experienced heparin induced thrombocytopenia. as a result heparin was discontinued, and the patient was started on argatroban for anticoagulation therapy. the patient then developed a gastrointestinal bleed. the patient continued to decline, and it was reported the patient expired while on impella rp support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia. this report is being filed as part of a retrospective review of historical records.
19248 | 1220648-2024-18512 | 20302937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20302937&pc=PYX | 07/18/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 07/18/2022 | 09/25/2024 |  |  | 03/31/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023210713 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella rp device for mechanical circulatory support. while on support, patient had bleeding that later resolved, but partial thromboplastin time was subtherapeutic, so heparin was noted added to the system due to previous bleeding risk. additionally, patient experienced hemolysis. the patient was later intubated, no change in drips and started patient on continuous renal replacement therapy (crrt). the patient later expired on support. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
19249 | 1220648-2024-18516 | 20303000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20303000&pc=PYX | 07/29/2022 | 04/03/2025 |  | 09/25/2024 |  |  | 07/29/2022 | 09/25/2024 |  |  | 04/30/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023218032 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 22 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Thrombocytopenia | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella rp device for mechanical circulatory support. while on support, patient experienced thrombocytopenia. due to suspected heparin-induced thrombocytopenia, staff was running g5 with sodium bicarbonate and systemic argatroban. it was also reported that the chief physician believed the patient was experiencing sepsis or worsening cardiogenic shock. urgent need to escalate care to bipella or va-ecpella was discussed. three days later, while it is stated that the impella rp was running without any issues, the patient expired due to multi-organ failure. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella rp smart assist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
19250 | 1220648-2024-18519 | 20303045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20303045&pc=OZD | 09/07/2024 | 11/08/2024 |  | 09/25/2024 |  |  | 09/07/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520971 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | UNKNOWN. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was not able to get flows above 1l/min. the therapeutic activated clotting time and correct positioning were confirmed. impella cp was removed and a new impella was placed. there was no patient harm. | [5 empty] | The investigation for the low pump flow has been completed. data logs were reviewed which confirmed the failure mode and clinical narrative. logs showed after pump started and run for 42 minutes, motor current (mc) spiked followed pump stop but could restart and low flow that triggered impella stop-mc high, auto suction response and suction alarms. this event remained to the rest of the case. visual inspection found the fracture on both impeller blades. no other issue was found on the rest of the pump. the cause of the low flow was fractured impeller blades. the fracture was characteristic of an interaction with another device but that was unable to be definitively determined since fluoroscopic imaging and/or additional clinical details were not provided. the cause of temporary the pump stop was most likely an interaction with another device. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
19251 | 1220648-2024-18518 | 20303206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20303206&pc=OZD | 09/06/2024 | 11/01/2024 |  | 09/25/2024 |  |  | 09/06/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1402264 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella support and the automated impella controller (aic) had a charging issue. the aic plug was not recognizing outlets when the console was plugged in and notification stated it was running on battery. multiple outlets were tried. the aic was replaced. no patient harm has been reported. | [5 empty] | The investigation for the console (aic) battery charging issues has been completed. data logs were reviewed finding multiple instances of ac power connections and disconnections throughout the case. additionally, "ac power disconnected - alarm issued" event (#109) occurred, which remained unresolved and the user requested a shutdown. the console was returned for evaluation finding that reported failure mode was successfully reproduced. console was connected to ac and no power was detected (ac power disconnected - alarm issued" event # 109), and console was operating on battery power. the root cause for the ac power not being recognized was traced to a broken ac cord socket, which had damaged pins. added- h6 impact code 4629. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
19252 | 1220648-2024-18520 | 20303245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20303245&pc=OZD | 06/20/2023 | 11/04/2024 |  | 09/25/2024 |  |  | 08/30/2024 | 09/25/2024 |  |  | 02/28/2025 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024332416 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hypovolemia | Renal Impairment | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient with a two-day impella cp support, shortly after placement the impella began to have suction events. it was found the impella had migrated near the mitral valve. this led to further hemodynamic instability in the patient who was already suffering from severe cardiac dysfunction. this further cardiac dysfunction caused hypotension, resulting in poor end organ perfusion, further damaging the patient already injured kidney and liver. the patient also experienced hemolysis due to the impella device. additionally, there was significant bleeding from the impella access site. this led to severe hypotension. the patient was transfer to a higher level of care and continued to hemorrhage. when the impella was removed, the patient was found to have a large pseudoaneurysm in the right axillary artery caused by the impella. the patient underwent additional intervention days later to fix the pseudoaneurysm. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the positioning issues, hemolysis, and the access site bleed has been completed. data logs showed the pump ran without issue during support. there were suction alarms and impella stop alarm was triggered during the case but resolved quickly. no positioning alarm was triggered. the product was not returned for review. the root cause of positioning issue was most likely use related. the root cause of access site bleeding was most likely anticoagulation management related. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
19253 | 1220648-2024-18521 | 20303436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20303436&pc=PYX | 01/31/2022 | 05/21/2025 |  | 09/25/2024 |  |  | 01/31/2022 | 09/25/2024 |  |  | 02/28/2022 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1469832 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella rp device for mechanical circulatory support. following insertion of impella rp pump, the decision was made to also implant impella cp despite iliacs not being able to accommodate sheath given poor cardiac function. it was noted that once impella rp and cp were in place, a lot of blood was coming from the endotracheal tube and physician noted a hole in the pulmonary artery. this was repaired and the patient was weaned off cardiopulmonary bypass (cpb) with both the rp and cp flowing nearly 3.5 l/min each. the patient continued to have significant bleeding from the chest. lots of blood products were given including factor vii without much resolution in the bleeding. because of the bleeding, the rp introducer sheath was left in place for fear of bleeding from the groin as well. the patient was transferred to the icu where the prognosis was very poor. there were no pulses to the right lower extremity, and it was beginning to mottle. the patient continued to bleed post-op and stopped responding to the treatment. survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella rp smart assist system with automated impella controller (limb ischemia) section: contraindications (united states). ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ impella rp with smart assist system with automated impella controller (bleeding) section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19254 | 1220648-2024-18522 | 20303584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20303584&pc=PYX | 02/15/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 02/15/2022 | 09/25/2024 |  |  | 11/30/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022153847 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with pulmonary embolism was attempted to be implanted with an impella rp device for mechanical circulatory support. dring attempted implant, impella rp had delivery issues. the patient went into pulseless electrical activity (pea). cardiopulmonary resuscitation (cpr) was done for 45 minutes with attempted rp insertion. impella rp pump was unable to pass into pulmonary artery (pa) via. 0.027 wire even with the use of buddy wire. multiple attempts were made but after doing cpr for 45 minutes, the decision was made to stop the treatment. survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella rp smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
19255 | 1220648-2024-18523 | 20303672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20303672&pc=PYX | 02/20/2022 | 11/23/2024 |  | 09/25/2024 |  |  | 02/20/2022 | 09/25/2024 |  |  | 05/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022103164 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 12 MO |  | Male |  |  | White | Death | Yes | Yes | Increase in Pressure | Failure to Advance | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was additionally reported that after the impella rp was successfully placed, the outflow cage remained low in the pulmonary artery. the physician was unable to advance the pump further without risk for buckling in the right ventricle. the following day on (b)(6) 2024, transesophageal echocardiography was performed and proper position of the impella rp was verified. it was noted that the patient arrested, the patient's chest was reopened but the patient expired. there is an unknown relationship between the impella rp and the patient's outcome. | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella rp device for mechanical circulatory support. while on suppport, impella rp had high purge flow. purge flow remained high, and nurse was instructed on purge cassette changes if there was a leak found. purge was changed to dextrose 10% due to low purge alarm. survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ additional information for the initial mfr 1220648-2024-18523 was provided in sections b5 and h6. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
19256 | 1220648-2024-18524 | 20303699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20303699&pc=PYX | 03/30/2022 | 05/20/2025 |  | 09/25/2024 |  |  | 03/30/2022 | 09/25/2024 |  |  | 03/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2022073154 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. during support, the patient suffered an embolism, began to decompensate and the decision was made to place an impella cp for additional support. then an air embolus occurred, and patient decompensated, the echocardiogram showed that the right ventricle was still functioning, however patient continued to code. survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19257 | 1220648-2024-18526 | 20303801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20303801&pc=PYX | 04/04/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 04/04/2022 | 09/25/2024 |  |  | 01/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2022049584 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | American Indian Or Alaskan Native | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, the patient had multiple episodes of ventricular tachycardia and was shocked twice. survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp system with automated impella controller. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records.
19258 | 1220648-2024-18525 | 20303853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20303853&pc=OZD | 05/05/2021 | 10/15/2024 |  | 09/25/2024 |  |  | 08/30/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Company representation | [5 empty] | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhagic Stroke | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the morning after impella 5.5 was placed, the patient suffered a stroke. a computed tomography scan of the brain showed a large subarachnoid hemorrhage most prominent in the basilar cisterns extending into the cervical spine, large intraventricular hemorrhage, and multiple small likely acute to subacute infarcts involving the supratentorial and infratentorial brain. the patient was unable to recover and expired two days after the impella was placed. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. | A.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. instructions for use state the following: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
19259 | 1220648-2024-18527 | 20303933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20303933&pc=PYX | 04/19/2022 | 05/22/2025 |  | 09/25/2024 |  |  | 04/19/2022 | 09/25/2024 |  |  | 06/30/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2022109648 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Hispanic | American Indian Or Alaskan Native | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, the device exhibited low flow. action taken is unknown. survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19260 | 1220648-2024-18528 | 20303936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20303936&pc=PYX | 07/02/2021 | 05/22/2025 |  | 09/25/2024 |  |  | 07/02/2021 | 09/25/2024 |  |  | 11/30/2021 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1467630 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support the patient kept going into ventricular fibrilation. the patient was shocked multiple times and was moved to intensive care unit. the patient coded again during the move to the unit. in the unit the patient had significant bleeding out of the chest tube. heparin in purge was reduced. additional information noted care was withdrawn and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp system with automated impella controller & impella rp flex with smartassist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19261 | 1220648-2024-18529 | 20303960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20303960&pc=PYX | 05/25/2022 | 05/21/2025 |  | 09/25/2024 |  |  | 05/25/2022 | 09/25/2024 |  |  | 01/31/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2023179111 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during delivery of the device lactate was noted to be 14, hemoglobin dropped to 7. the medical staff initiated a massive blood transfusion protocol, and 3 units were administered. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19262 | 1220648-2024-18530 | 20304184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304184&pc=PYX | 06/04/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 06/04/2022 | 09/25/2024 |  |  | 12/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2022156317 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Sepsis | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, the patient experienced atrial fibrillation and was administered amiodarone. the patient continued to experience atrial fibrillation after the medication was administered. additionally, the patient's cultures were positive for sepsis and antibiotics were administered. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp smart assist system with automated impella controller. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer. ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
19263 | 1220648-2024-18532 | 20304240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304240&pc=PYX | 11/09/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 11/09/2022 | 09/25/2024 |  |  | 08/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022131255 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support after decompensating post aortic valve procedure. while on impella rp support, the patient experienced oozing at the impella access site. as a result, heparin was removed from the purge solution. it was also reported that the patient's urine output decreased, and they were placed on continuous renal replacement therapy. it was noted that the patient expired while on impella rp support. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. "be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open." "make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound." this report is being filed as part of a retrospective review of historical records.
19264 | 1220648-2024-18535 | 20304314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304314&pc=PYX | 06/22/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 06/22/2022 | 09/25/2024 |  |  | 04/30/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2023208355 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support it was reported that the patient had a stroke. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
19265 | 1220648-2024-18537 | 20304335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304335&pc=PYX | 01/27/2022 | 05/21/2025 |  | 09/25/2024 |  |  | 01/27/2022 | 09/25/2024 |  |  | 10/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022143589 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. while on support, the patient with an impella 5.5 that was inserted in the axillary, and had the impella cp removed and repaired surgically, also had the impella rp inserted and started. there were excellent flows on p9, and patients' blood pressure improved immediately, although swan numbers remained unchanged. during peel away, the pump dislodged and was unable to advance. no additional impella rp sheaths to be found, so the pump was removed, and mattress suture and pressure was held. survival outcome at explant noted the patient expired. medical safety review of the event noted that there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19266 | 1220648-2024-18531 | 20304399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304399&pc=OZD | 08/26/2024 | 10/29/2024 |  | 09/25/2024 |  |  | 08/26/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025457839 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Vascular Dissection | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the impella 5.5 implant procedure for 67-year-old female patient, the 5.5 was being used to cross the aortic valve without the wire and the innominate artery was dissected. the right axillary access was aborted. the doctor opted to place the impella direct aortic with full sternotomy. the patient received 4 units of blood products, 1 platelet and 1 liter of cell saver. additionally, once on support the pump could not be repositioned and the flows were suboptimal. the patient expired after 13.95 hours of support. the relationship between the cause of death and the impella is unknown. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci & impella 5.5 with smartassist for use during. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the vascular damage, low pump flow, and positioning issues has been completed. no product was returned for evaluation. clinical details noted that physician denied without ultrasound or angiogram for right axillary artery measurement, attempted to place diagnostic catheter and guidewires without fluro and only using echo and was unsuccessful in crossing av with catheter and wires after multiple attempts. also attempted to cannulate aorta and insert impella 5.5 across av without a wire, resulting in dissection of innominate artery later found on angiogram. using a wire for impella 5.5 axillary insertion is recommended per ifu 10003049. right axillary artery access was aborted and direct aortic access was obtained and pump was able to be placed without incident. pump position was unable to be properly identified due to patient anatomy and was not clear on position of inlet to av. impella flows also not able to flow accurately as position of inlet was not clear. once patient was taken to icu, pump was found to be deep in papillary muscle and was unable to be re-positioned. data logs from case were reviewed and live time data logs show persistent "impella position unknown" alarms with fluctuating placement signal. additionally, pump flow also was very variable with mostly being within the proper flow range for p-level. the root cause of the vascular damage - peripheral vessel was due to a use issue as physician damaged vessel when attempting to insert impella 5.5 wirelessly. the root cause of the low pump flow was due to positioning issues as the pump was positioned deep in papillary muscle. the root cause of the positioning issues was most likely due to a use issue as pump position was not confirmed after initial placement.
19267 | 1220648-2024-18533 | 20304425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304425&pc=PYX | 06/17/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 06/17/2022 | 09/25/2024 |  |  | 12/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2023160804 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. during support, the patient went into ventricular fibrillation and the medical staff performed cardiopulmonary resuscitation (cpr) and the patient coded for 35 minutes without achieving return of spontaneous circulation (rosc). survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp system with automated impella controller & impella rp flex with smartassist system with automated impella controller instructions for use & clinical reference manual. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records.
19268 | 1220648-2024-18536 | 20304435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304435&pc=PYX | 12/02/2021 | 05/21/2025 |  | 09/25/2024 |  |  | 12/02/2021 | 09/25/2024 |  |  | 03/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2022070440 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the physician repaired several areas of bleeding. once bleeding was under control, the patient was weaned off bypass as impella 5.0 and rp flows were increased. the rp to p6 with flows of 3.1 l/min without suction. the patient was watched for the next hour, but very little respiratory effort, heart movement or sustainable blood pressure were noted. additional information noted care was withdrawn and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19269 | 1220648-2024-18540 | 20304461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304461&pc=OZD | 07/15/2024 | 09/25/2024 |  | 09/25/2024 |  |  | 07/16/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025507790 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella cp support, there was a leak at the purge arm at the yellow luer lock. the staff performed a purge sidearm bypass and cleaned the sidearm retainer. the purge cassette was removed and reinserted and the event resolved successfully. no patient harm has been reported. | [5 empty] | A.1, a.4 and a.5 are unknown. the impella device was returned and evaluated. visual inspection found a crack on the luer. the leak was confirmed from the crack of the yellow luer of the pss under leak test. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
19270 | 1220648-2024-18541 | 20304464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304464&pc=PYX | 10/14/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 10/14/2022 | 09/25/2024 |  |  | 02/29/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023184994 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to being implanted with an impella rp device for mechanical circulatory support. during implantation, the impella rp was not advanced to the left pulmonary artery (lpa). attempts for transcutaneous and trans venous pacing were unsuccessful, so no further attempts were made. survival outcome at explant noted the patient expired. medical safety review of the event noted that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller instructions for use for the related event are as follows: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
19271 | 1220648-2024-18545 | 20304520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304520&pc=PYX | 04/19/2022 | 05/14/2025 |  | 09/25/2024 |  |  | 04/19/2022 | 09/25/2024 |  |  | 01/31/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023167665 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, the rp catheter was unable to pass through the pulmonary artery (pa), as the left pa was being compressed by the mass and also compressing the main pa not allowing the catheter to pass to the right pulmonary artery (rpa) that was full of clot. survival outcome at explant noted the patient expired. medical safety review of the event noted that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19272 | 1220648-2024-18542 | 20304549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304549&pc=PYX | 12/08/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 12/08/2022 | 09/25/2024 |  |  | 08/31/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023258573 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella rp device for mechanical circulatory support. during attempted implant, impella rp pump had delivery issues. sheath kinked and had significant difficulty attempting to advance into pulmonary artery (pa) despite buddy wire. catheter was also discovered to be kinked. survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella rp smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
19273 | 1220648-2024-18538 | 20304593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304593&pc=PYX | 12/23/2021 | 05/21/2025 |  | 09/25/2024 |  |  | 12/23/2021 | 09/25/2024 |  |  | 11/30/2022 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2021019101 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | User facility |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support of rp the patient was critically ill with an open chest and requiring multiple blood products, with primary hemodynamic influence being hypovolemia. chest tubes have outputs > 3 liters. the patient was administered two units of packed red blood cells and platelets. staff reviewed reduction in purge heparin. additional information noted care was withdrawn and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19274 | 1220648-2024-18548 | 20304673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304673&pc=PYX | 11/11/2022 | 05/21/2025 |  | 09/25/2024 |  |  | 11/11/2022 | 09/25/2024 |  |  | 08/31/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2023253338 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. during support, the device exhibited high purge issues, and the medical staff was able to decrease the pressure successfully. additionally, a short ramp study was performed, and the patient became hypotensive and experienced ventricular fibrillation. the resolution for this issue is unknown. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp with smartassist system. section: using the automated impella controller with the impella rp with smartassist. system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19275 | 1220648-2024-18534 | 20304674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304674&pc=PYX | 10/13/2021 | 05/21/2025 |  | 09/25/2024 |  |  | 10/13/2021 | 09/25/2024 |  |  | 01/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022049588 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced chest bleeding and tamponade requiring return to the operating room (or) for wash out. the chest was left open and an additional chest tube was placed. one unit of blood products were administered. operating room repair was performed. the nurse called to confirm that the heparin could be as low as 6.25u/ml. the physician wanted it reduced. during support, the patient was receiving antibiotics and was administered 9.5 units of blood. the patient went back to operating room for attempt to close. additional information noted care was withdrawn and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19276 | 1220648-2024-18547 | 20304698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304698&pc=PYX | 12/08/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 12/08/2022 | 09/25/2024 |  |  | 06/30/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP SMART ASSIST |  |  | ABIOMED, INC. |  |  | US |  | 2023222769 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Ischemia | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post-op cardiothoracic surgery was implanted with an impella rp device for mechanical circulatory support. while on support, impella rp had ongoing high purge pressure alarm. guidance was given to assess the system for any kinks or issues with tubing. no issues were found. it was discovered that tissue plasminogen activator (tpa) was previously run through the purge due to a previous high purge pressure. it was discussed to repeat the process again and advised to obtain another bag of tpa. additionally, it was noted that patient's liver was failing, and patient was put on continuous renal replacement therapy [crrt]. it was also noted that patient's bilateral lower extremity pulses were absent indicating limb ischemia. survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella rp smart assist system with automated impella controller (limb ischemia) section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ impella rp with smartassist system (high purge pressure). section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
19277 | 1220648-2024-18539 | 20304723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304723&pc=OZD | 09/20/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 09/20/2022 | 09/25/2024 |  |  | 03/31/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207148 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced bleeding from an unknown source for which four units of packed red blood cells were required. at the time of the bleed, the patient's hemoglobin dropped from 9.4 to 6.5. the cause for the bleed could not be determined and further treatment was not possible due to the patient's unstable condition. it was reported care was withdrawn, and the patient expired within 10 minutes of withdrawal of care. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
19278 | 1220648-2024-18546 | 20304798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304798&pc=PYX | 12/18/2021 | 09/25/2024 |  | 09/25/2024 |  |  | 12/18/2021 | 09/25/2024 |  |  | 04/30/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2022087533 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. during support, the device exhibited suction issues which indicates low flow. additionally, the patient experienced hemolysis. the resolution for both issues is unknown. the procedural outcome noted that the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
19279 | 1220648-2024-18543 | 20304872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304872&pc=PYX | 02/24/2022 | 05/22/2025 |  | 09/25/2024 |  |  | 02/24/2022 | 09/25/2024 |  |  | 09/30/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2022137125 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, the device exhibited suction issues which indicate low flow. the medical staff administered albumin and infused blood; however, the problem was not resolved. the patient also experienced bleeding from an unknown location, the resolution for this issue is unknown. the procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19280 | 1220648-2024-18549 | 20304912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304912&pc=PYX | 03/03/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 03/03/2022 | 09/25/2024 |  |  | 11/30/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2022152178 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that the impella rp was placed was placed using best practice techniques. accessed in the left femoral vein (lfv). the peel away sheath was left in place due to concerns of oozing at the insertion site. the sheath was secured with sutures and a tegaderm. the procedural outcome noted that the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type.
19281 | 1220648-2024-18550 | 20304915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304915&pc=PYX | 06/06/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 06/06/2022 | 09/25/2024 |  |  | 01/31/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023167841 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during implantation, the impella rp was inserted to right femoral vein (rfv). the pump was taken out twice after the implant failed. proper position was verified by fluoroscopy and the physician was pleased with position. swan placed after impella inserted and the physician decided to leave peel away sheath in place. survival outcome at explant noted the patient expired. medical safety review of the event noted that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ this report is being filed as part of a retrospective review of historical records.
19282 | 1220648-2024-18552 | 20304941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20304941&pc=PYX | 02/28/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 02/28/2022 | 09/25/2024 |  |  | 02/28/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2022124663 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella rp pump for mechanical circulatory support. following implant, it was noted that high purge pressure and low purge flows were encountered. the purge cassette and bag were changed, but the issue persisted. to counteract the issue, d10 with heparin was ordered from the pharmacy. care was eventually withdrawn the following day and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not returned by the customer and therefore, evaluation of the device was not possible. should any new information be acquired, a supplemental mdr will be filed. as noted in the impella ifu: impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
19283 | 1220648-2024-18557 | 20305012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305012&pc=PYX | 01/16/2022 | 05/27/2025 |  | 09/25/2024 |  |  | 01/16/2022 | 09/25/2024 |  |  | 07/31/2022 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1493711 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for pulmonary circulation support. while on support, the patient had access site oozing and mattress sutures were placed. the patient¿s left leg had also been very ischemic despite having no femoral access. survival outcome at explant noted the patient expired. medical safety review of the event noted that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp flex with smartassist instructions for use & clinical reference manual instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19284 | 1220648-2024-18553 | 20305020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305020&pc=PYX | 05/14/2021 | 09/25/2024 |  | 09/25/2024 |  |  | 05/14/2021 | 09/25/2024 |  |  | 05/31/2021 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1427640 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support, there were difficulties in stopping the bleeding at the insertion site. in total, six z-sutures were placed to stop the bleeding. survival outcome at explant noted the patient expired. medical safety review of the event noted that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp flex with smartassist instructions for use & clinical reference manual instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19285 | 1220648-2024-18554 | 20305033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305033&pc=PYX | 07/29/2021 | 09/25/2024 |  | 09/25/2024 |  |  | 07/29/2021 | 09/25/2024 |  |  | 01/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022049580 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella rp device for hemodynamic support. while on support, the patient was heparin-induced thrombocytopenia+ (hit+), and all heparin was removed. survival outcome at explant noted the patient expired. medical safety review of the event noted that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
19286 | 1220648-2024-18555 | 20305105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305105&pc=PYX | 09/09/2021 | 05/12/2025 |  | 09/25/2024 |  |  | 09/09/2021 | 09/25/2024 |  |  | 10/31/2022 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2021010708 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock infarction was implanted with an impella rp device for mechanical circulatory support. it was reported while on impella rp support, the mattress sutures at the impella insertion site failed and the site began bleeding. three sutures were placed to obtain hemostasis. it was noted at procedural outcome that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19287 | 1220648-2024-18551 | 20305111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305111&pc=PYX | 04/29/2021 | 05/21/2025 |  | 09/25/2024 |  |  | 04/29/2021 | 09/25/2024 |  |  | 07/31/2022 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1493711 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the heparin concentration was reduced due to bleeding from chest tubes. the family withdrew care, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19288 | 1220648-2024-18556 | 20305115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305115&pc=PYX | 12/20/2021 | 05/12/2025 |  | 09/25/2024 |  |  | 12/20/2021 | 09/25/2024 |  |  | 05/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2022099454 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. during support, the patient experienced access site bleeding. to resolve this issue, blood products were administered, and heparin was withheld. the procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19289 | 1220648-2024-18559 | 20305226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305226&pc=PYX | 05/13/2022 | 06/02/2025 |  | 09/25/2024 |  |  | 05/13/2022 | 09/25/2024 |  |  | 09/30/2022 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 1502413 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, the patient experienced limb ischemia. the resolution for this issue is unknown. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp smart assist system with automated impella controller: section: contraindications (united states): ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ this report is being filed as part of a retrospective review of historical records. | Medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. the following has been corrected from the initial medical device report 1220648-2024-18559: section h6: health effect - impact code corrected to 4648.
19290 | 1220648-2024-18560 | 20305279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305279&pc=PYX | 05/26/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 05/26/2022 | 09/25/2024 |  |  | 02/29/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2023184815 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. the impella rp was placed without issue into position. heavy blood was oozing and required a second mattress suture. a forward tension suture was placed with a repositioning sheath. the site was still oozing while trying to dress the site. a third mattress suture placed with a rotated stopcock used to maintain tension. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19291 | 1220648-2024-18561 | 20305281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305281&pc=PYX | 11/07/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 11/07/2022 | 09/25/2024 |  |  | 03/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022081667 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, the tissue plasminogen activator (tpa) was started due to low impella rp purge flows. survival outcome at explant noted the patient expired. medical safety review of the event noted that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
19292 | 1220648-2024-18562 | 20305301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305301&pc=PYX | 11/10/2022 | 05/13/2025 |  | 09/25/2024 |  |  | 11/10/2022 | 09/25/2024 |  |  | 06/30/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023221384 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in was implanted with an impella rp device for mechanical circulatory support. while on support, there was intermittent low purge pressures. there were no leaks in the system noted, and the purge cassette was changed. dextrose increased to d10 and the issue resolved. low purge pressures at 0100 were again reported the following night, no leaks in purge system. purge cassette changed without resolution. purge dextrose increased to d20 (from d10) with pp and pf returning within range. survival outcome at explant noted the patient expired. medical safety review of the event noted that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp & rp flex impella; rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller instructions for use for the related event are as follows: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (eg, infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19293 | 1220648-2024-18563 | 20305311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305311&pc=PYX | 03/18/2022 | 05/21/2025 |  | 09/25/2024 |  |  | 03/18/2022 | 09/25/2024 |  |  | 02/28/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022045556 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. while on support, the pump stopped. during placement there were many repositioning difficulties. the impella would not advance through the tv valve area. eventually after many attempts, the impella went in deflecting towards the right pulmonary artery. there were a few suctions events and a few high mc events leading to pump stopping. the pump and swan were repositioned a few more times and the issue was resolved. survival outcome at explant noted the patient expired. medical safety review of the event noted that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella; rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19294 | 1220648-2024-18565 | 20305397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305397&pc=PYX | 09/26/2021 | 05/22/2025 |  | 09/25/2024 |  |  | 09/26/2021 | 09/25/2024 |  |  | 04/30/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022094780 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella rp device for mechanical circulatory support. while on support, suction alarms were received despite the device being in a good position. the device was explanted, and care was withdrawn. survival outcome at explant noted the patient expired. medical safety review of the event noted that the use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19295 | 1220648-2024-18558 | 20305406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305406&pc=PYX | 05/03/2022 | 05/12/2025 |  | 09/25/2024 |  |  | 05/03/2022 | 09/25/2024 |  |  | 01/31/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2023173479 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. it was reported that the mattress suture broke and the patient experienced a significant bleed. the physician was able to purse string the sheath and the patient was able to maintain hemostasis. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller: section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound." ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19296 | 1220648-2024-18566 | 20305408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305408&pc=PYX | 07/18/2021 | 05/21/2025 |  | 09/25/2024 |  |  | 07/18/2021 | 09/25/2024 |  |  | 10/31/2021 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1457586 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, the patient had bleeding episode, had to transfuse two (2) units of pack red blood cells (prbc's). the impella was removed the next day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp flex with smartassist instructions for use & clinical reference manual instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp flex smartassist system catheter: use of the repositioning sheath and the 23 fr peel-away introducer: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ impella rp with smart assist system with automated impella controller: section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound." ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19297 | 1220648-2024-18568 | 20305471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305471&pc=PYX | 02/01/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 02/01/2022 | 09/25/2024 |  |  | 07/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022131543 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, it was noted that the patient had a small hematoma. in addition, the team removed heparin in purge for high aptt (>200) and bleeding from a-line site and mouth. the plan was to restart heparin at a lower dose and recheck aptt. the patient was successfully weaned and explanted in icu. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp flex with smartassist instructions for use & clinical reference manual instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19298 | 1220648-2024-18564 | 20305519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305519&pc=PYX | 07/02/2021 | 05/21/2025 |  | 09/25/2024 |  |  | 07/02/2021 | 09/25/2024 |  |  | 05/31/2022 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1484873 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, the patient had some bleeding around the endotracheal [et] tube and that the systemic heparin was stopped. the patient was successfully weaned from the impella. | [5 empty] | D6 explant date unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp flex with smartassist instructions for use & clinical reference manual instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ impella rp with smart assist system with automated impella controller. section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19299 | 1220648-2024-18569 | 20305580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305580&pc=PYX | 12/18/2021 | 05/21/2025 |  | 09/25/2024 |  |  | 12/18/2021 | 09/25/2024 |  |  | 02/28/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2022056174 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. it was reported while on impella rp support, the patient experienced bleeding from their chest tube. as a result, heparin was reduced from the purge solution. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined.
19300 | 1220648-2024-18571 | 20305591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305591&pc=PYX | 11/08/2021 | 05/21/2025 |  | 09/25/2024 |  |  | 11/08/2021 | 09/25/2024 |  |  | 06/30/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022100752 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. the impella rp was placed without issue. while on support, the patient was returned to the operating room for a cardiac tamponade. a significant amount of blood was noted in the chest. heparin was withheld. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19301 | 1220648-2024-18570 | 20305592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305592&pc=PYX | 11/05/2021 | 05/21/2025 |  | 09/25/2024 |  |  | 11/05/2021 | 09/25/2024 |  |  | 06/30/2022 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1489693 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support it was reported that the patient had significant bleeding from chest tubes, so no heparin was in place. the patient was taken to the operating room for a washout of blood and clots. the patient was unstable; however, the impella was explanted and the patient continued on ecmo support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19302 | 1220648-2024-18567 | 20305639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305639&pc=PYX | 07/30/2021 | 05/21/2025 |  | 09/25/2024 |  |  | 07/30/2021 | 09/25/2024 |  |  | 02/28/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022043145 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella rp device for mechanical circulatory support. while on support, the patient had a possible gi bleed. the patient was not on heparin due to bleeding. the patient was also confirmed for heparin-induced thrombocytopenia (hit). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp flex with smartassist instructions for use & clinical reference manual instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19303 | 1220648-2024-18572 | 20305675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305675&pc=PYX | 11/28/2021 | 05/20/2025 |  | 09/25/2024 |  |  | 11/28/2021 | 09/25/2024 |  |  | 08/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022132542 |  | 004334 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic was implanted with an impella rp device for mechanical circulatory support. while on support the patient was found to be positive for a brain bleed. heparin was removed and bicarb was added to the purge. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19304 | 1220648-2024-18573 | 20305787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20305787&pc=PYX | 12/08/2021 | 05/21/2025 |  | 09/25/2024 |  |  | 12/08/2021 | 09/25/2024 |  |  | 10/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022143614 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella rp for mechanical circulatory support. the rp was placed post procedure of a mitral valve and aortic valve replacement through the right femoral vein utilizing the cannulation site. the sheath was removed, and a purse string suture was used to maintain hemostasis around the rp. the patient remained in the or with chest open to manage bleeding at surgical site. the patient was completely sedated overnight due to massive bleeding issues. the rv on an echocardiogram looked good. the impella was weaned and explanted at bedside using purse string suture that was placed at time of implant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp system with the automated impella controller instructions for use for the related event are as follows: section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19305 | 1220648-2024-18574 | 20306356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20306356&pc=OZD | 07/29/2024 | 11/14/2024 |  | 09/25/2024 |  |  | 08/29/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025511335 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant in japan reported that the impella cp was exchanged due to the patient developing hemolysis. patient expired improved with the pump replacement. however, patient ultimately expired from acute heart failure. impella functioned until the end without issue. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The investigation for the hemolysis has been completed. product and data logs were not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. this pump was used beyond indication. the impella cp intended design life is 8 days. the instructions for use for the impella cp with smartassist system states the following: section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ corrections to the initial mfr report: b5: it was stated in the initial report that " patient expired improved with the pump replacement." however, it should have said "patient hemolysis improved with the pump replacement."
19306 | 1220648-2024-18575 | 20306376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20306376&pc=OZD | 08/29/2024 | 11/11/2024 |  | 09/25/2024 |  |  | 08/29/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526454 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported that during impella cp support, the 76-year-old male patient had stenosis in the right femoral artery and had developed leg ischemia due to the impella. limb ischemia was treated independently and successfully with a fem-fem bypass. however, patient expired on support. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: b2, d4-udi number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
19307 | 1220648-2024-18576 | 20307003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20307003&pc=OZD | 08/27/2024 | 11/14/2024 |  | 09/25/2024 |  |  | 08/27/2024 | 09/25/2024 |  |  |  |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488390 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, extensive bleeding was noted on the access side after transfer to intensive care unit (icu). the icu staff could not handle the bleeding, and a surgical intervention was needed. afterwards, a sandbag was placed on the groin to prevent further bleeding. patient eventually expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The patient expired while on support. per medical safety officer review there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical details were provided. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.2 revised as a selection was missing from initial manufacturer device report number 1220648-2024-18576. b.5 revised as medical safety review was missing from initial manufacturer device report number 1220648-2024-18576. d.4 revised unique identifier (udi) # as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-18576. e.1 revised name prefix/title, facility name, address line 1 and city since initial manufacturer device report number 1220648-2024-18576 was submitted. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-18576 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-18576 in accordance with updated procedures.
19308 | 1220648-2024-18577 | 20307023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20307023&pc=OZD | 10/29/2021 | 09/25/2024 |  | 09/25/2024 |  |  | 04/05/2022 | 09/25/2024 |  |  | 05/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089524 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of hemolysis. the device was not received for evaluation. the root cause of the hemolysis was unable to be determined since insufficient information was provided. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci . section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
19309 | 1220648-2024-18579 | 20307038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20307038&pc=OZD | 11/10/2021 | 09/25/2024 |  | 09/25/2024 |  |  | 11/15/2021 | 09/25/2024 |  |  | 08/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022126529 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. the complainant reported a failure to deliver the device due to patient's anatomy. as a result of this complication the impella was removed, an intra-aortic balloon pump was placed. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the product was returned. the root cause of the delivery issues is most likely patient condition. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
19310 | 1220648-2024-18578 | 20307040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20307040&pc=OZD | 11/06/2021 | 09/25/2024 |  | 09/25/2024 |  |  | 11/06/2021 | 09/25/2024 |  |  | 08/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121495 |  | 000000000001000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. multiple attempts were made before successful access was achieved. after successful pump insertion, the repository sheath began to ooze. the patient was taken back to the catheterization lab for an exchange to a new 14 french peel-away sheath, which stopped the bleeding. the impella was explanted and replaced with a new impella cp to address the issue. the patient survived the procedure. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding was most likely due to lack of vessel recoil as the bleeding occurred immediately after the transition from the 14fr introducer to the gwru. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
19311 | 1220648-2024-18580 | 20307043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20307043&pc=OZD | 11/10/2021 | 05/14/2025 |  | 09/25/2024 |  |  | 03/29/2022 | 09/25/2024 |  |  | 07/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110893 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced hemolysis. as a result of the hemolysis the impella was weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | The relationship between the hemolysis and the use of impella was unable to be determined as no product, logs, or clinical details were provided.
19312 | 1220648-2024-18582 | 20307046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20307046&pc=OZD | 11/14/2021 | 09/25/2024 |  | 09/25/2024 |  |  | 11/17/2021 | 09/25/2024 |  |  | 07/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114727 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella cp device for mechanical circulatory support during percutaneous coronary intervention (pci). after the pci was completed, bleeding and a hematoma was reported at the impella access site. as a result of this complication the impella was removed, an intra-aortic balloon pump was placed. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of access site bleeding. the device was not received for evaluation. the root cause of the access site bleeding was unable to be determined as no product or logs were returned for analysis. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
19313 | 1220648-2024-18581 | 20307047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20307047&pc=OZD | 11/12/2021 | 10/01/2025 |  | 09/25/2024 |  |  | 11/24/2021 | 09/25/2024 |  |  | 08/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022122134 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. while on support, hemolysis was diagnosed through colored urine. repositioning was done to address the issue, unknown if the issue was resolved. the patient continued support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined.
19314 | 1220648-2024-18583 | 20307049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20307049&pc=OZD | 11/14/2021 | 09/25/2024 |  | 09/25/2024 |  |  | 04/05/2022 | 09/25/2024 |  |  | 07/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105948 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient experienced hemolysis. action taken to resolve the issue are unknown. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of hemolysis. the device was not received for evaluation. the root cause of the hemolysis and its relation to the impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
19315 | 1220648-2024-18584 | 20307050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20307050&pc=OZD | 11/15/2021 | 09/25/2024 |  | 09/25/2024 |  |  | 11/16/2021 | 09/25/2024 |  |  | 07/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022115207 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during repo sheath insertion, the catheter kinked and then had continuous suction alarms. consequently, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was not received for evaluation. the root cause of the low pump flow was due to a kinked cannula, as the clinical states the pump moved as the gwru was inserted. this report is being filed as part of a retrospective review of historical records
19316 | 1220648-2024-18585 | 20307051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20307051&pc=OZD | 08/02/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 08/02/2022 | 09/25/2024 |  |  | 04/30/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213605 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. prior to insertion, a low purge pressure alarm was noted. the purge cassette and line were inspected for leaks or disconnections. the purge system was re-primed and reconnected. however, the purge pressure rapidly fell to the low 100s, triggering a low alarm. after multiple attempts to resolve the low purge pressure alarm, the impella was explanted and replaced with a new impella cp. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of purge leak. the product was returned and was the purge leak was confirmed. the root cause of the purge leak was sidearm reservoir damage. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
19317 | 1220648-2024-18586 | 20307052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20307052&pc=OZD | 08/02/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 08/02/2022 | 09/25/2024 |  |  | 04/30/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219071 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on day 5 of support, a sudden pump stop occurred with no noticeable changes in pump metrics. normal troubleshooting attempts to restart the device were unsuccessful. the physician determined the patient's hemodynamics were stable enough to remove the impella with no replacement device required. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the product was returned. the root cause of the pump stop was bearing wear. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
19318 | 1220648-2024-18587 | 20307053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20307053&pc=OZD | 08/02/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 08/02/2022 | 09/25/2024 |  |  | 03/31/2024 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211498 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the physician reported that the cannula kinked just distal to the motor housing after delivery. an attempt to reposition the pump was made but the cannula remained kinked. suction alarms and decoupled waveforms were observed. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of low pump flow. the device was returned for evaluation and the flow pump flow was confirmd. the root cause of the low pump flow was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
19319 | 1220648-2024-18588 | 20307056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20307056&pc=OZD | 08/02/2022 | 09/25/2024 |  | 09/25/2024 |  |  | 08/02/2022 | 09/25/2024 |  |  | 12/31/2023 |  | 09/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172936 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, suction alarms were observed. the physician opted to remove the pump and replace it with an intra-aortic balloon pump and administer vasopressors. the procedural outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records
19320 | 1220648-2024-18383 | 20289526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20289526&pc=PYX | 05/06/2021 | 09/24/2024 |  | 09/24/2024 |  |  | 05/06/2021 | 09/24/2024 |  |  | 05/31/2022 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 1484873 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. during support, the patient developed hemolysis. it is unknown how this issue was resolved. additional information noted the family decided to withdraw care and the patient expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. patient's peri-procedural condition was critical. | [5 empty] | A.5 ethnicity is unknown. the impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records.
19321 | 1220648-2024-18384 | 20289557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20289557&pc=PYX | 05/06/2021 | 09/24/2024 |  | 09/24/2024 |  |  | 05/06/2021 | 09/24/2024 |  |  | 11/30/2022 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2021016762 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock infarction was implanted with an impella rp device for mechanical circulatory support. it was reported while on impella rp support, the patient experienced oozing at the impella access site. an additional mattress suture was placed and pressure was held until hemostasis was achieved. the procedural outcome noted that the patient expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp with smart assist system with automated impella controller section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19322 | 1220648-2024-18385 | 20289592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20289592&pc=PYX | 05/11/2021 | 05/20/2025 |  | 09/24/2024 |  |  | 05/11/2021 | 09/24/2024 |  |  | 01/31/2023 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2022049585 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. during support, the patient developed heparin-induced thrombocytopenia (hit). the staff added bicarbonate to the purge solution to resolve this issue. additional information noted care was withdrawn and the patient expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19323 | 1220648-2024-18386 | 20289614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20289614&pc=PYX | 05/20/2021 | 09/24/2025 |  | 09/24/2024 |  |  | 05/20/2021 | 09/24/2024 |  |  | 10/31/2021 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 1462891 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that the impella rp had a malfunctioning tuohy that would not lock. the pump remained on for support. the care team had to slide the sterile sleeve and reposition sheath all the way back on impella rp catheter prior to peel away removal (too much resistance). additionally, the patient experienced bleeding post peel away sheath removal and reposition sheath replacement. manual pressure was applied for 30 minutes and another mattress suture was added. heparin in the purge was decreased and no systemic heparin on board after catheterization lab. the patient expired while on support. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. | The investigation of the product damage (sterile sleeve damage) and the access site bleeding has been completed. the root cause of the product damage was not determined since product was not returned. clinical details were provided which were sufficient to determine the root cause. the root cause of the access site bleeding issue was high (300) patient activated clotting time (act) values. corrections to the initial mfr report: g1 mdr reporting contact email updated. h6 conclusion code 19 added.
19324 | 1220648-2024-18387 | 20289647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20289647&pc=PYX | 06/11/2021 | 09/24/2024 |  | 09/24/2024 |  |  | 06/11/2021 | 09/24/2024 |  |  | 09/30/2022 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 1502413 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, the device exhibited high pressure flow which to resolve this issue tissue plasminogen activator was added to the purge. it was reported that the patient experienced gastrointestinal bleeding and hemolysis. the resolution for these issues is unknown. the patient expired while on support. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. | [5 empty] | A.5 ethnicity is unknown. the impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records.
19325 | 1220648-2024-18388 | 20289680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20289680&pc=PYX | 06/08/2021 | 05/14/2025 |  | 09/24/2024 |  |  | 06/08/2021 | 09/24/2024 |  |  | 12/31/2022 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2022027419 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, the patient was actively bleeding from multiple areas and was unable to start continuous renal replacement therapy (hemodialysis). heparin was removed from the purge and the staff monitored the pump metrics. the patient expired while on support. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. | [5 empty] | A.5 ethnicity is unknown. the impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19326 | 1220648-2024-18389 | 20289713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20289713&pc=PYX | 06/18/2021 | 05/20/2025 |  | 09/24/2024 |  |  | 06/18/2021 | 09/24/2024 |  |  | 10/31/2021 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 1458488 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. during support, the patient experienced bleeding resulting in a subdural hematoma. systemic heparin was withheld. the patient expired while on support. medical safety review of the event noted that there is not enough evidence to exclude the impella as an associated factor in the patient's (death) outcome. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19327 | 1220648-2024-18390 | 20289771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20289771&pc=PYX | 06/23/2021 | 09/24/2024 |  | 09/24/2024 |  |  | 06/23/2021 | 09/24/2024 |  |  | 02/28/2023 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2022065736 |  | 004334 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, the patient developed a hematoma. decreasing heparin in the purge and reducing the dosage were discussed. additional information noted the patient expired within three hours of admiision to the unit. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. | [5 empty] | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19328 | 1220648-2024-18391 | 20291059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20291059&pc=PYX | 05/15/2022 | 09/24/2024 |  | 09/24/2024 |  |  | 05/15/2022 | 09/24/2024 |  |  | 02/29/2024 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023184844 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with post-op left ventricular assist device (lvad) was implanted with an impella rp device for mechanical circulatory support. while on support, patient experienced ventricular tachycardia (vt). patient was given amiodarone drip to resolve vt. additional information noted that the care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp system with automated impella controller. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records.
19329 | 1220648-2024-18392 | 20291067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20291067&pc=PYX | 09/20/2022 | 09/24/2024 |  | 09/24/2024 |  |  | 09/20/2022 | 09/24/2024 |  |  | 12/31/2023 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023160873 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella rp device for pulmonary circulatory support. it was reported that a diagnostic cath was performed to assess coronaries. the diagnostic revealed an embolus from native valve. a stent was placed followed by the cardiac surgeon deciding to remove impella rp due to higher flows on tandem. patient survived the explant procedure but subsequently expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
19330 | 1220648-2024-18393 | 20291104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20291104&pc=PYX | 05/25/2021 | 09/24/2024 |  | 09/24/2024 |  |  | 05/25/2021 | 09/24/2024 |  |  | 10/31/2022 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED,INC. |  |  | US |  | 2021007584 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support, partial thromboplastin time (ptt) was noted to be less than 100 with multiple sites bleeding. the purge was d5w without heparin. the physician was advised to add heparin once oozing slowed and ptt lowered. it was also noted that rp flows suddenly dropped. decision was made to exchange catheter for new rp. new rp catheter inserted over .027 wire, increased p-levels to p-8 with current flows of 3.1 l/min. procedure outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ impella rp with smartassist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19331 | 1220648-2024-18394 | 20291120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20291120&pc=PYX | 05/28/2021 | 09/24/2024 |  | 09/24/2024 |  |  | 05/28/2021 | 09/24/2024 |  |  | 11/30/2022 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2021016868 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support the patient experienced access site bleeding. it was noted that the physician performed a figure 8 suture, rather than mattress suture. ct surgeon came by and performed proper mattress suture technique and achieved hemostasis. additional information noted care the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿. impella rp with smartassist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿. ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿. this report is being filed as part of a retrospective review of historical records.
19332 | 1220648-2024-18395 | 20291191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20291191&pc=PYX | 07/10/2021 | 05/23/2025 |  | 09/24/2024 |  |  | 07/10/2021 | 09/24/2024 |  |  | 02/28/2023 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022064748 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Asystole | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. while on support flows dropped to 1.5. x-ray showed device malposition. the patient was taken to cath lab to attempt to reposition under fluoro. after multiple attempts to reposition the impella, the patient was coding, requiring multiple shocks. the patient was found to be in asystole and expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp system with automated impella controller & impella rp flex with smartassist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19333 | 1220648-2024-18397 | 20291205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20291205&pc=PYX | 01/05/2022 | 05/21/2025 |  | 09/24/2024 |  |  | 01/05/2022 | 09/24/2024 |  |  | 06/30/2023 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022117701 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Perforation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella rp device for mechanical circulatory support. while on support, patient had bleeding from the chest tube. heparin was stopped and blood products were administered to the patient. additionally, the patient had a right ventricle perforation and was sent to the operating room to repair the ventricle and ten (10) units of red blood cells were administered. additional information noted the care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu (bleeding): impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ impella rp smartassist system with automated impella controller (ventricle perforation) section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19334 | 1220648-2024-18396 | 20291220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20291220&pc=PYX | 01/05/2022 | 09/24/2024 |  | 09/24/2024 |  |  | 01/05/2022 | 09/24/2024 |  |  | 01/31/2023 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022031814 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, it was noted that access site had some oozing. the oozing was controlled with mattress sutures. additional information noted that the patient¿s family opted to have care withdrawn and the patient expired thereafter. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19335 | 1220648-2024-18398 | 20291255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20291255&pc=PYX | 08/12/2021 | 05/14/2025 |  | 09/24/2024 |  |  | 08/12/2021 | 09/24/2024 |  |  | 02/28/2023 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022056124 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for hemodynamic support. while on support, continuous suction was noted. it was also noted that impella slipped back across pulmonary artery valve which ballooned and the patient coded. the patient¿s condition continued to deteriorate, and the decision was made to withdraw care. survival outcome at explant noted the patient expired. medical safety review of the event noted that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp system with automated impella controller & impella rp flex with smartassist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19336 | 1220648-2024-18399 | 20291476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20291476&pc=PYX | 01/21/2022 | 05/21/2025 |  | 09/24/2024 |  |  | 01/21/2022 | 09/24/2024 |  |  | 05/31/2022 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1485147 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. it was reported while on impella rp support, the patient experienced bleeding from the chest. as a result of the bleeding, heparin was reduced. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp flex with smartassist instructions for use & clinical reference manual instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
19337 | 1220648-2024-18400 | 20291929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20291929&pc=OZD | 07/20/2024 | 06/30/2025 |  | 09/24/2024 |  |  | 07/20/2024 | 09/24/2024 |  |  |  |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025485829 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, the patient had increased runs of ventricular tachycardia which required more support. additionally, purge blocked alarm occurred. tissue plasminogen activator (tpa) protocol was initiated and purge flow returned to within normal limits after 20 hours of tpa protocol. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider. | The investigation of high purge pressure and vf/vt/af/at has been completed. the impella device was not returned for evaluation. high purge pressure: data logs were returned and there were no suction alarms or motor current spikes prior to the high purge pressure alarms. the purge pressure was seen rising gradually over the span of 10 minutes. the pressure reached 1080mmhg and triggered the high purge pressure alarms and purge system blocked alarms. as per the clinical information, tissue plasminogen activator (tpa) protocol was initiated which helped return the purge flow to normal. these symptoms are similar to that of biomaterial buildup. therefore, the root cause of the high purge pressure issue was most likely biomaterial buildup. vf/vt/af/at: as the necessary information was not provided, the root cause of the vt/vf was not determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-18400 in accordance with updated procedures.
19338 | 1220648-2024-18402 | 20292067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20292067&pc=OZD | 08/15/2024 | 09/18/2025 |  | 09/24/2024 |  |  | 08/15/2024 | 09/24/2024 |  |  |  |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025478386 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Infusion or Flow Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the impella 5.5 had impella in aorta alarm with spike of motor current and left ventricle (lv) waveform. impella flows and waveforms were consistent with flow saturation. the impella 5.5 was confirmed in proper position. later in the morning, the patient¿s urine turned dark red. another echocardiogram was requested to assess lv thrombus. the impella 5.5 was exchanged for right axillary impella replacement. the procedural outcome was the patient survived the surgery. | It was reported that the device exhibited a systolic placement signal. no further information provided. |  |  |  |  | The investigation of hemolysis and thrombus has been completed. the cause of the hemolysis issue was biomaterial, based on return product investigation and data log review. as the necessary information was not provided, the root cause of the thrombus was not determined. | The investigation into the saturated flow has been completed. however, the investigation into hemolysis and thrombus is pending. the impella pump was returned for investigation. no physical abnormalities were reported on the returned pump. biomaterial was flushed out while the pump was run in a bucket of water for flow testing. the pump flow was within limits. according to the clinical report, placement signal not reliable alarms was triggered, and there were no issues with motor current or pump flow at the start of the case. after 14 days, a good pump position was confirmed during a reported saturated pump flow. the pump was then explanted. the data log showed similar trending placement signal and optical signal parameters (snr and contrast) while running at a higher p-level (p8), with reported low pump flow. this trend returned to normal after reducing the p-level. a motor current spike was reported at the end of the case run, followed by saturated pump flow. the root cause of the saturated pump flow issue was biomaterial, based on return product investigation and data log review. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation for the optical signal issue has been completed. the data log shows similar trending placement signal and optical signal parameters (snr and contrast) while running at a higher p-level (p8), with reported low pump flow. this trend returned to normal after reducing the p-level. a motor current spike was reported at the end of the case run, followed by saturated pump flow. the root cause of the optical signal issue was patient condition related, based on return product investigation and data log review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. g1 updated with correct mdr reporting contact email h.6. medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-18402.
19339 | 1220648-2024-18401 | 20292154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20292154&pc=PYX | 08/31/2024 | 10/30/2024 |  | 09/24/2024 |  |  | 08/31/2024 | 09/24/2024 |  |  |  |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025518073 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella rp flex placed to support a patient with multiple cardiac and systemic risk factors/comorbidities. the patient had suffered from an access site bleed and hypovolemia with low flows. the patient was treated with 2 units of red blood cells and platelets. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | The investigation for the low pump flow and the access site bleed has been completed. no product was returned for investigation. the data logs confirm low pump flows for matching p-level. the data logs show a downward trend of noisy placement signal and negative cvp during low flow events. the flows increase and recover to expected levels at p6. clinical details state that the patient was hypovolemia and received a blood transfusion. therefore, the root cause of the low pump flow was most likely patient condition related. the root cause of the access site bleeding - major was not determined as limited clinical information was provided.
19340 | 1220648-2024-18403 | 20292270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20292270&pc=OZD | 08/30/2024 | 10/31/2024 |  | 09/24/2024 |  |  | 08/30/2024 | 09/24/2024 |  |  |  |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Other | Company representation |  |  |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Mitral Perforation | [11 empty] | Life Threatening | Required Intervention | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received information from abiomed legal on (b)(6) 2024 regarding a impella cp patient in michigan with reported hemolysis transferred to cleveland clinic for higher level of care. in the operating room, the discovery was made of perforation and effusion as well as bruising and mitral valve regurgitation. patient had surgical repair to the mitral valve and coronary sinus. | [5 empty] | The investigation for the hemolysis and the heart valve damage has been completed. the pump was not returned for investigation, and the pump's serial number was not provided. data logs were not retrieved due to the unknown serial number. according to the provided information, hemolysis was reported during impella use, and it was noted that mitral valve regurgitation required a surgical approach. the cause of the heart valve damage issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. added- a2, a3, h6 impact code 4627. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
19341 | 1220648-2024-18404 | 20292347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20292347&pc=OZD | 09/02/2024 | 12/11/2025 |  | 09/24/2024 |  |  | 09/02/2024 | 09/24/2024 |  |  | 02/28/2026 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474069 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant had placed a impella 5.5 pump for care escalation from the impella cp to support a patient admitted in acute myocardial infarction/cardiogenic shock. the impella 5.5 was placed and clamped with the soft jaw. the pump then had saturated flows and new valve injury of the aortic insufficiency. these issues prompted the team to remove the impella 5.5 pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (unites states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter. ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter. ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation for the saturated flow and the heart valve damage has been completed. real time data logs were reviewed which showed slight flow saturation at p-5. flow seems to be close to 4l/min at p-5 and greater than 4.4 l/min at one instance. normal flows at p-5 is 2.1 to 3.6l/min. also, placement signal was trending down. no ingestion trend seen or motor current spike. saturated flows is seen for approximately 18 mins. the pump was returned and there were no visual abnormalities. the pump was run and saturated flow was reproduced with an upward trend of purge pressure. a tear down of the pump was performed and there was material around the bearing. the root cause of saturated flow was most likely biomaterial. the root cause of the heart valve damage is not determined as limited information related to failure was provided. added d9 device return date. | Additional information was received in the form of medical safety review and noted the following: medical safety reviewed the event and noted there is not enough information to exclude impella 5.5 pump 3 as an associated factor in the patient's outcome. impella cp pump 1 and impella 5.5 pump 2 events are serious injury. this is one of three device reports associated with the event. refer to the following manufacturer report numbers: 1220648-2025-48263 (impella cp pump 1), 1220648-2024-18404 (impella 5.5 pump 2), and 1220648-2025-48473 (impella 5.5 pump 3).
19342 | 1220648-2024-18405 | 20292434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20292434&pc=OZD | 09/02/2024 | 09/24/2024 |  | 09/24/2024 |  |  | 09/02/2024 | 09/24/2024 |  |  |  |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025526002 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed a impella cp to support a patient admitted for a st elevation myocardial infarction. the pump was placed however the patient does not have pedal pulses and hemolyzed labs. the decision was made to explant the impella cp and place the impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
19343 | 1220648-2024-18359 | 20292446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20292446&pc=PYX | 09/17/2021 | 09/24/2024 |  | 09/24/2024 |  |  | 09/17/2021 | 09/24/2024 |  |  | 12/31/2022 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED. INC. |  |  | US |  | 2022024701 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, bleeding was noted at the groin access site and a mattress suture was placed to manage the complication. upon removal of the impella rp, a thrombus was observed. actions taken to resolve the issue are unknown. the procedural outcome noted that the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp with smart assist system with automated impella controller. section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
19344 | 1220648-2024-18406 | 20292630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20292630&pc=OZD | 08/27/2024 | 11/19/2024 |  | 09/24/2024 |  |  | 08/27/2024 | 09/24/2024 |  |  |  |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025506327 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant placed an impella cp for support of a patient admitted. the medical doctor stated that the patient was bleeding heavily from the groin and became tachycardiac which prompted the team to explant the impella cp. the patient eventually expired after care was withdrawn. medical safety review of the event noted that there is not enough evidence to exclude the impella as an associated factor in the patient's (death) outcome. | [5 empty] | A.5 is unknown. b.2 the exact date of death is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | D.9 updated as the purge cassette was returned for investigation. the investigation of access site bleeding ¿ major has been completed. with only the purge cassette being returned for investigation, the root cause of the access site bleeding could not be determined. b.2 revised since initial manufacturer device report number 1220648-2024-18406 was submitted to add required intervention as the device was explanted. h.6 code 4115 was entered incorrectly on initial manufacturer device report number 1220648-2024-18406 as a product (purge cassette) has been returned for evaluation.
19345 | 1220648-2024-18407 | 20292742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20292742&pc=OZD | 08/29/2024 | 03/25/2025 |  | 09/24/2024 |  |  | 08/29/2024 | 09/24/2024 |  |  |  |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512633 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 79-year-old male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, purge pressure high alarm and purge system block alarm was present. tissue plasminogen activator (tpa) protocol from the facility¿s previous experience with an alarm was initiated and seven hours later, the alarm was still present. the pump was removed and replaced. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | Upon review, manufacturer device report number 1220648-2025-26906 was found to be a duplication of manufacturer device report number 1220648-2024-18610. please refer to manufacturer device report number 1220648-2024-18610 for information.
19346 | 1220648-2024-18408 | 20292900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20292900&pc=OZD | 08/30/2024 | 03/27/2025 |  | 09/24/2024 |  |  | 09/04/2024 | 09/24/2024 |  |  |  |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103719 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 support being prepped when the automated impella controller (aic) failed. the team replaced the aic and support proceeded without harm or injury from delay. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation has been completed. the console had a controller error due to a bad lamp on (b)(6) 2024 and then again on the complaint date of (b)(6) 2024. the console was investigated. the console was confirmed to have an intermittently failing lamp. the cause of the controller error was a bad lamp on the optical bench. e.1 added fax number as it was inadvertently omitted on the initial manufacturer device report number 1220648-2024-18408. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-18408 was submitted in accordance with updated procedures.
19347 | 1220648-2024-18410 | 20292957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20292957&pc=OZD | 09/04/2024 | 11/06/2024 |  | 09/24/2024 |  |  | 09/04/2024 | 09/24/2024 |  |  |  |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025516514 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while the pump was inserted, flows started at 1l/min and immediate suction alarms were present. a kink was noted at the at the proximal end of the cannula. multiple maneuvers were attempted to reposition and correct the kink but were unsuccessful. decision was made to remove the impella. no patient harm has been reported. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the low pump flow and the cannula kink has been completed. the impella cp pump was returned for evaluation and was visually inspected. cannula kink near outflow cage was observed. no biomaterial or broken impeller blade was observed. no other abnormality found. the cause of the low pump flow was the cannula kink as a result of improper technique. added- h6 impact code 4628.
19348 | 1220648-2024-18411 | 20293006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20293006&pc=OZD | 08/30/2024 | 10/31/2024 |  | 09/24/2024 |  |  | 08/30/2024 | 09/24/2024 |  |  |  |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531302 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support due to partial st elevation myocardial infarction. while on support an echocardiogram was performed and the inlet was found to be 3.5cm to aortic valve but remarked that inlet may be close to papillary structures. the patient's urine was dark and they were eventually put on dialysis. it is unknown if the pump was repositioned and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of hemolysis, positioning issues, and renal failure has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
19349 | 1220648-2024-18412 | 20293038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20293038&pc=OZD | 08/29/2024 | 11/12/2024 |  | 09/24/2024 |  |  | 08/29/2024 | 09/24/2024 |  |  |  |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025538498 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 46-year-old male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella cp support, the patient had limb ischemia. a femoral to femoral bypass was performed. support was continued and the patient survived the explant. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
19350 | 1220648-2024-18413 | 20293098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20293098&pc=OZD | 09/05/2024 | 11/06/2024 |  | 09/24/2024 |  |  | 09/05/2024 | 09/24/2024 |  |  |  |  | 09/24/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147478 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that during preventative maintenance, the automated impella controller (aic) displayed a controller failure issue. after consulting with engineers and troubleshooting, replacing the purge pressure sensor controller resolved the issue. purge pressure sensor was changed on the console. no patient involvement. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) red alarm has been completed. data logs were reviewed which confirmed the reported failure. there were numerous instances of purge pressure sensor defective controller failure (event set # (b)(4)) found on the reported occurrence date. the console was returned for evaluation finding that the reported controller failure was able to be reproduced. the controller failure was determined to be caused by a defective purge pressure sensor.
19351 | 1220648-2024-18414 | 20293199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20293199&pc=OZD | 08/29/2024 | 11/20/2024 |  | 09/24/2024 |  |  | 08/29/2024 | 09/24/2024 |  |  |  |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524140 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant placed an impella cp to support a patient admitted in for high risk percutaneous coronary intervention. the pump was placed but due to bleeding at the access site, one bag of blood was delivered and the issue resolved. the pump remained on and the patient expired the following day. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. patient's peri-procedural condition was critical. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
19352 | 1220648-2024-18415 | 20293217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20293217&pc=OZD | 09/05/2024 | 11/06/2024 |  | 09/24/2024 |  |  | 09/05/2024 | 09/24/2024 |  |  |  |  | 09/24/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1548488 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a temperature high alarm that self-resolved. the alarm occurred when the patient was returning from ambulation and plugging aic back into outlet. no patient harm has been reported. | It was further reported that care was withdrawn and the patient expired. the patient did not lose support at any time despite the alarm. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the console (aic) battery failure alarm has been completed. data logs were reviewed which revealed the ¿battery temperature high¿ alarm (#49, due to battery temperature over 60degc) that self-resolved in 10 seconds. there was no preceding battery temperature alarm #48, which occurs when battery temperature is between 50degc and 60degc. battery and power data embedded in ltlogs do not show any temperature rise or elevated battery temperatures, as ltlog data is recorded every 60s. the console was returned for evaluation. upon functional testing, the batteries and power management system on ic8539 were tested, and no issues were observed. the root cause of the battery temp high alarm was likely due to a momentary communication glitch between the batteries and the pbm, where a single sample from battery data (in this case value of battery temperature) was corrupt. b5-mso review.
19353 | 1220648-2024-18416 | 20293250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20293250&pc=OZD | 08/13/2024 | 11/21/2024 |  | 09/24/2024 |  |  | 09/05/2024 | 09/24/2024 |  |  | 05/31/2025 |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350439 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 37 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Renal Failure | Tachycardia | Ventricular Fibrillation | [9 empty] | Required Intervention | [5 empty] | RENAL REPLACEMENT THERAPY. | VENOUS-ARTERIAL EXTRA CORPOREAL MEMBRANE OXYGENATI. | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Information from literature was received. after impella cp was placed and approximately 3 hours after admission to the intensive care unit, an electrocardiogram showed patient with ventricular tachycardia. impella support was maintained during p9 without any suction alarms. attempts were made to suppress the ventricular tachycardia by repositioning the impella cp, managing sedation, and administering amiodarone, but the patient's condition deteriorated, and the patient developed persistent ventricular fibrillation. therefore, venous-arterial extra corporeal membrane oxygenation (va ecmo) was added for cardiopulmonary support. due to pulmonary edema and anuria, continuous hemofiltration was initiated 12 hours after admission. cardiac and renal function gradually improved, and the patient was successfully weaned from va-ecmo on day 4, from impella cp on day 5, and from renal replacement therapy on day 23 after admission. during hospitalization, the patient's electrocardiogram showed dramatic changes along with the patient health condition. after rehabilitation, the patient was discharged on day 41 without any neurological, cardiovascular, respiratory, or renal sequelae. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported renal failure and ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the renal failure and vt could not be determined.
19354 | 1220648-2024-18418 | 20293412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20293412&pc=OZD | 08/30/2024 | 10/24/2024 |  | 09/24/2024 |  |  | 08/30/2024 | 09/24/2024 |  |  |  |  | 09/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Other | Company representation |  |  |  |  | 54 YR |  | Male |  |  |  | Death | Yes | No | Pumping Stopped | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complaints team was forwarded information by abiomed legal regarding a pump stop with an impella cp with a patient located in wisconsin in 2022. while on support, the impella alarmed, "pump off; motor current high". the pump was restarted twice but could not be restarted a third time when the impella stopped. the patient began to decompensate within 10 minutes. the patient expired 24 minutes after the impella stopped. | [5 empty] | The investigation of pump stop has been completed. the pump was not returned for the investigation. pump serial number information was not provided. data logs could not be retrieved without pump information. pump stop was reported on impella cp during use. the cause of the issue pump stop issue was not determined since product was not returned and necessary clinical information was not provided. e4 should have been blank on manufacturer device report 1220648-2024-18418. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿