26131 | 1220648-2024-11680 | 19398650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19398650&pc=OZD | 10/25/2021 | 05/25/2024 |  | 05/25/2024 |  |  | 10/25/2021 | 05/25/2024 |  |  | 07/31/2022 |  | 05/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1499354 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 10/25/2021 |  |  |  |  | The user facility reported that a patient was implanted with an impella cp for mechanical circulatory support due to an st elevated myocardial infarction (stemi). while on support, it was reported that the patient had a fever and that there were diminished pulses in the lower leg. the patient had a cold leg with diminished pulses, indicative of limb ischemia. the patient was successfully weaned from the impella and the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26132 | 1220648-2024-11681 | 19398653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19398653&pc=OZD | 10/26/2021 | 10/13/2025 |  | 05/25/2024 |  |  | 10/26/2021 | 05/25/2024 |  |  | 02/28/2023 |  | 05/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022054066 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 10/26/2021 |  |  |  |  | The user facility reported a 68-year-old patient implanted with an impella cp for mechanical circulatory support due to acute myocardial infarction/cardiogenic shock. while on support, the patient experienced hemolysis with no specified resolution. additionally, the patient experienced access site bleeding with no indication of intervention needed. the patient was successfully weaned and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined e4 should have been left blank on the initial manufacturer device report number 1220648-2024-11681 as it is unknown if the initial reporter also sent the report to the food and drug administration. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-11681 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-11681 in accordance with updated procedures. h6 component code revised from the initial submission in accordance with updated procedures.
26133 | 1220648-2024-11611 | 19390734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19390734&pc=OZD | 04/26/2024 | 06/11/2024 |  | 05/24/2024 |  |  | 04/26/2024 | 05/24/2024 |  |  | 06/30/2025 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024427795 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient experiencing acute st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was intervened upon in the cardiac catheterization lab at the tertiary center and had reperfusion ectopy. the runs of ectopy, ventricular tachycardia were stabilized. later the patient complained of chest pain and vision problems. computed tomography scan confirmed basilar ischemic stroke. the team and family placed the patient on care, comfort measures for end-of-life transition. medical safety review outcome noted there is not enough information to associate use of the device with the patient's transition. | [5 empty] | The impella device has been received from the customer, however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the stroke/cva issue has been completed since the original report was filed. the pump was returned for investigation. as clinical information related to the stroke/cva was not provided, the root cause could not be determined.
26134 | 1220648-2024-11612 | 19390819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19390819&pc=OZD | 04/29/2024 | 03/19/2025 |  | 05/24/2024 |  |  | 04/29/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025458223 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Pumping Stopped | [6 empty] | Atrial Fibrillation | Atrial Flutter | Tachycardia | No Clinical Signs, Symptoms or Conditions | [9 empty] | Required Intervention | Death | [23 empty] | A |  |  |  | Initial submission | Followup | Followup | Followup |  | 04/29/2024 |  |  |  |  | The complainant reported a 68-year-old female patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported the pump stopped during support. attempts to restart the pump were unsuccessful and the decision was made to explant the pump. upon removal, a large clot was noted at the inlet of the pump. there was no patient harm. | Additional information received from abiomed¿s long-term outcome and quality indicator (loqi) impella registry that roughly two months after explant, the patient experienced paroxysmal atrial fibrillation / atrial flutter and ventricular tachycardia. the relationship to the prior impella support was unknown / undetermined. the patient outcome was fatal. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the pump stop has been completed. the pump was not returned for investigation. data logs do not show any pump stops. the root cause of the pump stop was not determined since product was not returned and data logs are inconclusive. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist for use during cardiogenic shock. | The investigation of atrial fibrillation/atrial flutter and ventricular tachycardia and pump stop was completed. the pump was not returned for investigation. clinical information stated that the pump stopped on (b)(6)2024 and was unable to be restarted, but patient was unable to tolerate anticoagulation wean so the pump was left in until explant on (b)(6)2024. data logs for (b)(6)2024 did not show any pump stops. the pump stop was reported to be on (b)(6)2024. the root cause of the pump stop was not determined since the product was not returned and data logs are inconclusive. as the necessary information was not provided, the root cause of the atrial fibrillation/atrial flutter and ventricular tachycardia was not determined. h.6 code 114 was reported incorrectly on follow up 1 for manufacturer device report number 1220648-2024-11612. a new code was added to investigation findings code. code 4112 was inadvertently omitted from follow up 1 for manufacturer device report number 1220648-2024-11612 and was added. h.10 part of the investigation results for the pump stop were inadvertently omitted from follow up 1 for manufacturer device report number 1220648-2024-11612 and were added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. b1 updated to adverse event and product problem. b2 death added, exact date of death is unknown. h1 type of reportable event updated. h6 health effect - clinical code and health effect - impact code updated to reflect new information the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11612. g1 reporting contact email. g1 reporting contact fax number missing. g2 report source; study was missing.
26135 | 1220648-2024-11614 | 19391067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19391067&pc=PYX | 09/15/2019 | 07/17/2024 |  | 05/24/2024 |  |  | 05/02/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The abiomed complaints team received a publication entitled: "clinical indications of impella short-term mechanical circulatory support in a tertiary centre" in which 57 impella patients over 2 years reviewed. seven of the patients were implanted with an impella rp. out of the seven, there were two patients who endured hemolysis. no further information is available. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. added- h6 component code 925 f6/f8 should have been left blank in the initial report.
26136 | 1220648-2024-11617 | 19391207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19391207&pc=PYX | 03/17/2019 | 06/20/2024 |  | 05/24/2024 |  |  | 05/02/2024 | 05/24/2024 |  |  | 07/31/2020 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1363686 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella rp pump was found to be mispositioned during patient support. it was noted that the outflow was at the level of the pulmonic valve and not in the main artery. retrograde right ventricular outflow tract flow with tricuspid regurgitation present at the time of the issue. the device was subsequently explanted as a result of the positioning issue. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. product was not returned for investigation. data logs were reviewed and did not find any positioning issue or any other abnormality noted on the data logs. therefore, the root cause of the positioning issue was not determined since product was not returned and insufficient information provided. d4-corrected model number.
26137 | 1220648-2024-11618 | 19391251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19391251&pc=OZD | 04/01/2022 | 07/25/2024 |  | 05/24/2024 |  |  | 05/02/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | On 2024-05-02 complaints forwarded abstract from journal heart and lung transplantation entitled: "¿long-term¿ use of impella - safe to do?" in which impella supports from 5/2017 to 5/2021 were analyzed. the complications noted during impella cp support were device migration, pump thrombosis, and general device malfunction. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ zeschky, c., ahmed, h., alayyar, m., jothidasan, a., husain, m., stock, u., & smail, h. (2022). ¿long-term¿ use of impella - safe to do? the journal of heart and lung transplantation, 41(4). https://doi.org/10.1016/j.healun.2022.01.1603 | This supplemental report is being written to retract the initial manufacturer device report 1220648-2024-11618. it was determined that manufacturer device report 1220648-2024-11618 is a duplication. the event was captured on manufacturer device report number 1220648-2024-11449.
26138 | 1220648-2024-11619 | 19391252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19391252&pc=OZD | 10/30/2021 | 03/17/2025 |  | 05/24/2024 |  |  | 10/30/2021 | 05/24/2024 |  |  | 07/31/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121563 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/30/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient went into cardiac arrest and cardiac pulmonary resuscitation was performed. the impella device was not able to obtain a pump speed level above p4 without suction alarms. impella device was weaned and explanted. patient tolerated the procedure and condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26139 | 1220648-2024-11620 | 19391476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19391476&pc=OZD | 04/26/2024 | 06/19/2024 |  | 05/24/2024 |  |  | 04/26/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025455157 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2024 |  |  |  |  | The user facility reported an 81-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced a gastrointestinal bleed and developed hemolysis. the patient was transfused with one unit of packed red blood cells. the decision was made to withdraw care and the patient expired. per medical safety officer (mso) review not enough information to associate use of the device with outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: hemolysis hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ section: potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the hemolysis and the vascular damage - peripheral vessel issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided.
26140 | 1220648-2024-11621 | 19391511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19391511&pc=OZD | 05/02/2024 | 07/03/2024 |  | 05/24/2024 |  |  | 05/02/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  |  |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Renal Failure | Ischemia Stroke | Thrombosis/Thrombus | [6 empty] | Life Threatening | Required Intervention | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | On 2024-05-02 the complaints team forwarded publications. the entitled: "deploying mechanical circulatory support via the axillary artery in cardiogenic shock and high-risk percutaneous coronary intervention" details retrospective review from january 2016 to september 2018 - 48 patients in which 4 had upper limb ischemia, 14 had acute kidney injury requiring renal replacement therapy, 1 had care escalated to ecmo (extracorporeal membrane oxygenation), 1 had ischemic stroke, 6 had bleeding barc >2, hemolysis in 4, and finally 1 had axillary artery thrombosis. the devices were 2.5, cp, rp. unspecified allegations for the one patient care escalated to ecmo. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the reported access site bleed, hemolysis, limb ischemia, renal failure, stroke, thrombus, and ventricular arrhythmia has been completed. the impella devices were not returned for evaluation. clinical details were not sufficient to determine the root cause. therefore the root cause of the reported issues could not be determined. citation: kajy, m., laktineh, a., blank, n., tayal, r., tanveer, s., mohamad, t., elder, m., schreiber, t., & kaki, a. (2020). deploying mechanical circulatory support via the axillary artery in cardiogenic shock and high-risk percutaneous coronary intervention. the american journal of cardiology, 128, 127¿133. https://doi.org/10.1016/j.amjcard.2020.04.039. d4-corrected model number. d4-corrected serial number. h6-added component codes. f6 and f8 should have been blank.
26141 | 1220648-2024-11622 | 19391531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19391531&pc=OZD | 05/02/2024 | 07/03/2024 |  | 05/24/2024 |  |  | 05/02/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Renal Failure | Ischemia Stroke | Thrombosis/Thrombus | [6 empty] | Disability | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | On 2024-05-02 the complaints team forwarded publications. the entitled: "deploying mechanical circulatory support via the axillary artery in cardiogenic shock and high-risk percutaneous coronary intervention" details retrospective review from january 2016 to september 2018 - 48patients in which 4 had upper limb ischemia, 14 had acute kidney injury requiring renal replacement therapy, 1 had care escalated to ecmo (extracorporeal membrane oxygenation) , 1 had ischemic stroke, 6 had bleeding barc >2, hemolysis in 4, and finally 1 had axillary artery thrombosis. the devices were 2.5, cp, rp. unspecified allegations for the one patient care escalated to ecmo. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis .¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. section: contraindications (united states). ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿. | The investigation for the reported access site bleed, hemolysis, limb ischemia, renal failure, stroke, thrombus, and ventricular arrhythmia has been completed. the impella device was not returned for evaluation. clinical details were not sufficient to determine the root cause. therefore the root cause of the reported issues could not be determined. citation: kajy, m., laktineh, a., blank, n., tayal, r., tanveer, s., mohamad, t., elder, m., schreiber, t., & kaki, a. (2020). deploying mechanical circulatory support via the axillary artery in cardiogenic shock and high-risk percutaneous coronary intervention. the american journal of cardiology, 128, 127¿133. https://doi.org/10.1016/j.amjcard.2020.04.039 d4-corrected model number. d4-corrected serial number. f6 and f8 should have been blank. h6 added component codes.
26142 | 1220648-2024-11623 | 19391754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19391754&pc=OZD | 11/03/2021 | 03/25/2025 |  | 05/24/2024 |  |  | 11/03/2021 | 05/24/2024 |  |  | 12/31/2022 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022028971 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2021 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced slight oozing from all line insertion sites. it was also noted that the impella cp device was inserted too deep. heparin was removed from purge. the impella site secured with angle matched. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26143 | 1220648-2024-11624 | 19391844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19391844&pc=OZD | 01/03/2022 | 03/25/2025 |  | 05/24/2024 |  |  | 01/03/2022 | 05/24/2024 |  |  | 07/31/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022123394 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the femoral angio revealed a clot in the profunda and possibly in the superficial femoral artery. it was noted to not be flow limiting and peripheral pulses were palpable. systemic heparin was started. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26144 | 1220648-2024-11625 | 19391927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19391927&pc=OZD | 01/03/2022 | 05/24/2024 |  | 05/24/2024 |  |  | 01/03/2022 | 05/24/2024 |  |  | 07/31/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022115123 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/03/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a retroperitoneal bleed and oozing at the impella sites was noted. heparin was removed from purge and blood products were given to the patient. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records.
26145 | 1220648-2024-11626 | 19392004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19392004&pc=OZD | 11/04/2021 | 03/25/2025 |  | 05/24/2024 |  |  | 11/04/2021 | 05/24/2024 |  |  | 07/31/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110545 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/04/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient's pump was explanted due to cheater found in inlet causing a pump stop. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk cpi instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26146 | 1220648-2024-11627 | 19392056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19392056&pc=OZD | 05/10/2024 | 06/19/2024 |  | 05/24/2024 |  |  | 05/10/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314945 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support, the automated impella controller (aic) displayed ¿impella stopped¿. the motor current was too high. the impella was attempted to be restarted several times without success. the console was exchanged, and the issue resolved. the impella remained implanted. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the aic - pump stop issue has been completed since the initial report was submitted. the console was returned, tested, and visually inspected, but the failure mode could not be reproduced. due to the malfunction seen in the field resolving when switching consoles, but then the failure not reproducing when returned from the field, the most likely cause of the malfunction was a use issue.
26147 | 1220648-2024-11628 | 19392171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19392171&pc=PYX | 12/24/2021 | 06/13/2024 |  | 05/24/2024 |  |  | 04/25/2024 | 05/24/2024 |  |  | 03/31/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022071492 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, the patient exhibited access site bleeding. the bleeding was controlled with a mattress suture and manual pressure, and heparin was withheld. additionally, limb ischemia was noted in the patient's lower extremities. after ventricular tachycardia/pulseless electrical activity arrest, the patient expired. there is not enough information to associate use of the device with patient outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ section: contraindications (united states). ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the reported limb ischemia, access site bleed, and the ventricular tachycardia (vt) has been completed. the impella has not been returned for evaluation. the root cause of access site bleeding, limb ischemia, and the vt was unable to be determined as limited clinical details were provided and no product was returned for review. impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ h5 changed to yes for labeled for single use in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
26148 | 1220648-2024-11629 | 19392281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19392281&pc=OZD | 11/04/2021 | 03/25/2025 |  | 05/24/2024 |  |  | 11/04/2021 | 05/24/2024 |  |  | 09/30/2022 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022130492 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/04/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the device was at a pump speed of p2 and a flow of 2.3 liters. the patient was administered low doses of catecholamines due to observed limb ischemia. the impella cp device was weaned and explanted. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
26149 | 1220648-2024-11630 | 19392444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19392444&pc=OZD | 12/29/2021 | 03/25/2025 |  | 05/24/2024 |  |  | 12/29/2021 | 05/24/2024 |  |  | 09/30/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142412 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/29/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was oozing noted at the insertion site. manual pressure was held for ten (10) minutes, and the site continued to ooze. the impella was weaned. patient remained hemodynamically stable on p2. the impella was pulled across the valve and the re-access port was used to rewire the artery. the pump was explanted, and the 14 fr x 25 cm sheath was inserted. hemostasis was achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26150 | 1220648-2024-11631 | 19392599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19392599&pc=OZD | 12/30/2021 | 03/25/2025 |  | 05/24/2024 |  |  | 12/30/2021 | 05/24/2024 |  |  | 10/31/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152092 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, access site bleeding was noted and retrograde dissection at first access site was revealed. the wire closed the dissection flap and flow was restored. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26151 | 1220648-2024-11632 | 19392675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19392675&pc=OZD | 01/18/2022 | 03/25/2025 |  | 05/24/2024 |  |  | 01/18/2022 | 05/24/2024 |  |  | 09/30/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022145828 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced a rectal bleed and ooze from impella cp groin site which was resolved with angle-matching. intermittent rounds of advanced cardiac life support protocol and mass transfusion protocol were initiated. the patient stabilized before being transferred to the operating room and remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26152 | 1220648-2024-11633 | 19392676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19392676&pc=OZD | 04/26/2024 | 06/11/2024 |  | 05/24/2024 |  |  | 04/26/2024 | 05/24/2024 |  |  | 06/30/2025 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024361100 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2024 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed but explanted and the patient was escalated to the impella 5.5 pump after two days. the team observed vascular damage at the site and associated bleeding. the surgeon conducted surgical repair of the site. the injury was linked to the original impella cp implant. the patient's family decided to withdraw care and the patient expired. per the medical safety office review, the patient required surgical correction - intervention could not be excluded from the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation into the vascular damage - peripheral vessel has been completed since the original report was submitted. the device was not returned for investigation. as per the clinical, significant bleeding occurred when the pocket was opened during explant, requiring a vascular surgeon's intervention. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided.
26153 | 1220648-2024-11634 | 19392720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19392720&pc=OZD | 01/18/2022 | 03/25/2025 |  | 05/24/2024 |  |  | 01/18/2022 | 05/24/2024 |  |  | 09/30/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022145518 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was reported that the echocardiogram showed the patient had a thrombus in the left ventricle. the decision was made to remove the impella. the patient was taken back to the cath lab and the impella was removed, and hemostasis was achieved via x2 percloses. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | Thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
26154 | 1220648-2024-11635 | 19392731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19392731&pc=OZD | 12/30/2021 | 03/25/2025 |  | 05/24/2024 |  |  | 12/30/2021 | 05/24/2024 |  |  | 09/30/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140936 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a small hematoma was noted at the groin. no active bleed was observed. the physician decided to place the patient on argatroban per extracorporeal membrane oxygenation protocol and leave purge solution as dextrose 5% in water. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26155 | 1220648-2024-11636 | 19392846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19392846&pc=OZD | 12/31/2021 | 05/24/2024 |  | 05/24/2024 |  |  | 12/31/2021 | 05/24/2024 |  |  | 10/31/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151457 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/31/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had bleeding complication over two days. the second day of bleeding at the impella access site was also noted to have systemic heparin on hold so heparin was stopped. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26156 | 1220648-2024-11637 | 19393074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393074&pc=OZD | 01/01/2022 | 05/24/2024 |  | 05/24/2024 |  |  | 01/01/2022 | 05/24/2024 |  |  | 06/30/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022101726 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/01/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and after the pump was repositioned, the patient had continued bleeding at the access site. manual pressure was held, and a hematoma was observed. the device was explanted and replaced with an intra-aortic balloon pump via the right femoral artery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26157 | 1220648-2024-11638 | 19393282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393282&pc=OZD | 01/01/2022 | 05/24/2024 |  | 05/24/2024 |  |  | 01/01/2022 | 05/24/2024 |  |  | 10/31/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022153479 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/01/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there was bleeding from the groin, so heparin was removed from purge fluid and the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26158 | 1220648-2024-11639 | 19393339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393339&pc=OZD | 05/03/2024 | 07/19/2024 |  | 05/24/2024 |  |  | 05/03/2024 | 05/24/2024 |  |  | 06/30/2025 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024393455 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Unknown |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Disability | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp placement, the pump was retracted back into the descending aorta, and the other impella pump was implanted through the other femoral access for support. the first impella was pulled out from the ventricle with an additional loop catheter. the pump had been positioned incorrectly and too deeply in the left ventricle for unknown reasons and could not be easily repositioned. during the removal of the first impella pump, according to the implanter, the papillary muscle was torn. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the heart valve damage and the positioning issue has been completed. an impella cp was returned for evaluation. upon visual inspection, a kink in the cannula was noted. no data logs were returned for evaluation. clinical details noted that pump was placed too deep in ventricle and was unable to be easily repositioned. upon removing the pump from the ventricle, the papillary muscle was torn. it is unclear if this tear was due to the impella. the root cause of the heart valve damage was most likely due to improper positioning of the pump too deep in the ventricle. the root cause of the positioning issues was most likely due to use with the pump being placed too deep in the ventricle. the instructions for use referenced in the initial report is for the impella cp with smart assist in the following sections: section: potential adverse events (united states), section: warnings & cautions: warnings, section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves g1-corrected mfr site name and address.
26159 | 1220648-2024-11640 | 19393398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393398&pc=OZD | 04/01/2023 | 07/12/2024 |  | 05/24/2024 |  |  | 05/02/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | In (b)(6) 2017 to (b)(6) 2021 impella pumps supported patients at medical center in the uk. in (b)(6) 2024 the publication entitled: "¿long-term¿ use of impella - safe to do?" in which impella supports were analyzed. the complications noted during impella cp support were hemolysis, access site bleeding, thrombocytopenia, vascular complications, device migration, pump thrombosis, device malfunction, ventricular arrhythmias, access site infection, sepsis, and cva. the adverse events listed in the above mentioned article are reported in complaint (b)(4). | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ zeschky, c., ahmed, h., alayyar, m., jothidasan, a., husain, m., stock, u., & smail, h. (2022). ¿long-term¿ use of impella - safe to do? the journal of heart and lung transplantation, 41(4). https://doi.org/10.1016/j.healun.2022.01.1603. | The investigation for the reported pump stop has been completed. the pump nor the data logs were returned for evaluation. clinical details were not sufficient to determine the root cause. therefore, the root cause of the pump stop could not be determined. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high-risk cpi from the following sections: section: using the automated impella controller with the impella catheter section: using the automated impella controller with the impella rp with smartassist system catheter section: warnings, cautions, and precautions g1-corrected mfr site name and address.
26160 | 1220648-2024-11641 | 19393409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393409&pc=OZD | 01/03/2022 | 05/24/2024 |  | 05/24/2024 |  |  | 01/03/2022 | 05/24/2024 |  |  | 09/30/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140088 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/03/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had bleeding at the groin access site. manual pressure was held, a stitch was placed, and the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26161 | 1220648-2024-11642 | 19393443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393443&pc=OZD | 01/11/2024 | 07/11/2024 |  | 05/24/2024 |  |  | 05/06/2024 | 05/24/2024 |  |  | 11/30/2025 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024425075 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured thrombocytopenia with a "possible" relationship to the impella device used: ¿patient was monitored for thrombocytopenia after impella implant. lowest platelet count was 40. highest was 341. baseline unknown but admission was 111. ¿it was reported that the patient/event has not recovered/not resolved. | It was further reported that care was withdrawn from the patient and the patient subsequently expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for thrombocytopenia has been completed. the impella device was not returned for investigation. thrombocytopenia was reported during pump use. the cause of the thrombocytopenia issue was not determined since product was not returned and sufficient clinical information was not provided. added- b5 mso review, h6 component code 925 d4-corrected model number f6/f8 should have been left blank in the initial report.
26162 | 1220648-2024-11643 | 19393492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393492&pc=OZD | 05/09/2024 | 06/24/2024 |  | 05/24/2024 |  |  | 05/09/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025441799 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient experienced ventricular fibrillation during impella 5.5 support. at the time of the noted issue, the impella alarmed with a position in aorta alarm and support was dropped to p2. the patient was defibrillated, and support levels were increased following the intervention. a second episode of tachycardia was triggered overnight coinciding with a loss of consciousness. the patient was once again defibrillated until stabilized. it was noted that weaning attempts were being made with the impella at the time of the noted episodes. the power level was increased, but intermittent episodes of vt (ventricular tachycardia) persisted until the pump was eventually explanted eight days later. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation into the vf/vt/af/at (torsades/vfib) issue has been completed since the original report was submitted. the product was not returned for investigation. as the necessary information was not provided, the root cause of the vt/vf was not determined.
26163 | 1220648-2024-11644 | 19393640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393640&pc=OZD | 08/31/2019 | 07/26/2024 |  | 05/24/2024 |  |  | 05/02/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaint from the literature review reported 130 patients supported by 2.5, cp, 5.0 or rp, complications noted during hospital stay, acute kidney injury occurred in 56.7% of the patients; need of renal replacement therapy (rrt) 31.7%; access site-related bleeding 14.3%; life-threatening bleeding 31.5%; acute limb ischemia 14.5%; hemolysis 33.3%. the rate of all-cause mortality at 30 days was 39.7%. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ pazzanese, v., ancona, m. b., bertoldi, l. f., pagnesi, m., marini, c., gramegna, m., montorfano, m., chieffo, a., pappalardo, f., & camici, p. g. (2019). p1712the impella percutaneous mechanical circulatory support device in cardiogenic shock: a single-center, real-world, observational experience. european heart journal, 40(supplement_1). https://doi.org/10.1093/eurheartj/ehz748.0467 | D.4 expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. the investigation for the access site bleeding, limb ischemia, and hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding and hemolysis could not be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. a.3 revised as selection was inadvertently omitted on initial manufacturer device report number 1220648-2024-11644. b.3 date of event revised as previously entered incorrectly on initial manufacturer device report number 1220648-2024-11644. g.1 revised manufacturing site name and address section as previously entered incorrectly on initial manufacturer device report number 1220648-2024-11644. h.6 added code 925 as it was inadvertently left off the previously submitted manufacturer device report number 1220648-2024-11644. code 4620 was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-11644. h.10 d.4, d.6a, d.6b and h.4 unknown statements were inadvertently omitted from the initial manufacturer device report 1220648-2024-11644. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report 1220648-2024-11644 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
26164 | 1220648-2024-11645 | 19393664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393664&pc=OZD | 04/11/2022 | 05/28/2025 |  | 05/24/2024 |  |  | 04/11/2022 | 05/24/2024 |  |  | 12/31/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023175251 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient was found to be bleeding from the mouth and was oozing from the access sites. as a result, heparin was stopped in the purge. intermittent albumin was added, the bleeding issues were resolved and heparin was added back to purge. as the patient's platelet level was down to 102, a heparin-induced thrombocytopenia panel was taken, and the patient was given one unit of packed blood cells overnight. the patient continued on support until being bridged to a left ventricular assist device. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ thrombocytopenia: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ vascular dissection: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26165 | 1220648-2024-11646 | 19393666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393666&pc=OZD | 05/02/2024 | 07/15/2024 |  | 05/24/2024 |  |  | 05/02/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 EU |  |  | OZD |  |  | IMPELLA 5.0 EU |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaint from the literature review reported 130 patients supported by 2.5, cp, 5.0 or rp, complications noted during hospital stay, acute kidney injury occurred in 56.7% of the patients; need of renal replacement therapy (rrt) 31.7%; access site-related bleeding 14.3%; life-threatening bleeding 31.5%; acute limb ischemia 14.5%; hemolysis 33.3%. the rate of all-cause mortality at 30 days was 39.7%. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ pazzanese, v., ancona, m. b., bertoldi, l. f., pagnesi, m., marini, c., gramegna, m., montorfano, m., chieffo, a., pappalardo, f., & camici, p. g. (2019). p1712the impella percutaneous mechanical circulatory support device in cardiogenic shock: a single-center, real-world, observational experience. european heart journal, 40(supplement_1). https://doi.org/10.1093/eurheartj/ehz748.0467. | The investigation for the access site bleed, limb ischemia, hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed, limb ischemia, hemolysis could not be determined. g1-corrected mfr site name and address.
26166 | 1220648-2024-11647 | 19393695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393695&pc=PYX | 03/10/2020 | 06/24/2024 |  | 05/24/2024 |  |  | 05/02/2024 | 05/24/2024 |  |  | 03/31/2021 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1421755 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Publication received indicated a patient experiencing postcardiotomy cardiogenic shock was implanted with an impella rp device for mechanical and circulatory support and experienced ventricular fibrillation. details of the event from the publication noted following mitral valve replacement repair and then air embolism, the right ventricle collapsed and right ventricle impella rp pump was placed for support. during the placement of the impella rp pump, the patient had refractory ventricular fibrillation and an electrical storm despite 20 electrical cardioversions. the patient had pulse oximetry readings of 100% but had right-sided jugular venous distention along with a cyanotic face and upper extremities. immediately after placement of the impella rp, the patient's ventricular fibrillation ended. adjunctively the patient was treated with medications. the patient made a full recovery and was discharged to an acute rehabilitation facility. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation into the vf/vt/af/at (refractory ventricular fibrillation) has been completed since the original report was submitted. the device was not returned for investigation. the most likely cause of the guidewire damage was use related. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
26167 | 1220648-2024-11648 | 19393772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393772&pc=OZD | 07/25/2021 | 07/29/2024 |  | 05/24/2024 |  |  | 05/02/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | On (b)(6) 2024 complaints team forwarded publication out of bonn regarding 60 impella patients, 40 patients for cp . the entitled: "role of acquired von willebrand syndrome in the development of bleeding complications in patients treated with impella rp devices" detailed complications as 9 major bleed, 5 ecmo placements, 0 access site bleed, and 10 vw syndrome. patients were those implanted from (b)(6) 2019 to (b)(6) 2020. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ oezkur, m., reda, s., rühl, h., theuerkauf, n., kreyer, s., duerr, g. d., charitos, e., silaschi, m., medina, m., zimmer, s., putensen, c., & treede, h. (2021). role of acquired von willebrand syndrome in the development of bleeding complications in patients treated with impella rp devices. scientific reports, 11(1). https://doi.org/10.1038/s41598-021-02833-8. | A.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. the investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed was unable to be determined as insufficient clinical details were provided. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.2 revised as previously entered selection on the initial manufacturer device report number 1220648-2024-11648 was incorrect. b.3, e.1 zip code and g.1 manufacturing site name and address section were revised as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-11648. h.6 code 4620 was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-11648. a new code was added to health effect - impact codes. code 925 was added to component code since the initial manufacturer device report number 1220648-2024-11648 in accordance with updated procedures. h.10 a.1, a.4, a.5,d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) #, d.6a, d.6b and h.4 are unknown statements were inadvertently omitted from the initial manufacturer device report 1220648-2024-11648. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-11648 in accordance with updated procedures.
26168 | 1220648-2024-11649 | 19393784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393784&pc=OZD | 04/11/2022 | 03/25/2025 |  | 05/24/2024 |  |  | 04/11/2022 | 05/24/2024 |  |  | 01/31/2024 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182516 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support.after support, the device was successfully weaned and explanted. the patient experienced an access site bleed. preclose tightened and manual pressure for 10 mins with femostop at the end to achieve hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26169 | 1220648-2024-11650 | 19393796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393796&pc=OZD | 05/02/2024 | 07/15/2024 |  | 05/24/2024 |  |  | 05/02/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Ventricular Fibrillation | Thrombocytopenia | Thrombosis/Thrombus | [6 empty] | Disability | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | In 5/2017 to 5/2021 impella pumps supported patients at medical center in the uk. in (b)(6) 2024 the publication entitled: "¿long-term¿ use of impella - safe to do?" in which impella supports were analyzed. the complications noted during impella cp support were hemolysis, access site bleeding, thrombocytopenia, vascular complications, device migration, pump thrombosis, device malfunction, ventricular arrhythmias, access site infection, sepsis, and cva. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ zeschky, c., ahmed, h., alayyar, m., jothidasan, a., husain, m., stock, u., & smail, h. (2022). ¿long-term¿ use of impella - safe to do? the journal of heart and lung transplantation, 41(4). https://doi.org/10.1016/j.healun.2022.01.1603. | The investigation for the hemolysis, access site bleeding, thrombocytopenia, vascular complications, device migration, pump thrombosis, device malfunction, ventricular arrhythmias, access site infection, sepsis, and cva has been completed. without additional clinical details nor the returned device, the root cause of the issues could not be determined. the instructions for use referenced in the initial report is for the impella 5.0 in the following sections: section: general patient care considerations. section: entering cardiac output. section: potential adverse events (united states). section: hemolysis. section: table 3.2 impella catheter components. section: warnings & cautions: cautions.
26170 | 1220648-2024-11651 | 19393811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393811&pc=PYX | 04/01/2023 | 07/29/2024 |  | 05/24/2024 |  |  | 05/02/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP EU |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Ventricular Fibrillation | Cardiac Perforation | Thrombocytopenia | Thrombosis/Thrombus | [5 empty] | Required Intervention | Life Threatening | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | On 2024-05-02 complaints forwarded abstract from journal heart and lung transplantation entitled: "¿long-term¿ use of impella - safe to do?" in which impella supports from 5/2017 to 5/2021 were analyzed. the complications noted during impella rp support were hemolysis, access site bleeding, thrombocytopenia, vascular complications, ventricular arrhythmias, access site infection, sepsis, and cva. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ zeschky, c., ahmed, h., alayyar, m., jothidasan, a., husain, m., stock, u., & smail, h. (2022). ¿long-term¿ use of impella - safe to do? the journal of heart and lung transplantation, 41(4). https://doi.org/10.1016/j.healun.2022.01.1603. | A.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. the investigation for the hemolysis, access site bleeding, thrombocytopenia, vascular damage, ventricular tachycardia/ventricular fibrillation (vt/vf), infection/sepsis, thrombus and stroke/cerebrovascular accident (cva). has been completed. the root cause of the hemolysis, access site bleeding, vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the cause of the thrombocytopenia was not determined due to insufficient clinical details. as the necessary clinical information was not provided, the root cause of the vt/vf, infection/sepsis, stroke/cva and thrombus was not determined. as noted in the impella instructions for use: impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella rp smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ b.2 required intervention was previously selected incorrectly on the initial manufacturer device report number 1220648-2024-11651. b.3 and g.1 manufacturing site name and address section were revised as previously entered incorrectly on the initial manufacturer device report number 1220648-2024-11651. h.6 code 4620 was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-11651. a new code was added to health effect - impact codes. code 925 was added to component code since the initial manufacturer device report number 1220648-2024-11651 in accordance with updated procedures. h.10 a.1, a.4, a.5,d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) #, d.6a, d.6b and h.4 are unknown statements were inadvertently omitted from the initial manufacturer device report 1220648-2024-11651. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-11651 in accordance with updated procedures.
26171 | 1220648-2024-11653 | 19393909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19393909&pc=OZD | 05/16/2024 | 07/10/2024 |  | 05/24/2024 |  |  | 05/16/2024 | 05/24/2024 |  |  | 09/30/2025 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024397036 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | DISTAL PERFUSION CATHETER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The user facility reported a 55-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had minimal distal flows following impella cp placement. a distal perfusion catheter was placed to help manage the condition. support with the implanted pump continued following the intervention. the patient was reported as stable. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
26172 | 1220648-2024-11654 | 19394041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19394041&pc=OZD | 05/06/2024 | 07/08/2024 |  | 05/24/2024 |  |  | 05/06/2024 | 05/24/2024 |  |  | 09/30/2025 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025481503 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during implant procedure, the impella cp could not be started. manual pump start was not possible, and automated impella controller (aic) displayed 'impella stopped" error message. the pump was therefore explanted and new one was placed without any issue. there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the pump did not start has been completed. the impella cp was returned for evaluation. upon visual inspection of the pump, the motor cable lines at the base of the motor housing were found to be twisted over each other and one of the motor cable lines was found to be necking. electrical testing was performed and was found to be in range at the time of testing with the pump sitting untouched on the lab bench. the wire inside the necked was found to be broken and when tension was pulled on the motor cable line, the inner wire would fully separate and result in an open line on electrical testing. once tension in wire was removed, electrical testing of the motor cable line would read in range. pump was run in lab in a bucket of water for approximately 5 minutes, when pump was being repositioned in bucket, a pump stop occurred. pump was able to be restarted and each time the pump catheter was manipulated, a pump stop occurred. data logs from the field showed the pump was inserted into the patient and attempted to be restarted 6 times, each time alarm #119 "impella stopped" alarm was triggered. the root cause of the pump not starting was due to an intermittent electrical open motor cable line. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high risk cpi in following sections: section: section: using the automated impella controller with the impella catheter section: using the automated impella controller with the impella rp with smartassist system catheter section: impella stopped added- d9 device return date, h6 impact code 4629 g1 corrected mdr reporting address g1-updated mfr site name and address.
26173 | 1220648-2024-11655 | 19394166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19394166&pc=OZD | 05/13/2024 | 06/24/2024 |  | 05/24/2024 |  |  | 05/13/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025485737 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella cp support had no distal flow requiring a fem-fem bypass. it was further reported that the patient is oozing from all access sites including impella site. oozing was managed by target act management. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into limb ischemia ¿ reversible and the access site bleeding ¿ major issues has been completed since the original report was submitted. the product was not returned for investigation. the root cause of access site bleeding was determined to be patient condition because the bleeding was reported from all the access sites. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
26174 | 1220648-2024-11656 | 19394313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19394313&pc=OZD | 05/10/2024 | 07/17/2024 |  | 05/24/2024 |  |  | 05/10/2024 | 05/24/2024 |  |  | 11/30/2025 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430442 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the 68-year-old male patient experienced oozing from the impella groin when patient moved leg because of symptomatic syndrome. the patient's hemoglobin dropped to 9.6 from 13. the sheath was repositioned to resolve the issue. no blood products were given. the device was not removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Section h6 has been updated, as the investigation has been completed. according to clinical, there was oozing from impella groin due leg movement by the patient because of symptomatic syndrome. hemoglobin dropped to 9.6 from 13 and repositioning of sheath was completed. impella running stable and was not removed. the impella device was not returned and therefore, an evaluation of the device was not possible. product history review was completed, and the pump passed all sterile inspection checks. the root cause of access site bleeding is determined to be patient movement because the bleeding was caused by patient leg movement.
26175 | 1220648-2024-11657 | 19394384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19394384&pc=OZD | 05/10/2024 | 07/09/2024 |  | 05/24/2024 |  |  | 05/10/2024 | 05/24/2024 |  |  | 09/30/2025 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024406313 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the female patient experienced ischemic signs. the patient's feet were marbled, and lactate was 1.2. no actions have been taken, and the device remained implanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the section: potential adverse events (united states). g1-corrected mfr site name, address line 1, mfr site city. removed mfr address line 2.
26176 | 1220648-2024-11658 | 19394510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19394510&pc=OZD | 05/08/2024 | 06/26/2024 |  | 05/24/2024 |  |  | 05/10/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1474181 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a battery error is displayed on the automated impella controller (aic). the battery was not charging. the aic was previously sent for maintenance. when the aic arrived back at the hospital, it was not connected to the socket. when the aic was then plugged into the socket, it was noticed that the aic was not charging, and the battery switch was not set to "off". even after a longer charging process (3 days), the battery could not be charged, even when attempts were made to switch the battery switch back and forth between "on" and "off" several times. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation has been completed since the initial report was submitted. the console was tested and battery failure alarm issue was able to be reproduced. after reviewing the logs, the battery error issue was confirmed. no logs were found from the reported occurrence date but there were logs from two days prior. there were numerous instances of battery failure alarm (transport connection blown/missing) found from this date. results of running batttest on telnet revealed that ¿fu¿ was missing on both batteries. issue was determined to be caused by depleted batteries. console was last disconnected from alternating current power on 2023.09.19 and was never plugged back to ac power since then. these depleted batteries could not be charged back up because they were locked out due to low cell voltage (less than 2.3v). the root cause of this issue was the automated impella controller not being routinely charged. a.3 revised from the initial submission in accordance with updated procedures. b.3 date of event revised as previously entered incorrectly. d.1 brand name and d.4 serial name were revised as previously entered incorrectly. d.2 common device name revised from the initial submission in accordance with updated procedures. g.1 revised manufacturing site name, address line 1 and 2 and city as previously entered incorrectly. h.6 medical device problem code was revised in accordance with updated procedures. h.8 usage of device was revised as previously entered incorrectly.
26177 | 1220648-2024-11659 | 19394579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19394579&pc=OZD | 04/26/2024 | 06/30/2024 |  | 05/24/2024 |  |  | 04/26/2024 | 05/24/2024 |  |  | 08/31/2025 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024390842 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2024 |  |  |  |  | The complainant reported a 78-year-old female patient with post-cardiotomy cardiogenic shock had an attempted implant with an impella 5.5 for mechanical circulatory support. the complainant reported delivery issues during the implant ¿ the doctor had a difficult time getting a wire across the anastomosis and valve, but eventually succeeded. the impella would not cross the anastomosis of graft to axillary artery after multiple attempts. the patient eventually expired. there is not enough information to associated the use of the device with the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related. the impella would not cross the anastomosis of graft to the axillary artery after multiple attempts as per the clinical description. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ b.5 the patient's age was reported incorrectly on the previous report. the patient's age is 63 years old. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 codes 463 and 3038 were reported incorrectly on the previously submitted report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
26178 | 1220648-2024-11652 | 19394638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19394638&pc=OZD | 08/23/2022 | 03/25/2025 |  | 05/24/2024 |  |  | 08/23/2022 | 05/24/2024 |  |  | 07/31/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022105858 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/23/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced limb ischemia and required surgical closure of the puncture site to stop bleeding. it is unknown if the ischemia was resolved and whether the device was explanted due to the reported issues, but the patient survived. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26179 | 1220648-2024-11660 | 19394750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19394750&pc=OZD | 05/10/2024 | 07/29/2024 |  | 05/24/2024 |  |  | 05/10/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025488232 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that 60-year-old female patient was implanted with impella rp and impella cp in bail-out situation, as otherwise the patient would have expired. the patient was hemodynamically stable at very low level under mcs and received massive administration of inotropics and vasopressors. the patient's liver function was noted to be already severely impacted. the partial thromboplastin time (ptt) spontaneously increased, and the patient began to severely bleed from the puncture site, leading to the patient receiving several emergency blood bags. the situation did not improve and decision was made to withdraw care. additional information was received, indicating that bleeding was due to existing multi-organ failure, most specifically, liver failure. the devices will not be returned as they were discarded by the site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was anticoagulation management due to high patient act values, heparin was stopped to achieve hemostasis as per the clinical description. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. added- h6 component code 925 g1-corrected mfr site name and address.
26180 | 1220648-2024-11661 | 19395818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19395818&pc=OZD | 09/05/2022 | 05/22/2025 |  | 05/24/2024 |  |  | 09/05/2022 | 05/24/2024 |  |  | 07/31/2024 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237708 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | Insufficient Information | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/05/2022 |  |  |  |  | The complainant reported that a patient was implanted with an impella cp support for postcardiotomy cardiogenic shock / low cardiac output syndrome. while on support, the pump was had positioning issues with no specified resolution reported. the patient continued on support until the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26181 | 1220648-2024-11662 | 19396212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19396212&pc=OZD | 05/15/2024 | 06/27/2024 |  | 05/24/2024 |  |  | 05/15/2024 | 05/24/2024 |  |  | 11/30/2025 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025436809 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp and ecmo (extracorporeal membrane oxygenation) supporting a patient admitted in cgs (cardiogenic shock) and cardiomyopathy. the patient had the cp placed and the patient was then transferred out to the tertiary center for care escalation and consideration of vad (ventricular assist device)/transplant. after 82 hours the pump was explanted due to a cold leg. the limb ischemia had been followed and troubleshot but, the pump was explanted and amputation above knee (aka) performed. the pump was explanted and hd catheter placed. no new pump was placed as axillary not an option nor was transplant or vad - patient not a candidate. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into the limb ischemia ¿ irreversible issue been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
26182 | 1220648-2024-11664 | 19396915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19396915&pc=OZD | 09/06/2022 | 05/24/2024 |  | 05/24/2024 |  |  | 09/06/2022 | 05/24/2024 |  |  | 12/31/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023179541 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | Hemolysis | Perforation of Vessels | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/06/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bloody nasal/oral secretions and hemolysis. additionally, delivery issues were noted with low pump flow. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26183 | 1220648-2024-11663 | 19397053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19397053&pc=OZD | 05/16/2024 | 07/22/2024 |  | 05/24/2024 |  |  | 05/16/2024 | 05/24/2024 |  |  |  |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025465400 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a left ventricle thrombus was observed and a heparin bolus was given. additionally, the impella cp pump was observed close to the mitral valve, and repositioning was suggested. it is unknown if the pump was successfully repositioned as the pump was shortly removed after the positioning issue was discovered. the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system-section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). | The investigation for the thrombus and positioning issue has been completed. the impella device was not returned for investigation. data logs were reviewed and found many suction alarms occurred. positioning issue alarms occurred at the time of reported issue. the cause of the positioning issue most likely use related due to improper technique. as the necessary information was not provided, the root cause of the thrombus was not determined. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." d.4 revised model number from the initial submission of manufacturer device report 1220648-2024-11663 in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report 1220648-2024-11663 in accordance with updated procedures.
26184 | 1220648-2024-11665 | 19397820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19397820&pc=OZD | 01/18/2022 | 03/25/2025 |  | 05/24/2024 |  |  | 01/18/2022 | 05/24/2024 |  |  | 10/31/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154788 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was reported that the patient was having hematuria. the physician decreased the p level and an echocardiogram was performed, which showed that the device was in good position. fluids were also given, and it was noted that the patient had pink urine. staff turned the impella to p2 and there were no alarms, and an echocardiogram now showed that the impella was 3.2cm from the aortic valve. the patient had a deep left ventricle cavity, and the physician suspected the device could be advanced approximately 1cm further. it is unknown whether there was any attempt to advance the device, but as the patient was hemodynamically stable, the physician then decided to explant the device. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: hemolysis: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26185 | 1220648-2024-11666 | 19397843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19397843&pc=OZD | 01/18/2022 | 03/25/2025 |  | 05/24/2024 |  |  | 01/18/2022 | 05/24/2024 |  |  | 10/31/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156302 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was noted that the groin access site was oozing and the patient was experiencing hemolysis, so heparin was withheld and a femstop was placed. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26186 | 1220648-2024-11667 | 19397870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19397870&pc=OZD | 01/20/2022 | 05/13/2025 |  | 05/24/2024 |  |  | 01/20/2022 | 05/24/2024 |  |  | 04/30/2022 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1488338 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/20/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the sheath was removed with the remaining sheath and bleeding was observed at the access site. one pre-sutured and two post-sutured sutures were applied, pressure was applied to stop the bleeding for several minutes, and the procedure was completed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26187 | 1220648-2024-11668 | 19397885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19397885&pc=OZD | 04/15/2022 | 03/25/2025 |  | 05/24/2024 |  |  | 04/15/2022 | 05/24/2024 |  |  | 01/31/2024 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181522 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, groin site appears soft with very minimal amount of oozing. hemostasis achieved with perclose x2 and one 8fr angioseal. the patient was successfully weaned off of support and the device explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26188 | 1220648-2024-11669 | 19397904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19397904&pc=OZD | 09/11/2022 | 05/24/2024 |  | 05/24/2024 |  |  | 09/11/2022 | 05/24/2024 |  |  | 04/30/2024 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226476 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/11/2022 |  |  |  |  | The complainant reported that a 76 year old male patient was implanted with an impella cp for mechancial circulatory support. while on support the patient had access site bleeding and gastrointestinal bleeding. the patient was given two units of blood, with no systemic heparin at the time. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26189 | 1220648-2024-11670 | 19397928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19397928&pc=OZD | 09/14/2022 | 05/20/2025 |  | 05/24/2024 |  |  | 09/14/2022 | 05/24/2024 |  |  | 05/31/2024 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230286 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2022 |  |  |  |  | The complainant reported a 72 year old female patient was implanted with an impella cp for mechanical circulatory support due to post cardiotomy cardiogenic shock and low cardiac output syndrome. while on support, the pump experienced positioning issues that were resolved by resetting with a new sheath. additionally, the patient experienced ventricular tachycardia and atrial fibrillation that were resolved with cardioversion. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26190 | 1220648-2024-11671 | 19397931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19397931&pc=OZD | 10/14/2021 | 03/25/2025 |  | 05/24/2024 |  |  | 10/14/2021 | 05/24/2024 |  |  | 07/31/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022115162 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, during and after introducer insertion the sheath was occlusive despite caliber being quite large. vessels did have a fair amount of tortuosity and it was thought that the use of the long sheath may have caused this issue. the case continued and the 5th stent was placed. the impella was weaned and removed without issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: limb ischemia impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
26191 | 1220648-2024-11672 | 19397967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19397967&pc=OZD | 10/14/2021 | 03/25/2025 |  | 05/24/2024 |  |  | 10/14/2021 | 05/24/2024 |  |  | 07/31/2023 |  | 05/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110520 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during closure, 1st perclose suture was lost, 2nd perclose was unsuccessful at achieving complete hemostasis, a 3rd perclose was attempted and the footplate pinched a segment of the cfa and ripped out a 1 1/2 inch portion of the artery. 18fr cook sheath was inserted and surgeon called for surgical closure. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26192 | 1220648-2024-08589 | 19380430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19380430&pc=OZD | 03/13/2024 | 08/15/2024 |  | 05/23/2024 |  |  | 03/13/2024 | 05/23/2024 |  |  | 09/30/2025 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444666 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and the patient had a posterior wall injury and the sheath was through the ligament. the injury caused retroperitoneal bleeding. the bleed went unnoticed until it was extremely large. the pump was removed, vascular surgeons completed the closure and multiple blood products were provided. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the vascular damage was most likely due to the user, as the angle of sheath insertion was too high. d6a and d6b added. h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-08589. d4 model number. f6/f8-should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
26193 | 1220648-2024-11550 | 19380619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19380619&pc=OZD | 10/30/2021 | 03/31/2025 |  | 05/23/2024 |  |  | 10/30/2021 | 05/23/2024 |  |  | 07/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022123873 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/30/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella access site. the patient's heparin was reduced until bleeding was controlled. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26194 | 1220648-2024-10662 | 19380693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19380693&pc=PYX | 03/01/2024 | 05/30/2024 |  | 05/23/2024 |  |  | 04/23/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Study | Literature |  |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | A publication was shared with abiomed ¿ benefits of a novel percutaneous ventricular assist device for right heart failure: the prospective recover right study of the impella device. the publication contained a retrospective review of 175 total patients screened for 30 patient enrollments in recover right. the publication stated bleeding was the most prevalent adverse event in the study. device-access related bleeding was observed in one patient. bleeding during insertion of the device was low (93% of the patients lost 100 ml of blood for the combined introducer sheath placement and the impella rp insertion). | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is for the impella rp with smart assist system with automated impella controller in section: warnings and cautions. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿
26195 | 1220648-2024-11551 | 19380695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19380695&pc=OZD | 10/31/2021 | 05/23/2024 |  | 05/23/2024 |  |  | 10/31/2021 | 05/23/2024 |  |  | 07/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022115456 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/31/2021 |  |  |  |  | The complainant reported a patient, demographics unknown, was taken to the cardiac catheterization laboratory (ccl) for impella placement. it was reported the implant procedure was aborted due to the patient¿s anatomy and an existing clot in the left ventricle (lv). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions. as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluation" this report is being filed as part of a retrospective review of historical records.
26196 | 1220648-2024-11552 | 19381064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19381064&pc=OZD | 11/01/2021 | 05/23/2024 |  | 05/23/2024 |  |  | 11/01/2021 | 05/23/2024 |  |  | 08/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022132905 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/01/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella insertion was aborted and explanted due to the cheater, a small tapered device used to guide wires into the sheaths, was found in the inlet passage causing the pump to stop. the impella was replaced with a new impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter as noted in the impella ifu ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smart assist system catheter as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
26197 | 1220648-2024-10897 | 19381264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19381264&pc=OZD | 05/02/2024 | 07/03/2024 |  | 05/23/2024 |  |  | 05/02/2024 | 05/23/2024 |  |  | 07/31/2025 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388881 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Damaged by Another Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing st-segment elevation myocardial infarction was prepared for an impella cp device implant for mechanical circulatory support. during the procedure, the incorrect guidewire was handed off for the impella cp pump implant and the guidewire became stuck, could not be removed. both the impella cp pump and guidewire were removed entirely, and a new impella cp pump was successfully implanted. there was no report of harm as a result of this event. | [5 empty] | The investigation for the product damage (bond failure) has been completed. according to the clinical review, during impella cp insertion the wrong guidewire was used. an attempt was made to remove it from the pump unsuccessfully. the decision was then made to remove the pump and replace it with another impella cp device. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the delivery issues was use related. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
26198 | 1220648-2024-11553 | 19381289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19381289&pc=OZD | 10/22/2021 | 03/31/2025 |  | 05/23/2024 |  |  | 10/22/2021 | 05/23/2024 |  |  | 07/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022111931 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2021 |  |  |  |  | The complainant reported that the patient was placed on impella cp support for percutaneous coronary intervention. while on support, the patient developed oozing post redressing in the catheterization lab. the site was re-sutured and looks better. the patient was in atrial fibrillation and an amio bolus given x2 and a drip was started, vasopressin was added. the patient was successfully weaned from support. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26199 | 1220648-2024-11554 | 19381322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19381322&pc=OZD | 10/22/2021 | 03/31/2025 |  | 05/23/2024 |  |  | 10/22/2021 | 05/23/2024 |  |  | 03/31/2022 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1480854 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. while on support, there was a perforation during left femoral artery access, and hematoma, which prevented impella placement. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26200 | 1220648-2024-11555 | 19381396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19381396&pc=OZD | 04/16/2024 | 06/21/2024 |  | 05/23/2024 |  |  | 04/27/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443374 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. while on support, an access site hematoma was observed; the site was iced along with temporary stoppage of heparin. further information noted, some days later, micro-bubbles in the aorta were noticed on echocardiogram. no intervention for the micro bubbles was provided; a repeat echocardiogram was planned. it was noted there was no neurological deficit, and the patient was stable. | [5 empty] | The investigation for the reported access site bleed and embolism has been completed. the pump and purge cassette were returned for investigation. no physical abnormalities were observed on the returned product. priming was performed without issue. the pump was run as expected during the test in a bucket of water. no purge leak was reported during the test. data logs show a air-in-purge alarm during the bag change procedure. no other purge-related abnormalities were reported in the data logs. according to clinical details, a microbubble was noted in the aorta. also, a hematoma was observed, and heparin was on hold. the air in purge alarm occurred on a different day than the air in the ventricle observation. the cause of the embolism issue was not determined since sufficient clinical information was not provided, and its relation to impella was unknown. the cause of the access site bleeding issue was not determined since sufficient clinical information was not provided. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is cw observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
26201 | 1220648-2024-11556 | 19381427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19381427&pc=OZD | 11/01/2021 | 05/23/2024 |  | 05/23/2024 |  |  | 11/01/2021 | 05/23/2024 |  |  | 08/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022122897 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/01/2021 |  |  |  |  | The complainant reported a patient was taken to the cardiac catheterization laboratory (ccl) for impella placement. while on support, the patient experienced access site bleeding which was resolved by redressing the site in the unit and hemostasis achieved. the patient received fluids and an unknown amount of packed red blood cells. it was further reported the team was unable to palpate pedal pulses bilaterally and it was also difficult to find the radial pulses, however they were able to find distal pulses via doppler, but they were fleeting. it was reported that the patients¿ limbs were both ischemic. the team reportedly monitored the distal pulses and lactate levels in the cardiovascular unit. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related events are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26202 | 1220648-2024-11557 | 19381474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19381474&pc=OZD | 04/28/2024 | 06/21/2024 |  | 05/23/2024 |  |  | 04/28/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025461220 |  | 1000100 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy waiting for a heart transplant was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient had a lot of ectopy. the impella pump was repositioned and helped the ectopy. it was noted from moment of impella 5.5 implant, patient appeared under-supported, and physician keep saying "it's like the impella isn't even there." eventually, the team tried bipella support with an impella rp device before adding extracorporeal membrane oxygenation (ecmo) for additional left-sided support. impella and ecmo support was noted to have nicely supported the patient who subsequently successfully receive a heart transplant. | [5 empty] | The investigation into the reported issues has been completed since the original report was filed. the pump was returned for investigation. the root cause of the low flow and the of the vascular damage - peripheral vessel was unable to be determined as limited clinical details were provided, and no data was returned for review. as the necessary information was not provided, the root cause of the vt/vf was not determined. because evidence related to infection was not provided, the root cause cannot be determined. the root cause of positioning issue was most likely use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26203 | 1220648-2024-11558 | 19381530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19381530&pc=OZD | 10/23/2021 | 03/31/2025 |  | 05/23/2024 |  |  | 10/23/2021 | 05/23/2024 |  |  | 07/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022116688 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/23/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was positioned under the papillary muscle. re-positioning was attempted but not successful, so the device was explanted due to pump stoppage. a clot was also observed. patient remained on support and was successfully weaned later. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smart assist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
26204 | 1220648-2024-11559 | 19381569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19381569&pc=OZD | 05/13/2024 | 07/16/2024 |  | 05/23/2024 |  |  | 05/13/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489601 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. however, the impella cp pump began to experience persistent suction issues due to poor positioning. attempts to reposition the device were unsuccessful, and the decision was ultimately made to replace the device. there was no patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined. the instructions for use for the impella cp with smartassist system in section: use of echocardiography for positioning of the impella catheter to ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ d4-corrected model number. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26205 | 1220648-2024-11560 | 19381677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19381677&pc=OZD | 11/01/2021 | 05/23/2024 |  | 05/23/2024 |  |  | 11/01/2021 | 05/23/2024 |  |  | 08/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022129405 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/01/2021 |  |  |  |  | The complainant reported a patient, race unknown, noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a lot of oozing at the end of the procedure, to resolve the surgeon placed purse string sutures around the device to resolve the bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26206 | 1220648-2024-11561 | 19381780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19381780&pc=OZD | 10/23/2021 | 05/27/2025 |  | 05/23/2024 |  |  | 10/23/2021 | 05/23/2024 |  |  | 08/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022126976 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombocytopenia | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/23/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had distal clots. also, there was bleeding, possibly due to cooling blanket used. and the patient¿s toes were cold and treated with plain old balloon angioplasty. moreover, the patient had heparin-induced thrombocytopenia. device weaned and explanted in cath lab. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26207 | 1220648-2024-11562 | 19381834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19381834&pc=OZD | 05/14/2024 | 07/23/2024 |  | 05/23/2024 |  |  | 05/15/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446423 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The user facility reported a patient of unknown age and gender with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that there was a kink in the impella cp. no resistance or excessive force was applied during insertion. the catheter was kinked, which made repositioning impossible. the impella cp was then explanted. the patient survived the event. | [5 empty] | The investigation for positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue was most likely use related due to catheter being kinked during support. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11562. f6/f8-should have been left blank. g1 reporting contact fax number missing. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26208 | 1220648-2024-11563 | 19381921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19381921&pc=OZD | 10/23/2021 | 03/25/2025 |  | 05/23/2024 |  |  | 10/23/2021 | 05/23/2024 |  |  | 07/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022112354 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/23/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, due to already being in the hospital when the patient started to decompensate and went into cardiogenic shock. the patient had rhythm changes in the morning to atrial fibrillation with bursts to 160s. the patient remained on support and was weaned successfully. | [5 empty] | To make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26209 | 1220648-2023-04823 | 19381973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19381973&pc=OZD | 10/08/2023 | 08/15/2024 |  | 05/23/2024 |  |  | 10/08/2023 | 05/23/2024 |  |  | 07/31/2025 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024381122 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/08/2023 |  |  |  |  | The user facility reported a 69 -year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. us complainant reported the 14fr sheath peel away sheath was noted to be 2-3 inches out of the body while in the cath lab. physician did not want to peel away and did not want to advance it. sheath was secured with sutures. a hematoma had formed. pressure was held but bleeding persisted. surgeon advanced the sheath and placed new forward tension sutures. bleeding resolved. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the cause of the access site bleeding issue was most likely use related per clinical review. d6a and d6b added. h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04823. f6/f8-should have been left blank. g1 reporting contact fax number missing. g1 manufacturing contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
26210 | 1220648-2024-11564 | 19382095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19382095&pc=OZD | 11/02/2021 | 03/25/2025 |  | 05/23/2024 |  |  | 11/02/2021 | 05/23/2024 |  |  | 07/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114050 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2021 |  |  |  |  | The user facility reported a 62-year-old female on impella cp for mechanical circulatory support. it was reported the patient had oozing from the impella cp site. systemic heparin was stopped and manual compression was applied. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26211 | 1220648-2024-11565 | 19382155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19382155&pc=OZD | 10/26/2021 | 03/25/2025 |  | 05/23/2024 |  |  | 10/26/2021 | 05/23/2024 |  |  | 07/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2022118248 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2021 |  |  |  |  | The complainant reported that while the patient was on impella cp support, the patient had bleeding at the site. therefore, site was tightened by suture and blood products were given. the patient continued on support until the device was successfully weaned. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore. evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26212 | 1220648-2024-11566 | 19382163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19382163&pc=OZD | 05/15/2024 | 07/12/2024 |  | 05/23/2024 |  |  | 05/15/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492951 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, purge fluid began leaking from the red plug. the site was taped to manage the leak and the pump was weaned and explanted the following day. there was no patient harm. | [5 empty] | The investigation for the reported purge leak-pump has been completed. the pump and purge cassette were returned for investigation. data logs show no trend related to purge pressure abnormalities. the product was received with the taped red handle. a catheter kink was observed after removing the tape. no other physical abnormalities were observed. the pump was primed without observing any purge-related alarms or issues. the pump ran without issue. the red handle was opened for purge leak investigation. purge leak was observed from the joint at inlet purge line kink protector inside the red handle. leaking was observed with very slow rate. the cause of the purge leak issue was bond failure at the inlet purge line kink protector inside the red handle. added- h6 component code 481 f6/f8 should have been left blank in the initial report. h6 impact code 2199 changed to 4641, 4627. h6 med dev prob code 1135 changed to 1250. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the following event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
26213 | 1220648-2024-11567 | 19382221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19382221&pc=OZD | 11/01/2021 | 05/23/2024 |  | 05/23/2024 |  |  | 11/01/2021 | 05/23/2024 |  |  | 02/28/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2022054461 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/01/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported that the impella was removed due to left ventricle thrombus. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves as noted in the impella ifu ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ as noted in the impella ifu "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
26214 | 1220648-2024-11568 | 19382377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19382377&pc=OZD | 10/24/2021 | 03/25/2025 |  | 05/23/2024 |  |  | 10/24/2021 | 05/23/2024 |  |  | 07/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121259 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2021 |  |  |  |  | It was reported the patient was on impella cp for hemodynamic support. while on support thrombocytopenia was noted and the patient was escalated to the impella 5.5 device. | [5 empty] | The cause of the issue was not determined due to insufficient clinical details. this report is being filed as part of a retrospective review of historical records. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26215 | 1220648-2024-11569 | 19382427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19382427&pc=OZD | 10/26/2021 | 05/20/2025 |  | 05/23/2024 |  |  | 10/26/2021 | 05/23/2024 |  |  | 08/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022128022 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2021 |  |  |  |  | The complainant reported the patient was on impella cp support and the pump had to be repositioned due to hemolysis. the patient continued on support until escalation to an impella 5.5. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ section: use of echocardiography for positioning of the impella catheter. impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26216 | 1220648-2024-11570 | 19382474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19382474&pc=OZD | 05/16/2024 | 07/05/2024 |  | 05/23/2024 |  |  | 05/16/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474224 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and while on support, the patient had ventricular tachycardia (vt). the patient was shocked three times. the impella was positioned shallow and it was repositioned. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. the instructions for use referenced in the initial report is for the impella 5.5 with smart assist for use during cardiogenic shock. added- h6 impact code 4628, h6 component code 925 d4-corrected model number f6/f8 should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
26217 | 1220648-2024-11571 | 19382600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19382600&pc=OZD | 10/24/2021 | 05/21/2025 |  | 05/23/2024 |  |  | 10/24/2021 | 05/23/2024 |  |  | 07/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022109413 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient did not have a distal pulse in the left foot, and the staff placed a fem-fem bypass. additionally, the impella exhibited an impella position notification. the device was replaced with an impella 5.5. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26218 | 1220648-2024-11572 | 19382631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19382631&pc=OZD | 10/26/2021 | 05/23/2024 |  | 05/23/2024 |  |  | 10/26/2021 | 05/23/2024 |  |  | 07/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022116088 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/26/2021 |  |  |  |  | The complainant reported that while the patient was on impella cp support, the patient had an episode of ventricular tachycardia. the impella was turned up to p6 to help support the ventricle. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26219 | 1220648-2024-11573 | 19382832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19382832&pc=OZD | 05/14/2024 | 07/17/2024 |  | 05/23/2024 |  |  | 05/15/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486776 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The user facility reported a 64-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was thrombosis and positioning issues noted. the team attempted to use blood thinners to lyse the clot burden, ultimately, the pump was explanted. the patient outcome is unknown. | [5 empty] | The investigation for the thrombosis and positioning issues has been completed. data logs were reviewed which confirmed the failure mode and clinical narrative. logs showed after the pump started and run for about 23 hours, a motor current (mc) spike followed low flow that triggered auto suction response and suction alarms. there was no positioning alarm found on the logs. product was not returned for review. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial ingestion. the root cause of positioning issue was most likely use related. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
26220 | 1220648-2024-11575 | 19382849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19382849&pc=OZD | 11/11/2021 | 05/23/2024 |  | 05/23/2024 |  |  | 11/11/2021 | 05/23/2024 |  |  | 08/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022130694 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/11/2021 |  |  |  |  | The complainant reported 61 year old female patient was placed on impella cp support. it was reported that after initiation of support, the patient experienced bleeding and atrial fibrillation. the patient received two units of packed red blood cells. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). this report is being filed as part of a retrospective review of historical records.
26221 | 1220648-2024-11574 | 19382854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19382854&pc=OZD | 10/24/2021 | 03/25/2025 |  | 05/23/2024 |  |  | 10/24/2021 | 05/23/2024 |  |  | 03/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022060274 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the medical staff was unable to achieve a pump speed level above p-2 without suction alarms and the patient experiencing episodes of ventricular tachycardia with loss of pusatility. an echocardiogram was performed in the catheterization laboratory to reposition the impella cp device. however, the device was unable to advance across the atrioventricular valve without contracting other cardiac structures. it is noted that the patient suffered a ruptured chordae. after numerous efforts to reposition, the decision was made to explant the impella cp device. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance." this report is being filed as part of a retrospective review of historical records.
26222 | 1220648-2024-11576 | 19382875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19382875&pc=PYX | 05/02/2024 | 07/16/2024 |  | 05/23/2024 |  |  | 05/02/2024 | 05/23/2024 |  |  | 10/31/2018 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP EU |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 1260071 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 37 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaint from the literature review reported a bi-pella approach in which a 37-year-old patient had rp-cp support post ischemic stroke. the complication noted for the impella support hemolysis was observed [lactate dehydrogenase (ldh) max value 1240 u/l] initially when impella pumps were at high rotational speed. the patient survived the event. | [5 empty] | The investigation for the hemolysis has been completed. the impella was not returned for evaluation. data logs were reviewed finding no suction alarms observed. no other issues observed to indicated to reported hemolysis issue. the cause of the hemolysis issue cannot be determined since complaint device was not returned for investigation and insufficient clinical data provided. the instructions for use referenced in the initial report is for the impella rp with smart assist system with automated impella controller. g1-corrected mfr site name and address g4-corrected pma number. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ pappalardo, f., scandroglio, a. m., & latib, a. (2018b). full percutaneous biventricular support with two impella pumps: the bi-pella approach. esc heart failure, 5(3), 368¿371. https://doi.org/10.1002/ehf2.12274.
26223 | 1220648-2024-11577 | 19382903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19382903&pc=PYX | 04/21/2024 | 06/04/2024 |  | 05/23/2024 |  |  | 04/21/2024 | 05/23/2024 |  |  | 11/30/2025 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024433041 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/21/2024 |  |  |  |  | The complainant reported a 56 year-old male patient with post-cardiotomy cardiogenic shock was implanted with an impella rp for mechanical circulatory support. the complainant reported high purge pressure. the patient survived the event and was weaned off the rp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. per data analysis, it was reported that there was an increase in purge pressure and a drop in purge flow which did not trigger an alarm. the root cause of the high purge pressure was most likely biomaterial. the root cause of the low pump flow issue was most likely biomaterial.
26224 | 1220648-2024-11578 | 19382984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19382984&pc=OZD | 10/24/2021 | 05/23/2024 |  | 05/23/2024 |  |  | 10/24/2021 | 05/23/2024 |  |  | 08/31/2024 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127528 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Sense | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/24/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient coded multiple times requiring cardio-pulmonary resuscitation. the patient was taken to the operating room where he coded again. placement signal unreliable was noted and the decision was made to pump into aorta and remove with surgical cutdown once weaned from bypass machine. additionally, the patient had a groin hematoma. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26225 | 1220648-2024-11579 | 19383044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383044&pc=OZD | 11/11/2021 | 03/25/2025 |  | 05/23/2024 |  |  | 11/11/2021 | 05/23/2024 |  |  | 07/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022111909 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Fever | Hemorrhage/Blood Loss/Bleeding | Sepsis | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2021 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced moderate to severe oozing even with 2 sutures deployed and the angle being maintained. additionally, medical staff expressed concern for sepsis as indicated by febrility, tachycardia, and increased cardiac output/cardiac index. it is unknown if these reported issues were resolved, but the patient survived the procedures. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding: impella cp with smart assist system. section: general patient care considerations . ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer . ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ infection/sepsis: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
26226 | 1220648-2024-11580 | 19383183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383183&pc=OZD | 05/17/2024 | 07/23/2024 |  | 05/23/2024 |  |  | 05/17/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433415 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The complainant reported a 62-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed hemolysis following the explant. one unit of blood products was given to the patient to resolve the hemolysis. the patient survived the event. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 component code added 925 the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11580: d6b explanted date. g1 reporting contact fax number missing.
26227 | 1220648-2024-11581 | 19383209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383209&pc=OZD | 10/29/2021 | 05/23/2024 |  | 05/23/2024 |  |  | 10/29/2021 | 05/23/2024 |  |  | 07/31/2024 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022119330 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/29/2021 |  |  |  |  | The complainant reported that the patient experienced oozing from both the impella and ecmo sites. the dressing was removed, and the angle of sheath increased; however, there was still slight oozing. the patient received one unit of blood products. the patient continued on support until the impella was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26228 | 1220648-2024-11582 | 19383217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383217&pc=OZD | 01/17/2022 | 03/25/2025 |  | 05/23/2024 |  |  | 01/17/2022 | 05/23/2024 |  |  | 07/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110537 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/17/2022 |  |  |  |  | The complainant reported a patient with myocarditis was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella access site. the oozing was initially control by using a circular mattress suture, but oozing returned. dressing changes and manual pressure was applied, and an additional suture was placed. systemic heparin was withheld and the patient was transfused with multiple units of packed red blood cells. the oozing continued until the impella cp and extracorporeal membrane oxygenation were explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26229 | 1220648-2024-11583 | 19383305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383305&pc=PYX | 02/28/2024 | 07/15/2024 |  | 05/23/2024 |  |  | 04/23/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | Abiomed complaints team received a publication entitled "short-term mechanical circulatory support for recovery from acute right ventricular failure: clinical outcomes", it was noted that two of four patients developed hemolysis during impella rp support. one instance resolved following impella explant and the other case resolved spontaneously after 72 hours of support. | [5 empty] | The impella devices were not received from the customer and therefore, an evaluation of the devices was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported hemolysis has been completed. the device was not returned for evaluation. clinical information was insufficient to determine the root cause. therefore, the root cause of the hemolysis could not be determined. added- h6 component code 925. added citation: cheung, a. w., white, c. w., davis, m. k., and freed, d. h. (2014). short-term mechanical circulatory support for recovery from acute right ventricular failure: clinical outcomes. the journal of heart and lung transplantation, 33(8), 794¿799. https://doi.org/10.1016/j.healun.2014.02.028 f6/f8-should have been left blank. h.6 codes 4641 and 4628 were reported incorrectly on manufacturer device report 1220648-2024-11583.
26230 | 1220648-2024-11584 | 19383370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383370&pc=OZD | 05/17/2024 | 07/15/2024 |  | 05/23/2024 |  |  | 05/17/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025446120 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported an 81-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the physician was aware of the patient¿s poor peripheral circulation and decided to proceed with impella cp implant. the patient¿s extremity became cool, and the pulse was absent. arterial study performed and confirmed decreased flow and the decision was made to explant the impella. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. added- h6 component code 925 f6/f8 should have been left blank in the initial report.
26231 | 1220648-2024-11585 | 19383411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383411&pc=OZD | 05/17/2024 | 09/19/2025 |  | 05/23/2024 |  |  | 05/17/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025478552 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | American Indian Or Alaskan Native | Malfunction | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/17/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge due durable ventricular assist device. the patient developed a hematoma at the cutdown site. heparin was withheld while the site was stabilizing. the patient remains on impella 5.5 support on the date of this report. | The user facility reported that a patient implanted with an impella 5.5 developed a hematoma at the cutdown site. heparin was withheld while the site was stabilizing. additionally, the pump had placement signal issues. support continued. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was most likely due to anticoagulation management since the hematoma was stabilized while heparin was on hold. added- d6b explant date, h6 component code 925 f6/f8-should have been left blank in the initial report. reporting contact fax number missing. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issue upon review of data logs, it is apparent that the console is the potential cause of the issue. please refer to 24-29515-2 for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
26232 | 1220648-2024-11586 | 19383537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383537&pc=OZD | 05/17/2024 | 07/10/2024 |  | 05/23/2024 |  |  | 05/18/2024 | 05/23/2024 |  |  | 06/30/2024 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235326 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2024 |  |  |  |  | The complainant reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had access site bleeding. the team held pressure, withheld heparin, redressed the site and two units of blood products were provided to the patient. additionally, the pump was repositioned and the patient had arrhythmia issues that required cardio pulmonary resuscitation and an increase in chemical support. furthermore, the patient's leg showed ischemia with no pulses. the impella was removed and replaced with the impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed, ventricular tachycardia, and limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed, ventricular tachycardia, and limb ischemia could not be determined. added-h6 component code 925 f6/f8 should have been left blank in the initial report.
26233 | 1220648-2024-11587 | 19383591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383591&pc=OZD | 05/18/2024 | 07/17/2024 |  | 05/23/2024 |  |  | 05/18/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461423 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | Difficult to Insert | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2024 |  |  |  |  | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the physician encountered delivery issues during implantation. staff was able to achieve left ventricle (lv) access with a with a pigtail and .035 wire. the impella was advanced into the artery, but the doctor was unable to make the turn down the subclavian into the aorta. the doctor removed the 5.5 and retried with a v18 wire for more stiffness. the 5.5 was re-advanced into the artery and was still unable to make the turn down to the aorta. the 5.5 was removed and lv access was lost with the wire. a stiff .035 amplatz wire was then advanced in a ¿single access¿ technique to serve as a buddy wire to help the impella down the tight curve of the subclavian. the impella was then advanced down the ascending and it stopped just before the aortic valve. after several attempts to remove slack, re-advance and pull wire tension, the doctor decided to reopen the chest to manipulate the impella past the aortic valve. once the chest was open, they quickly manipulated the impella across the valve into an excellent position. the wire was removed and the impella was started on p2. the implant resulted in significant blood loss. the patient was given four units of packed red blood cells and one liter of albumin. the patient remains on impella 5.5 support on the date of this report. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. impella 5.5 with smart assist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the access site bleed and the delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the delivery issue most likely was patient condition related as the patient was reported as having a tight curve of the subclavian that they were unable to pass the pump through. the cause of the access site bleeding- major cannot be determined since the complaint device not returned for investigation and insufficient clinical data provided. added- d6b, h6 impact code 4628 f6/f8 should have been left blank in the initial report. h6-med dev prod code 1250 removed from initial report. h6 component code 4744 removed from initial report.
26234 | 1220648-2024-11588 | 19383647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383647&pc=OZD | 05/18/2024 | 07/11/2024 |  | 05/23/2024 |  |  | 05/18/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504111 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced access site bleeding and left ventricle (lv) thrombus. the patient developed a hematoma of the right groin. hemostasis was obtained with manual pressure and transfusion of one unit of packed red blood cells. additionally, the lv thrombus was found on echocardiogram. consequently, the impella cp pump was explanted, and the patient was escalated to extracorporeal membrane oxygenation support instead of impella 5.5 due to the lv thrombus. the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ | The investigation for the reported access site bleed and thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the thrombus and access site bleed could not be determined. added- h6 component code 925 d4-corrected model number.
26235 | 1220648-2024-11589 | 19383702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383702&pc=OZD | 05/17/2024 | 07/23/2024 |  | 05/23/2024 |  |  | 05/17/2024 | 05/23/2024 |  |  | 06/30/2025 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024368804 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported a 58-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during attempted repositioning of an implanted impella cp pump, the device came back across the aortic valve. attempts to re-cross the valve were unsuccessful, and the pump subsequently kinked in the aorta. as a result of the kink, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | D.9 updated as the pump and purge cassette were returned for investigation. the investigation for the positioning issue has been completed. the returned impella cp pump was visually observed. a cannula kink near the pump housing was observed. no other abnormalities were found. the cause of the positioning issue was use related due to improper technique as the cannula trapped under the papillary with inflow sitting near the mitral valve and led to the cannula kink. b.7 revised as information was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-11589. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11589 and should not have been. h.6 code 4628 was added as it was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-11589.
26236 | 1220648-2024-11591 | 19383739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383739&pc=OZD | 05/19/2024 | 07/22/2024 |  | 05/23/2024 |  |  | 05/19/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025479502 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2024 |  |  |  |  | The user facility reported a 57-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed profuse access site bleeding. the physician removed the impella and inserted an intra-aortic balloon to resolve the issue. the patient was considered stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 revised as this information was inadvertently omitted from the initial manufacturer device report. f.6 and f.8 dates were reported on the initial manufacturer device report and should not have been. h.6 codes 4641 and 4114 were reported incorrectly on the previously submitted manufacturer device report. a new code has been added to type of investigation codes. h.6 added code 925 in accordance with updated procedures since manufacturer device report was submitted. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report in accordance with updated procedures. additional manufacturer narrative on the previously submitted manufacturer device report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded.
26237 | 1220648-2024-11592 | 19383754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383754&pc=OZD | 05/17/2024 | 07/18/2024 |  | 05/23/2024 |  |  | 05/17/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442100 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The complainant reported a 75 year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the 5.5 was pulled out of the left ventricular position in error and was subsequently replaced. the patient remains on support, and no patient harm has been reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issue most likely was use related due to improper technique as the impella pulled back back accidentally and replaced. f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing.
26238 | 1220648-2024-11593 | 19383793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383793&pc=OZD | 05/19/2024 | 06/27/2024 |  | 05/23/2024 |  |  | 05/19/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025447389 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2024 |  |  |  |  | The user facility reported a 52-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced bleeding and developed a hematoma at the groin site. the team attempted to troubleshoot with hold of pressure for an hour, perclose x2, and figure 8 stitch to resolve the issue. the patient was considered stable following the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
26239 | 1220648-2024-11594 | 19383852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383852&pc=OZD | 05/20/2024 | 06/06/2024 |  | 05/23/2024 |  |  | 05/20/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1090623 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/20/2024 |  |  |  |  | The user facility biomedical department reported that the automated impella controller (aic) had a broken purge disk. the aic was removed from service. no patient harm or involvement was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. added- d9 device return date.
26240 | 1220648-2024-11595 | 19383938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19383938&pc=OZD | 05/20/2024 | 07/09/2024 |  | 05/23/2024 |  |  | 05/20/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446043 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2024 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support for a patient admitted for ami/cgs (acute myocardial infarction¿/ cardiogenic shock) patient. the support allowed for a multi-vessel coronary pci (percutaneous coronary intervention) and was explanted after 29 hours. the pump was explanted but afterwards there were neuro changes and ct (computed tomography) imaging was performed which showed thrombosis. the team performed a thrombectomy for the cva (cerebral vascular accident) event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. added h6 4624 d4-corrected model number. f6/f8 should have been left blank in the initial report. g1 corrected reporting contact email.
26241 | 1220648-2024-11596 | 19384004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19384004&pc=OZD | 10/29/2021 | 05/23/2024 |  | 05/23/2024 |  |  | 10/29/2021 | 05/23/2024 |  |  | 09/30/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022132915 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Deliver | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/29/2021 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the physician noted resistance and fluoro revealed a kink in the sheath in the tortuous iliac. cp insertion was aborted and the doctor elected to attempt implant by lateral circumflex femoral artery (lcfa). however, the sheath would not advance through the iliac. the doctor changed to a different size sheath with struggle to insert it. it was reported that lcfa access was closed with perclose. consistent ooze was noted and fem-stop was placed with manual pressure. in addition, while on support, it was reported that the patient was having some rhythm changes and complete heart block so a temporary pacer was placed. it was also reported that patient experienced a right leg thrombus. the patient continued on support until the device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the immpella cp with smart assist catheter section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
26242 | 1220648-2024-11597 | 19385142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19385142&pc=OZD | 10/29/2021 | 05/14/2025 |  | 05/23/2024 |  |  | 10/29/2021 | 05/23/2024 |  |  | 04/30/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022074616 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/29/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was noted that the patient had a pericardial effusion.the patient was taken to the operation room for a pericardial window procedure. the fluid was drained however, the site continued to bleed. cardio thoracic surgery found that the source of the bleed was the left ventricle (lv) perforation from the impella. it is unknown if the lv perforation was repaired or how hemostasis was achieved. it was reported that surgical wound closure was successful, the device was explanted, and the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26243 | 1220648-2024-11599 | 19385599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19385599&pc=OZD | 10/29/2021 | 03/25/2025 |  | 05/23/2024 |  |  | 10/29/2021 | 05/23/2024 |  |  | 05/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083717 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/29/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hemolysis that required dialysis to resolve. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26244 | 1220648-2024-11600 | 19385674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19385674&pc=OZD | 05/19/2024 | 07/11/2024 |  | 05/23/2024 |  |  | 05/19/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025485745 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2024 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that in preparation for impella 5.5 insertion angios were taken of the artery, the md believed the artery was big enough to accommodate the 21fr cannula and believed it was 7mm. the graft was sewn on and impella wire placed in ventricle without issue. the 5.5 pump was inserted into the graft, but the inlet cage would not pass the arterial anastomosis. several attempts were made to pass through including the use of mineral oil to lube the tip, but the vessel would not accommodate. the 5.5 procedure was aborted, and a cp was placed instead. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation for the delivery issues has been completed. the impella device was returned for evaluation. upon evaluation of the device, there was no issues found with the device. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since vessel were too small to pass pump through. added- d9 device return date d4-corrected model number f6/f8 should have been left blank in the initial report. h6 impact code changed from 4627 to 4629.
26245 | 1220648-2024-11601 | 19385689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19385689&pc=OZD | 10/29/2021 | 03/25/2025 |  | 05/23/2024 |  |  | 10/29/2021 | 05/23/2024 |  |  | 08/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022125454 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/29/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, it was noted that there was a small bleed when perclose sutures were tightened so a third perclose was used. manual pressure was also applied for 30 minutes. although standard care was presented to the patient, it was decided to explant the device. the patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26246 | 1220648-2024-11598 | 19385690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19385690&pc=OZD | 05/18/2024 | 07/12/2024 |  | 05/23/2024 |  |  | 05/18/2024 | 05/23/2024 |  |  | 10/31/2025 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024339556 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2024 |  |  |  |  | The complainant reported a 61 year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported that the patient developed mottling and loss of pulses in the right leg. the cp was explanted to resolve the limb ischemia, and hemostasis was achieves with manual compression. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation and data logs were not made available. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. f6 and f8 should have been blank h6 health effect-impact code should be 4629 only. 4627 and 4625 should be removed from the initial report.
26247 | 1220648-2024-11604 | 19385809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19385809&pc=OZD | 10/29/2021 | 03/25/2025 |  | 05/23/2024 |  |  | 10/29/2021 | 05/23/2024 |  |  | 02/28/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022052916 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/29/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, oozing was noted from the groin access site. oozing was addressed by staff; however, specifics of the medical or surgical intervention are unknown. the patient continued on support until successful wean and impella removal. hemostasis was achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26248 | 1220648-2024-11605 | 19385834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19385834&pc=OZD | 05/15/2024 | 06/24/2024 |  | 05/23/2024 |  |  | 05/15/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025481908 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an attempted impella 5.5 implant and the implant was aborted due to the patient's anatomy. no known patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation into delivery issues has been completed since the original report was submitted. the product was not returned for investigation. the root cause of the access site bleeding was most likely patient condition since the implant was aborted due to patient's anatomy.
26249 | 1220648-2024-11602 | 19385909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19385909&pc=OZD | 05/20/2024 | 07/18/2024 |  | 05/23/2024 |  |  | 05/20/2024 | 05/23/2024 |  |  |  |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443057 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/20/2024 |  |  |  |  | The complainant reported a 49 year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported the 5.5 was showing purge blocked alarms. the alarms did not prompt any pump replacement or intervention. the patient remains on support despite the alarms. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure (hpp) has been completed. no product was returned for investigation. data logs show that 260 hrs into support there is a 13min rise in pp and drop in purge flow with hpp and purge system blocked alarms. the purge flow tick stalls. flow, placement signal (ps), and motor current (mc) are pulsatile throughout support. during this time of hpp, there are a few instances of the pressure purge recovering and then immediately returning to hpp; these periods do not correlate with bag or cassette changes. the root cause of the high purge pressure was most likely a kinked purge line due to excessive force by the user. the product was not returned therefore the location of the kink is unknown. f6/f8 should have been left blank in the initial report. g4-added pma number.
26250 | 1220648-2024-11606 | 19387969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19387969&pc=OZD | 05/15/2024 | 06/24/2024 |  | 05/23/2024 |  |  | 05/15/2024 | 05/23/2024 |  |  | 09/30/2025 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025444670 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a cp for support of a patient admitted in with a nstemi (non-st segment elevation myocardial infarction). the patient had the cp placed after code and cpr. the pump was placed but after 2 days there was thrombus noted, new clot, within the lv (left ventricle). this prompted the team to plan wean/explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombus issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
26251 | 1220648-2024-11607 | 19388630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19388630&pc=OZD | 06/28/2022 | 05/23/2024 |  | 05/23/2024 |  |  | 06/28/2022 | 05/23/2024 |  |  | 10/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022139834 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/28/2022 |  |  |  |  | The complainant reported that a73 year old female patient was implanted with an impella cp for mechanical circulatory support. while on support there was access site bleeding, noted by the hematoma. this was treated through a blood transfusion. the patient was successfully weaned from the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26252 | 1220648-2024-11608 | 19388689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19388689&pc=OZD | 09/27/2022 | 05/23/2025 |  | 05/23/2024 |  |  | 09/27/2022 | 05/23/2024 |  |  | 05/31/2024 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229851 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2022 |  |  |  |  | The complainant reported the 71 year old female patient was implanted with an impella cp for mechanical circulatory support as percutaneous coronary intervention. while on support, the pump experienced suction that was resolved by weaning from p3 to p2. the patient was successfully weaned and the pump explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26253 | 1220648-2024-11609 | 19388740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19388740&pc=OZD | 09/27/2022 | 05/23/2024 |  | 05/23/2024 |  |  | 09/27/2022 | 05/23/2024 |  |  | 08/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022119912 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/27/2022 |  |  |  |  | The complainant reported that a 72 year old male patient was implanted with an impella cp for mechanical circulatory support. while on support, it was reported that there was bleeding including a hematoma and hemolysis. there was no blood transfusion required. the patient was successfully weaned from the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ access site bleeding section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26254 | 1220648-2024-11610 | 19388826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19388826&pc=OZD | 10/25/2021 | 05/20/2025 |  | 05/23/2024 |  |  | 10/25/2021 | 05/23/2024 |  |  | 08/31/2023 |  | 05/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022125206 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was noted that the patient's urine was previously yellow now it's bloody/dark red. the staff was concerned that the impella was causing hemolysis. the impella device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26255 | 1220648-2024-11493 | 19369546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19369546&pc=OZD | 11/25/2021 | 03/27/2025 |  | 05/22/2024 |  |  | 11/25/2021 | 05/22/2024 |  |  | 09/30/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022136509 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/25/2021 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced blocked circulation down the impella leg. an external femoral to femoral bypass was attempted, but was unsuccessful. the medical team decided to explant the impella due to the limb ischemia as the patient was hemodynamically stable. the impella cp was weaned and successfully explanted from the patient. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26256 | 1220648-2024-11495 | 19369923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19369923&pc=OZD | 11/25/2021 | 03/25/2025 |  | 05/22/2024 |  |  | 11/25/2021 | 05/22/2024 |  |  | 07/31/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114361 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 11/25/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella access site after leaving the cardiac catheterization lab (ccl) and brought to their room. the patient was brought back to the ccl for unspecified interventions to achieve hemostasis. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26257 | 1220648-2024-08086 | 19369940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19369940&pc=OZD | 02/28/2024 | 05/24/2024 |  | 05/22/2024 |  |  | 02/28/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024436403 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient on impella support was brought to the operation room for a mitral valve replacement (mvr) and tricuspid ring. the impella was placed in surgical mode once fully on bypass. the surgeon evaluated motor housing position for cross clamp. the surgeon stated it was in the way and wanted to advance the pump some. the impella device was attempted to be advanced multiple times over several minutes but the impella was ¿stuck¿¿. the surgeon was able to compress the blue button, but the device would not move. the surgeon tried to move it in all directions with no luck. after battling with it for several minutes, the mechanism finally broke loose. the device was advanced. the surgeon lifted the heart. bleeding was noted in the chest cavity. left ventricle (lv) perforation was found. the impella was pulled out of body. lv was repaired successfully with a patch. mvr and tricuspid ring were successful. there was significant damage to right axillary/subclavian which the surgeon attempted to repair but was unable. the bypass was performed but was still unable to control the bleeding. the patient expired. | [5 empty] | The investigation for the reported ventricular perforation and the positioning issues has been completed. impella 5.5 pump was returned for investigation. visually inspected and no cannula, catheter kinks or other damages observed. the cath anchor was tested by pressing the blue button and was able to slide the catheter and move position without any resistance or abnormalities. no other physical abnormalities were found. data logs were reviewed, and no improper position alarms observed. no other issues were found in the logs. as per clinical the device catheter would not move further in either direction after compressing blue button. later doctor was battling with it for several minutes and attempts, the mechanism finally broke loose and device advanced further which likely led to lv perforation. the cause of the ventricle perforation issue was use related per clinical and return product investigation. the cause of the positioning issue was due to cath anchor use issue per clinical and return product investigation. the instructions for use referenced in the initial report is from ifu for impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system in section: potential adverse events (united states), which now includes a statement "cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
26258 | 1220648-2024-11496 | 19369960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19369960&pc=OZD | 04/24/2024 | 06/13/2024 |  | 05/22/2024 |  |  | 04/24/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025450025 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while priming the pump, the device would not zero correctly. pressures showed +/-6/6. staff attempted troubleshooting, but they were unsuccessful. the impella 5.5 was not implanted. a new pump was used successfully. | [5 empty] | The investigation for the reported priming issue has been completed. the impella pump was returned for evaluation. the root cause of the optical signal issue was not determined since product was not returned, data logs were not returned and insufficient clinical information provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26259 | 1220648-2024-11497 | 19370084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19370084&pc=OZD | 11/30/2021 | 03/25/2025 |  | 05/22/2024 |  |  | 11/30/2021 | 05/22/2024 |  |  | 07/31/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022116782 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/30/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella access site. it was unspecified as to how the bleeding was managed post operation, but the patient was successfully weaned and explanted. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26260 | 1220648-2024-11501 | 19370531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19370531&pc=OZD | 01/18/2022 | 04/02/2025 |  | 05/22/2024 |  |  | 01/18/2022 | 05/22/2024 |  |  | 06/30/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022107849 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2022 |  |  |  |  | The complainant reported a patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella was inadvertently pulled out and the impella was removed and reimplanted into position. patient's condition was stable post procedure. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26261 | 1220648-2024-11502 | 19370670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19370670&pc=OZD | 01/22/2022 | 04/02/2025 |  | 05/22/2024 |  |  | 01/22/2022 | 05/22/2024 |  |  | 10/31/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156946 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2022 |  |  |  |  | The complainant reported that a patient was presented for pre-operative coronary artery bypass graft (cabg). the impella was placed for support during cabg. it was observed that the patient lost bilateral pulses to the legs. the recommendation was made to utilize the rooke vascular boot to warm the lower extremities and redistribute pressure along the calf. it was noted that the right food was warm to the touch, but no doppler pulses were identified. the rooke boots were placed on the patient since the previous day and the patient¿s leg and foot are now warm to the touch. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26262 | 1220648-2024-11503 | 19370898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19370898&pc=OZD | 05/07/2024 | 06/13/2024 |  | 05/22/2024 |  |  | 05/07/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025457165 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Life Threatening | [23 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 05/07/2024 |  |  |  |  | The complainant reported a 64-year-old male patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed a thrombus ¿ a large clot in the descending aorta. the cp was removed and replaced. the cause of the clot remains unknown, as the patient was on anticoagulants. the patient survived the event. | [5 empty] | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. g1-corrected mfr site name and address. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
26263 | 1220648-2024-11504 | 19370933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19370933&pc=OZD | 05/09/2024 | 07/12/2024 |  | 05/22/2024 |  |  | 05/09/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025453632 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The complainant reported a 78-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support for st elevated myocardial infarction. the complainant reported the patient developed access site bleeding, which prompted the pump getting explanted after two hours. medical staff could not control the bleeding. the patient survived the event. | [5 empty] | The investigation into the access site bleeding - major issue has been completed. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details was provided and no product was returned for analysis. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.5 the patient was a male and not a female and was reported incorrectly on manufacturer device report number 1220648-2024-11504. b.7 revised as this information was inadvertently omitted from manufacturer device report number 1220648-2024-11504. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11504 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since manufacturer device report number 1220648-2024-11504 was submitted. h.6 revised component codes from manufacturer device report number 1220648-2024-11504 in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from manufacturer device report number 1220648-2024-11504 in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26264 | 1220648-2024-11505 | 19371014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371014&pc=OZD | 05/09/2024 | 07/05/2024 |  | 05/22/2024 |  |  | 05/09/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025457309 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The complainant reported a patient (race unknown) noted as high risk percutaneous coronary intervention (pci) was to be implanted with an impella cp device for mechanical circulatory support. the impella cp failed delivery. pci was performed via single access technique with intra-aortic balloon pump support. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because multiple attempts were done to implant the pump and were not successful. f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
26265 | 1220648-2024-11506 | 19371029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371029&pc=OZD | 11/10/2021 | 05/22/2024 |  | 05/22/2024 |  |  | 11/10/2021 | 05/22/2024 |  |  | 06/30/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022104488 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/10/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced oozing at the access site that was ongoing despite re-dressing, angle prop, and forward tension. in addition, the patient had oozing from other sites as well, so systemic heparin was put on hold. the physician decided to exchange the impella cp for an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26266 | 1220648-2024-11507 | 19371101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371101&pc=OZD | 05/09/2024 | 07/09/2024 |  | 05/22/2024 |  |  | 05/09/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025458587 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported there was an inability to deliver an impella 5.5 device for the support of a patient due to the patient's anatomy. consequently, the procedure was aborted. there was no report of patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related; the implant was aborted due to the patient's anatomy as per the clinical description. d4-corrected model number added- d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
26267 | 1220648-2024-11508 | 19371128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371128&pc=OZD | 11/13/2021 | 04/02/2025 |  | 05/22/2024 |  |  | 11/13/2021 | 05/22/2024 |  |  | 08/31/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022132989 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced oozing at the groin access site, which continued after applying manual pressure, but was subsequently treated with one unit of packed red blood cells and femstop. the patient then remained on impella support. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26268 | 1220648-2024-11509 | 19371281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371281&pc=OZD | 05/09/2024 | 07/11/2024 |  | 05/22/2024 |  |  | 05/09/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025474736 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The complainant reported a 61-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed oozing and a hematoma at the access site. medical staff held manual pressure, adjusted the sheath angle, and tightened sutures. medical staff then decided to explant the cp and provided two units of blood following the explant. the patient survived the event and has been reported as stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. d4-corrected model numberf6/f8-should have been left blank in the initial report. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
26269 | 1220648-2024-11510 | 19371297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371297&pc=OZD | 11/13/2021 | 04/02/2025 |  | 05/22/2024 |  |  | 11/13/2021 | 05/22/2024 |  |  | 08/31/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121525 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device experienced a placement signal alarm, so an echocardiogram was performed to confirm placement, and the placement signal was disabled. the device continued to show "position in aorta" alarms, despite the device having a pulsatile motor current. however, the patient also experienced multiple bouts of ventricular tachycardia requiring advanced cardiac life support protocol. hospital staff noted that the patient had a small left ventricle cavity and that this may have caused the pigtail catheter to have contact against the wall of the ventricle. the patient survived these procedures. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26270 | 1220648-2024-11511 | 19371339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371339&pc=OZD | 05/13/2024 | 06/27/2024 |  | 05/22/2024 |  |  | 05/13/2024 | 05/22/2024 |  |  | 07/31/2025 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024392214 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and lost pulses in the right distal limb. the impella cp was placed via the right femoral artery and during support, pulses became absent. the leg was cold and after 17 hours of support the impella cp pump was removed at bedside, and the patient was transported to the operating room for venous-arterial extra-corporeal membrane oxygenation. the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation into the limb ischemia ¿ reversible issue been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
26271 | 1220648-2024-11512 | 19371362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371362&pc=OZD | 05/12/2024 | 07/09/2024 |  | 05/22/2024 |  |  | 05/12/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025451694 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2024 |  |  |  |  | The user facility reported a 76-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during support the patient developed hemolysis and two units of packed red blood cells were provided. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section hemolysis: hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported hemolysis has been completed. the impella device was not returned for evaluation. data logs were reviewed which showed pump run without issue during support. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. f6/f8 should have been left blank in the initial report. g1 reporting contact fax number missing g1 reporting contact address line 1 should have been blank. corrected manufacturing site address line 1
26272 | 1220648-2024-11513 | 19371420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371420&pc=OZD | 05/15/2024 | 07/26/2024 |  | 05/22/2024 |  |  | 05/15/2024 | 05/22/2024 |  |  | 09/30/2025 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024393738 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The complainant reported a patient (race unknown) with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient experienced oozing from the graft. the event was resolved with anticoagulation management and four units of packed red blood cells. the patient was stable. the patient remains on impella 5.5 support on the date of this report. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was determined to be anti coagulation management and were able to achieve hemostasis. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26273 | 1220648-2024-11514 | 19371433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371433&pc=OZD | 05/14/2024 | 06/27/2024 |  | 05/22/2024 |  |  | 05/14/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST S2 SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2025478409 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for ongoing care escalation from an impella cp device. the 5.5 pump developed a hematoma at the access site. the site was treated by pressure, sandbag and reversal of heparin with protamine. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the access site bleeding ¿ major issue been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
26274 | 1220648-2024-11515 | 19371451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371451&pc=OZD | 11/20/2021 | 05/22/2024 |  | 05/22/2024 |  |  | 11/20/2021 | 05/22/2024 |  |  | 11/30/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022137352 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/20/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed oozing at the impella access site. systemic heparin drip was decreased as a result. the patient continued on support until successful weaning of the device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26275 | 1220648-2024-11517 | 19371501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371501&pc=OZD | 05/15/2024 | 06/26/2024 |  | 05/22/2024 |  |  | 05/15/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1810621 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The complainant reported a 36-year-old female patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported the auotmated impella controller (aic) had a broken part on the purge disc holder, so medical staff replaced the aic. the patient survived the event. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the logs are irrelevant to this complaint as the clinical staff only reported that the purge disc retainer was broken. the console was returned. upon examining the automated impella controller for any visible defects, it was evident that the blue purge disc retainer was broken. the cause of the issue was a broken purge retainer. d.2 common device name was revised from the initial submission in accordance with updated procedures. d.4 unique identifier (udi) # was revised as it was previously entered incorrectly. d.6a and d.6b implant and explant dates were reported in error on the previously submitted report. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
26276 | 1220648-2024-11516 | 19371502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371502&pc=OZD | 05/15/2024 | 06/20/2024 |  | 05/22/2024 |  |  | 05/15/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485878 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The complainant reported a 36-year-old female patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed limb ischemia ¿ the limbs were cold and pulses could not be located on the doppler. medical staff ordered the patient a bear hugger. the patient survived the event. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use referenced in the initial report is for the impella cp with smartassist in section: potential adverse events (united states). | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
26277 | 1220648-2024-11518 | 19371561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371561&pc=OZD | 05/15/2024 | 06/27/2024 |  | 05/22/2024 |  |  | 05/15/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448976 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had limb ischemia. when a peripheral angiogram was performed on the left leg, it showed total occlusion of the left femoral artery. the impella was removed. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿.
26278 | 1220648-2024-11519 | 19371572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371572&pc=OZD | 05/16/2024 | 07/12/2024 |  | 05/22/2024 |  |  | 05/16/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025468778 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The complainant reported a 39-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed limb ischemia despite assistance from a distal reperfusion sheath. medical staff explanted the pump, and the patient survived the event. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. d4-corrected model number f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing h.6 codes 4629 and 4625 were reported incorrectly on manufacturer device report 1220648-2024-11519. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
26279 | 1220648-2024-11520 | 19371580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371580&pc=OZD | 05/15/2024 | 07/10/2024 |  | 05/22/2024 |  |  | 05/16/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025474734 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | PHYSICIAN ASSISTANT | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The user facility reported a 53-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced oozing from their groin site and pericardial drain site during impella cp support. manual pressure was applied to the site and four units of packed red blood cells and two units of fresh frozen plasma were transfused to treat the bleed. support with the implanted pump remained ongoing following the intervention. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. d4-corrected model number f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26280 | 1220648-2024-11521 | 19371684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371684&pc=OZD | 05/16/2024 | 07/17/2024 |  | 05/22/2024 |  |  | 05/16/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025468623 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The user facility reported a 39-year-old female with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient developed hemolysis. after multiple failed attempts to resolve the hemolysis, the decision was made to explant the pump. an intra-aortic balloon pump was subsequently implanted for ongoing support. | [5 empty] | The investigation for the reported hemolysis has been completed. pump and data logs were not provided for investigation. according to clinical reports, hemolysis persisted even after repositioning the pump, which was found deep. fluid administration did not change the hemolysis. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. added-h6 impact code 4642, 4628 f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section hemolysis: hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
26281 | 1220648-2024-11522 | 19371743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371743&pc=OZD | 05/17/2024 | 07/30/2024 |  | 05/22/2024 |  |  | 05/17/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1116498 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) displayed a controller error alarm during impella 5.5 support. at the time of the noted alarm, the placement signal was lost. within a few seconds, the alarm resolved, and the placement signal re-appeared. there was no patient harm and support with the console continued. | [5 empty] | D.9 updated as the console was returned for investigation. the investigation into the automated impella controller error message has been completed. the logs confirmed the reported failure mode given. the controller error alarm triggered during the clinical procedure. the real time logs confirmed that all signals were received from the optical bench were represented as -(b)(6) values due to the error. the console was returned for evaluation. the controller error and loss of placement signal was due to an optical bench firmware communication protocol anomaly resulting in misalignment of data communication packets received by the embedded personal computer (epc). the aic software operated as designed. b.7 information was added as it was inadvertently omitted from the previously submitted manufacturer device report 1220648-2024-11522. d.4 revised model number from the initial submission of manufacturer device report 1220648-2024-11522 in accordance with updated procedures. d.6a date was reported on manufacturer device report number 1220648-2024-11522 and should not have been. h.3 "if the device was not evaluated, select from the approved fda codes or select other and enter text in h10". the incorrect selection was in the initial manufacturer device report 1220648-2024-11522. it should have been other (code unspecified, describe in h10). | The automated impella controller was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26282 | 1220648-2024-11523 | 19371745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371745&pc=OZD | 05/17/2024 | 07/16/2024 |  | 05/22/2024 |  |  | 05/17/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438729 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp being placed with single access with the 7fr companion sheath. the 7 failed and the team instead placed a 8fr. there was an access site bleed. the bleed was treated by manual hold, balloon tamponade. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the cause of the bleeding was most likely high patient act values (300 seconds) above the 160-180 recommended range. the instructions for use referenced in the initial report is for the impella cp with smart assist system. d4-corrected model number f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26283 | 1220648-2024-11494 | 19371756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371756&pc=OZD | 11/25/2021 | 04/02/2025 |  | 05/22/2024 |  |  | 11/25/2021 | 05/22/2024 |  |  | 08/31/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022135081 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/25/2021 |  |  |  |  | The complainant reported a patient in atrial fibrillation with rapid ventricular response was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced a large hematoma at the impella access site. systemic heparin was withheld and the patient received blood products to maintain hemoglobin, hematocrit, and platelets. hemostasis was achieved and heparin was restarted. the patient remained on impella support until they were successfully weaned and explanted. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26284 | 1220648-2024-11524 | 19371881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371881&pc=OZD | 05/17/2024 | 06/06/2024 |  | 05/22/2024 |  |  | 05/17/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1274262 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Component Incompatible | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had multiple pc replaced due to purge disc issues. the aic was replaced without harm or injury. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. upon examining aic for any visible defects, it was evident that the blue purge disc retainer was broken. the issue of broken purge disk retainer was confirmed. the cause of the issue was a broken purge retainer. in accordance with updated procedures, section d2 has been revised from the initial submission. | The impella device was received from the customer on (b)(6) 2024: data review: logs are irrelevant to this complaint as the clinical staff only reported that the purge disc retainer was broken. device analysis: upon examining aic for any visible defects, it was evident that the blue purge disc retainer was broken. the complaint console serial number was updated from "(b)(6)" to "(b)(6)" based on the unit returned for investigation. per the returned device investigation, the issue of broken purge disk retainer was confirmed in "(b)(6)", so it was evident that the intake has captured the incorrect complaint console number in the smartsolve system and also the console number "(b)(6)" reported in the event description is not related to the issue.
26285 | 1220648-2024-11525 | 19371931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371931&pc=OZD | 05/16/2024 | 07/23/2024 |  | 05/22/2024 |  |  | 05/16/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479234 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The complainant reported a 70-year-old male patient with elective placement, a high risk of coronary artery bypass graft, and valve, was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported failing to deliver the 5.5, so an impella cp was used instead. the patient survived the event. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since it was difficult advancing the impella 5.5. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d.4 revised model number from the initial submission of manufacturer device report 1220648-2024-11525 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11525 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-11525 was submitted. g.4 revised pma/510(k) number as it was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-11525. h.6 codes 4114, 463 and 3038 were reported incorrectly on the previously submitted manufacturer device report number 1220648-2024-11525. a new code has been added to type of investigation codes. additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. guidance.¿
26286 | 1220648-2024-11526 | 19371969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19371969&pc=OZD | 04/22/2024 | 05/22/2024 |  | 05/22/2024 |  |  | 04/22/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1132984 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 04/22/2024 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the automated impella controller (aic) generated purge disc not detected alarm. the console was replaced to resolve the issue. no patient harm was reported. | [5 empty] | The impella device was received from the customer and an investigation regarding the returned device has been completed. the logs were reviewed and the issues with the automated impella controller (aic) not being able to detect the purge disc was confirmed. the purge disc flag was found to be broken (missing blue disc retainer). the root cause of the issue was a broken purge disc flag.
26287 | 1220648-2024-11528 | 19372070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19372070&pc=PYX | 05/19/2024 | 07/17/2024 |  | 05/22/2024 |  |  | 05/19/2024 | 05/22/2024 |  |  | 05/31/2024 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2023213697 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2024 |  |  |  |  | The user facility reported an 83-year-old male with cardiomyopathy was implanted with an impella rp device for mechanical circulatory support. it was reported that an implanted impella rp pump triggered a purge pressure low alarm during patient support. the cassette was replaced in an attempt to troubleshoot the issue, but the failure remained. upon further inspection, a crack and leak were noted in the yellow luer. a purge bypass was completed to resolve the alarm. there was no patient harm and support with the implanted pump remained ongoing. | [5 empty] | The investigation for purge leak ¿ pump has been completed. the impella device was not returned for evaluation. low purge pressure reported and crack with leak observed. no product was returned. the data logs confirm 'purge pressure low' alarms. there was a downward trending of purge pressure before low purge pressure alarms were triggered and there were several 'purge system open' alarms during the purge cassette changes and 'de-airing' troubleshoot efforts. the purge pressure recovered after the purge reservoir and filter bypass, aligning with the clinical information, and there were no other purge related alarms. the root cause of the purge leak - pump was due to a damaged sidearm but the cause of the sidearm damage was not determined. d6b added f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26288 | 1220648-2024-11527 | 19372077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19372077&pc=OZD | 05/18/2024 | 07/19/2024 |  | 05/22/2024 |  |  | 05/18/2024 | 05/22/2024 |  |  | 07/31/2025 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432555 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2024 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to decompensating. after the implant, the impella was turned on auto, but the flows dropped from 3.1 to 2.1. the motor current flattened and suction was present. suction continued and flows dropped below 1. the pump was removed and replaced. | [5 empty] | The investigation for the low pump flow has been completed. data logs were reviewed which confirmed the failure mode and clinical narrative. logs showed after pump started, motor spiked followed low flow that triggered auto suction response and suction alarms. motor current spike greater than 1500 ma was found. logs also showed at the end of the case, high motor current #13 pump stop alarm. pump stopped and could not restart after multiple attempts. high flows of 3 l/min at lower p-level before pump stop was seen. pump was returned for analysis. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue under bench test. the biomaterial was likely expelled during support when pump was removed. based on clinical description and data review, the root cause of low flow was most likely biomaterial ingestion. the root cause of pump stop was most likely biomaterial ingestion. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high risk cpi in the following sections: section: using the automated impella controller with the impella catheter section: using the automated impella controller with the impella rp with smartassist system catheter. added- d9 device return date d4-corrected model number. f6/f8 should have been left blank in the initial report. h5 changed to yes. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
26289 | 1220648-2024-11529 | 19372198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19372198&pc=OZD | 05/15/2024 | 10/24/2025 |  | 05/22/2024 |  |  | 05/15/2024 | 05/22/2024 |  |  | 07/31/2025 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443826 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/15/2024 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp placed to support a patient admitted in with multivessel cad (coronary artery disease) and in need of a hrpci (high risk percutaneous coronary intervention) . while awaiting the pci (percutaneous coronary intervention) the team had to code the patient due to vt (ventricular tachycardia) arrest. the cp was placed but there was a persistent access site bleed that needed intervention in form of new suturing, manual pressure, and femostop. no blood products needed/given. the pump also had loss of ps but neither issue prompted pump explant. | [5 empty] | The investigation for the reported access site bleed has been completed. the pump was not returned for investigation. according to the clinical notes, the patient was awake and moving his limbs. access site bleeding was reported at this time. the access site was held with manual pressure until the doctor re-sutured and placed a fem-stop after repositioning the impella. arterial blood pressure improved after repositioned the device. the cause of the positioning issue was most likely patient movement, based on given clinical details. the cause of the access site bleeding issue was patient movement, based on clinical details. added- h6 component code 925, d7a is no d4-corrected model number f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | D6b "if explanted, give date" corrected.
26290 | 1220648-2024-11530 | 19372475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19372475&pc=OZD | 05/10/2024 | 07/10/2024 |  | 05/22/2024 |  |  | 05/10/2024 | 05/22/2024 |  |  | 07/31/2025 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024379158 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during support, there was a high-pressure purge alarm that was remedied with the use of tissue plasminogen activator (tpa). no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. pump was not returned for investigation. data logs show a gradual rise in purge pressure over two hours with purge pressure high alarms. this trend continued for 11 hours before returning to normal specification limits. the doctor-initiated tissue plasminogen activator (tpa) after observing the high purge pressure. the cause of the high purge pressure issue was most likely biomaterial. f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing.
26291 | 1220648-2024-11531 | 19372520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19372520&pc=OZD | 10/24/2021 | 05/22/2024 |  | 05/22/2024 |  |  | 10/24/2021 | 05/22/2024 |  |  | 07/31/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022118970 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/24/2021 |  |  |  |  | The user facility reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that following the exchange of the peel-away sheath the patient began having significant bleeding at the insertion site. adequate control of the bleeding was not achieved so the decision was made to remove the pump as patient was hemodynamically stable and not requiring pressors. the physician removed the pump and applied manual pressure. one unit of red blood cells was administered, and a fem-stop was applied. the patient's pulses were intact. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states).¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26292 | 1220648-2024-11532 | 19372693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19372693&pc=OZD | 05/13/2024 | 07/18/2024 |  | 05/22/2024 |  |  | 05/13/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484978 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella cp pump became clotted during patient support and was unable to properly flow. clots were subsequently seen in the tubing, and the decision was made to explant the device. there was no patient harm resulting from the noted issue. | [5 empty] | The investigation for the low pump flow has been completed. the impella device was not returned for investigation. data logs were reviewed finding no motor current spikes. many suction alarms occurred. p-level was unable to rise above p-6 and at last 24 hours dropped more between p-4 and p-2 and pump was unable to flow properly. the cause of the low pump flow issue cannot determined since the complaint device not available for investigation and insufficient clinical data provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26293 | 1220648-2024-11533 | 19372739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19372739&pc=OZD | 10/25/2021 | 05/22/2024 |  | 05/22/2024 |  |  | 10/25/2021 | 05/22/2024 |  |  | 06/30/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108584 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/25/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient developed oozing at the impella access site. the staff put a sandbag over the impella site and changed the dressing. the nurse was educated to keep the impella at an angle of insertion and not to place sandbag over the site directly. oozing minimalized, however the right fem swan sight was still oozing significantly. the patient received 1 unit of packed red blood cells and the heparin was cut in half. additionally, the patient's urine was pink tinged. it was reported that the patient survived the procedure and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26294 | 1220648-2024-11534 | 19372805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19372805&pc=OZD | 10/25/2021 | 05/22/2024 |  | 05/22/2024 |  |  | 10/25/2021 | 05/22/2024 |  |  | 06/30/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022107189 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/25/2021 |  |  |  |  | The user facility reported a 61-year-old patient placed on impella cp support for st-segment elevation myocardial infarction (stemi) and cardiogenic shock (cgs). no pulses were detected in impella leg. left foot showed signs of mottling. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26295 | 1220648-2024-11535 | 19372893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19372893&pc=OZD | 10/25/2021 | 05/22/2024 |  | 05/22/2024 |  |  | 10/25/2021 | 05/22/2024 |  |  | 07/31/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121687 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/25/2021 |  |  |  |  | The user facility reported a 79-year-old patient placed on impella cp support. it was reported that the patient was in rapid atrial fibrillation since the impella was implanted and experienced hemolysis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26296 | 1220648-2024-10582 | 19372970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19372970&pc=OZD | 04/23/2024 | 06/12/2024 |  | 05/22/2024 |  |  | 04/23/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024436420 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | The user facility reported a 54-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the impella 5.5 experienced an increase in high pressure, followed by a complete stoppage. after multiple unsuccessful attempts to restart the pump, the medical staff decided to explant the impella 5.5. the patient survived the event. | [5 empty] | The investigation into the high purge pressure issue has been completed since the original report was filed. the pump was returned for investigation. a kink was observed at the kink protector on the catheter side. the pump passed optical light testing but failed electrical testing on all three motor lines. the cause of the pump stop issue was an open electrical line due to excessive force. | The impella device was received from the customer and an evaluation of the device is pending. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
26297 | 1220648-2024-11536 | 19373069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19373069&pc=OZD | 11/10/2021 | 03/11/2025 |  | 05/22/2024 |  |  | 11/10/2021 | 05/22/2024 |  |  | 03/31/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022067872 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient was had red urine. the device was determined to be mispositioned at that time and was repositioned under echocardiogram, and the p level was reduced to p6. due to the hemolysis, the decision was made to explant the device. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
26298 | 1220648-2024-11499 | 19373080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19373080&pc=OZD | 01/13/2022 | 04/02/2025 |  | 05/22/2024 |  |  | 01/13/2022 | 05/22/2024 |  |  | 10/31/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151516 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | Patient-Device Incompatibility | [5 empty] | Unspecified Vascular Problem | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/13/2022 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that when the impella cp was explanted, 5 perclose devices failed to close the site due to the patient's body habitus. it was observed the perclose device unable to reach vessel tract, and the medical team placed a sheath as a bridge for surgical cutdown and closure. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26299 | 1220648-2024-11500 | 19373082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19373082&pc=OZD | 01/18/2022 | 04/02/2025 |  | 05/22/2024 |  |  | 01/18/2022 | 05/22/2024 |  |  | 11/30/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022150928 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/18/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was attempted to be implanted with an impella cp device for mechanical circulatory support and the impella cp implant was aborted related to ventricular thrombus. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
26300 | 1220648-2024-11498 | 19373242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19373242&pc=OZD | 01/13/2022 | 03/31/2025 |  | 05/22/2024 |  |  | 01/13/2022 | 05/22/2024 |  |  | 10/31/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149156 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/13/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. during delivery of the impella, the 14fr oscor sheath was inserted into the left femoral artery (lfa). lfa was occluded after the 14fr sheath was inserted. 14fr sheath was removed and the repositioning sheath was advanced into the lfa. patient's condition was stable post procedure. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
26301 | 1220648-2024-11537 | 19373303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19373303&pc=OZD | 04/22/2024 | 06/12/2024 |  | 05/22/2024 |  |  | 04/22/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456609 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/22/2024 |  |  |  |  | The user facility reported an 89-year-old male with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the patient developed access site bleeding, which the medical staff controlled with suture tightening, placing a femoral stop, and holding pressure. the patient eventually expired. there is not enough available information to associate the patient's event outcome with use of the impella device. | [5 empty] | The investigation into the access site bleeding - major and the product damage (sterile sleeve damage) issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was most likely use related, the bleeding was controlled by tightening the sutures, placing a fem stop, and holding pressure as per clinical description. the root cause of the sterile sleeve damage was unable to be determined as no product was returned and insufficient clinical details were provided. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26302 | 1220648-2024-11538 | 19373757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19373757&pc=OZD | 05/16/2024 | 07/22/2024 |  | 05/22/2024 |  |  | 05/16/2024 | 05/22/2024 |  |  | 07/31/2025 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024375529 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2024 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after 2 hours of support the pump was explanted due to high purge pressures. the pump was explanted and replaced without any harm or injury from interruption. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | D.9 updated as the pump was returned for investigation. the investigation for the high purge pressure has been completed. the pump was returned for investigation. no physical abnormalities were observed on the returned product. no biomaterial was observed in the purge gap and impeller. the pump was soaked for 15 minutes before purge testing using a test cassette. high purge pressure was not reproduced during the test. the pump ran within specification during test. data logs showed a gradual increase in purge pressure after implantation, with several high purge pressure alarms. purge pressure did not change after the cassette change procedure. no trend was noted in motor current and placement signal. saturated pump flow was noted while the pump was running on p-level 5. heparin was not administered in the purge fluid. as per clinical details, the pump was explanted after continuous high purge pressure alarms even after troubleshooting. the cause of the high purge pressure issue was most likely biomaterial. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ f.6 and f.8 dates were reported on the initial manufacturer device report number 1220648-2024-11538 and should not have been. h.10 additional manufacturer narrative instructions for use were revised from the initial manufacturer device report number 1220648-2024-11538 in accordance with updated procedures.
26303 | 1220648-2024-11539 | 19373815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19373815&pc=OZD | 05/16/2024 | 07/16/2024 |  | 05/22/2024 |  |  | 05/17/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1489955 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that automated impella controller (aic) failed to initiate the priming process during training. there was no patient involvement. | [5 empty] | The investigation for the console priming issue has been completed. data logs were reviewed which showed that purge disc couldn't be detected was confirmed. the console was returned for evaluation. the purge disc flag was found to be broken (missing at blue retainer). the root cause for the priming issue was a broken purge flag. added- d9 device return date d4-corrected model number | The automated impella controller has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
26304 | 1220648-2024-11540 | 19373901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19373901&pc=OZD | 05/16/2024 | 08/10/2024 |  | 05/22/2024 |  |  | 05/17/2024 | 05/22/2024 |  |  | 10/31/2025 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024409995 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted via direct access with an impella 5.5 for mechanical circulatory support. during attempted removal, the motor and cannula of the impella 5.5 pump broke off. the pump was removed from the aortic valve, but a subsequent attempt the following day was only successful in removing the motor. the pump was completely contained within the graft. attempts to remove the graft with pump were halted to allow the patient to rest. the team will then remove the patient's graft and imbedded pump. | [5 empty] | The investigation for product damage (detached inflow/outflow - great vessel) has been completed. the cause of the product damage (detached inflow/outflow - great vessel) issue was not determined since the product was not returned. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11540 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-11540 was submitted. h.6 code 4114 was reported incorrectly on the previously submitted manufacturer device report number 1220648-2024-11540. a new code has been added to type of investigation codes.h.10 additional manufacturer narrative on the previously submitted manufacturer device report number 1220648-2024-11540 stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: warnings, cautions, and precautions. ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ impella 5.5 with smart assist for use during cardiogenic shock. section: warnings & cautions: warnings ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
26305 | 1220648-2024-11541 | 19373951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19373951&pc=OZD | 05/17/2024 | 07/16/2024 |  | 05/22/2024 |  |  | 05/17/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025463840 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, a high purge pressure alarm was triggered. the purge cassette was replaced to troubleshoot the issue, but the high purge pressure remained. the physician declined use of tissue plasminogen activator and the decision was ultimately made to replace the pump. there was no patient harm. | [5 empty] | The investigation for the high purge pressure (hpp) has been completed. no pump was returned for investigation. data logs show that on the last day of support there is a 3 min rise in pp and drop in purge flow triggering hpp and purge system blocked alarms. at the same time, flow increases and becomes saturated with upward shift in motor current (mc). the p-level is constant at p-4. the root cause of the high purge pressure was most likely biomaterial. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
26306 | 1220648-2024-11542 | 19374394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19374394&pc=OZD | 05/18/2024 | 07/17/2024 |  | 05/22/2024 |  |  | 05/18/2024 | 05/22/2024 |  |  | 08/31/2025 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402812 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/18/2024 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp placed to replace another support but, this 2nd pc pump had low flows/suction. the flows prompted replacement and a new pump and va ecmo (venoarterial extracorporeal membrane oxygenation) were delivered. | [5 empty] | Returned complaint cp pump visually inspected. no broken blades found. cannula kink near inlet cage observed. no biomaterial observed. the cause of the low pump flow was cannula kink near inflow cage due to improper technique. d9 added d4 expiration date and unique identifier (udi) missing. f6/f8-should have been left blank in the initial report. h4 device manufacture date missing. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
26307 | 1220648-2024-11543 | 19375116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19375116&pc=OZD | 05/19/2024 | 07/17/2024 |  | 05/22/2024 |  |  | 05/19/2024 | 05/22/2024 |  |  | 08/31/2025 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMPVVVVVVVVVV |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385983 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Migration or Expulsion of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2024 |  |  |  |  | The complainant reported a 42-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported the 5.5 pump suddenly decreased flow and alarmed for suction. flows couldn¿t reach greater than 0.8 liters per minute. the patient was moved to the operating room, where medical staff discovered the pump was oriented towards the mitral valve. flow issues remained despite successful repositioning. medical staff then decided to replace the pump. the patient survived the event, and no harm has been reported. | [5 empty] | The investigation for the low pump flow has been completed. the returned complaint 5.5 pump visually inspected. no cannula kink observed. a big chunk of biomaterial observed in the outlet flow cage surrounding the impeller and blocked the purge gap. data logs were reviewed which found positioning alarms occurred, but no indication confirmed related of the stated positioning issue. the cause of the low pump flow was biomaterial ingestion. the cause of the positioning issue cannot be determined as it is unclear based on the clinical details when the pump came out of position. added- d6b, d9 device return date f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
26308 | 1220648-2024-11544 | 19377166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19377166&pc=OZD | 04/22/2024 | 05/22/2024 |  | 05/22/2024 |  |  | 04/22/2024 | 05/22/2024 |  |  |  |  | 05/22/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1442220 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission |  |  |  |  | 04/22/2024 |  |  |  |  | The user facility reported a 58-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during preparation for concomitant impella 5.5 implant, a yellow error appeared on the aic (automated impella controller) directing the user to switch to a back-up console. the aic was subsequently replaced, but the original aic would not power off. it was noted that the console was still sensing a connected device despite all plugs and components being removed. the planned implant was successful with the replacement aic. however, it was later noted that the patient's family opted for comfort care due to the patient's poor prognosis. the patient subsequently passed away following the decision. it is unknown if the malfunction contributed to the patient's passing. | [5 empty] | The impella device was received from the customer on 25-apr-2024. problem description from ci-26678 / salesforce: ¿in cvor with patient for an impella 5.5. aic turned on, steps utilized to prime purge cassette and impella catheter without issue. awaiting for placement for about 15-20mins. prior to device placement impella yellow alarm came on stating ¿controller error-switch to back up controller.¿ new aic obtained and primed without issue. aic with issue would not turn off, it kept sensing a device connected even though the white connector plug and purge cassette were out. csc called and walked through how to override and power off aic. malfunctioning aic marked and tagged by staff at facility.¿
26309 | 1220648-2024-11545 | 19377816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19377816&pc=OZD | 11/02/2021 | 05/22/2024 |  | 05/22/2024 |  |  | 11/02/2021 | 05/22/2024 |  |  | 01/31/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022062612 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/02/2021 |  |  |  |  | The user facility reported that a patient was implanted with an impella cp for mechanical circulatory support due to acute myocardial infarction/cardiogenic shock (ami/csg). while on support, it was reported the patient had thrombocytopenia. no other information was available regarding care presented to the patient to resolve the reported issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26310 | 1220648-2024-11546 | 19378171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19378171&pc=OZD | 11/10/2021 | 05/13/2025 |  | 05/22/2024 |  |  | 11/10/2021 | 05/22/2024 |  |  | 08/31/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022135050 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Ischemia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced an episode of ventricular tachycardia during stenting of the right coronary artery and was shocked. the patient's peel away sheath was removed, and a repositioning sheath was inserted due to unsuccessful attempts to insert a retrograde sheath. the patient developed limb ischemia and the right foot had no pulse. the patient was successfully weaned and the impella explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: limb ischemia: impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ delivery issues: impella cp with smart assist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluation." | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26311 | 1220648-2024-11547 | 19378358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19378358&pc=PYX | 04/22/2024 | 06/12/2024 |  | 05/22/2024 |  |  | 04/22/2024 | 05/22/2024 |  |  | 09/30/2024 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023271008 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | Life Threatening | Death | Disability | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/22/2024 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella rp was placed in the left femoral and notable kink was seen in the sheath in the femoral-iliac anatomical region. 027 impella guidewire was advanced through dual lumen pulmonary artery (pa) catheter to hold position, then pa catheter advanced successfully through this kink in sheath. pa catheter to left pa and 027 wire was buried very deep in left pa. rp was advanced over wire and to left pa. the impella wire was removed and rp started at p2. impella stopped alarm occurred and motor current high on aic (automated impella controller). the impella was restarted at p2 successfully. then slow gradual increase in p level until reaching p9 with flows of 4-4.1 lpm. the case continued with no further similar alarms, only suction alarms. the case was completed. it was further reported that the patient expired. we do not have enough information to exclude impella as an associated factor in the patient's outcome. the death event occurred and no alternative cause for the death event is identified as the likely cause for such event. | [5 empty] | The investigation into the pump stop and the introducer damage ¿ mechanical issue has been completed since the original report was filed. the root cause of the introducer damage - mechanical issue was not determined since the introducer was not returned. the pump was returned for investigation and was able to be restarted and continued to run. the returned pump was able to be run in the lab. there was no biomaterial found around the impeller. the root cause of the temporary pump stop issue was not determined since the temporary pump stop could not be reproduced. in accordance with updated procedures, sections d6 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
26312 | 1220648-2024-11548 | 19378359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19378359&pc=OZD | 11/10/2021 | 03/31/2025 |  | 05/22/2024 |  |  | 11/10/2021 | 05/22/2024 |  |  | 08/31/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022129512 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Failure | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that during insertion, the wire was difficult to visualize and there was difficulty crossing the aortic valve (av). the impella appeared to be orientated toward the mitral valve. the physician attempted to reposition by advancing and pulling back on the impella, however, orientation could not be corrected. the groin had minimal oozing initially but after further movement, the right femoral impella site had increased bleeding. the patient was returned to or for cutdown and arterial repair. the bleeding resolved after the patient was given blood products. the impella was explanted after hemolysis and raising creatinine was noted, due to difficulty with positioning. the impella was successfully explanted. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26313 | 1220648-2024-11549 | 19378609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19378609&pc=OZD | 05/27/2022 | 05/20/2025 |  | 05/22/2024 |  |  | 05/27/2022 | 05/22/2024 |  |  | 10/31/2023 |  | 05/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022148511 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/27/2022 |  |  |  |  | The complainant reported that a 80 year old male patient was implanted with an impella cp for mechanical circulatory support, due to heart disfunction. while on support there was some bleeding from the impella insertion site, but was improving. heparin was reported to have been stopped or withheld. patient was successfully weaned from the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26314 | 1220648-2024-11434 | 19358285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19358285&pc=OZD | 11/13/2021 | 03/25/2025 |  | 05/21/2024 |  |  | 11/13/2021 | 05/21/2024 |  |  | 07/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110245 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2021 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during support, the patient experienced bleeding at the impella access site, hemolyzed laboratory samples and blood in urine. the impella had alarms for suction, low purge flow, and high purge pressure. staff relieved the touhy pressure and changed the cassette to resolve the low purge flow, and high purge pressure. the pump was decreased to p7 to resolve the suction alarms. systemic heparin was discontinued and the patient was transfused with one unit of packed red blood cells. hemostasis was achieved after discontinuing heparin and changing dressings. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related events are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26315 | 1220648-2024-11435 | 19358607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19358607&pc=OZD | 11/16/2021 | 05/21/2024 |  | 05/21/2024 |  |  | 11/16/2021 | 05/21/2024 |  |  | 08/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022132510 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/16/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on support, the patient experienced bleeding at the impella access site. options to reduce partial thromboplastin time (ptt) and best practices for site management were discussed with the care team. the patient was given one (1) unit for red blood cells. the access site bleeding reportedly was resolved. it was reported that the pt was also bleeding from an unknown source. care team was waiting for ct chest/abdomen/pelvis. no systemic heparin was being given. heparin was restarted later after reportedly being stopped due to oral bleeding the day prior. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: use of impella with transcatheter aortic valves as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
26316 | 1220648-2024-11436 | 19358682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19358682&pc=OZD | 11/10/2021 | 05/21/2024 |  | 05/21/2024 |  |  | 11/10/2021 | 05/21/2024 |  |  | 07/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022119661 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/10/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had bleeding noted around the impella re-access sheath, manual pressure was held. perclose deployed in right femoral artery, hematoma noted above site. bleeding continued from the impella site, pressure was held. the physician opted to keep heparin out of the purge for 8-12 hours. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26317 | 1220648-2024-11439 | 19358927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19358927&pc=OZD | 11/10/2021 | 05/21/2024 |  | 05/21/2024 |  |  | 11/10/2021 | 05/21/2024 |  |  | 07/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114882 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/12/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced ventricular tachycardia (vt), received shock by automatic implantable cardioverter defibrillators (aicd), and stabilized after lido bolus and procedure initiated. the patient experienced vt again and received another shock by aicd. additionally, the device pulled out and had to be repositioned. a bedside echocardiogram revealed an inlet of 4.7cm, the physician pulled the device back to 3.4cm, and successfully re-locked 81.5cm at hub. the patient also experienced limb ischemia, one point of femoral pulses was unable to be dopplered. then the patient had access site bleeding and surgical wound closure was required for explant. the impella was removed via vascular cutdown by the physician, thrombus was identified and taken out of common femoral artery. the physician replaced the impella cp with the impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ limb ischemia impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26318 | 1220648-2024-11440 | 19358946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19358946&pc=PYX | 06/23/2016 | 06/13/2024 |  | 05/21/2024 |  |  | 05/02/2024 | 05/21/2024 |  |  | 02/28/2017 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1200787 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 30 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The complainant reported a 30-year old female patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. abiomed received a publication titled "biventricular impella placement via complete venous access¿ regarding a 2016 impella rp-5.0 transcaval support where the patient developed a hematoma. over two days, the patient developed anemia and a hematoma in the left groin. the complainant reported the hematoma required conservative therapy without blood transfusion. medical staff successfully implanted biventricular assist devices four days later, and, ultimately, the patient required a transfusion of two units of packed red blood cells due to the anemia. the patient survived the event. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: iimpella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and sufficient information is not provided in the clinical description.
26319 | 1220648-2024-11441 | 19358954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19358954&pc=OZD | 12/21/2021 | 03/25/2025 |  | 05/21/2024 |  |  | 12/21/2021 | 05/21/2024 |  |  | 08/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022129847 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/21/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there was an alarm on the console showing that the impella was in the aorta, and a nurse reported that the impella had moved during a dressing change. the pump was repositioned multiple times, and the patient remained on impella support. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26320 | 1220648-2024-11442 | 19358973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19358973&pc=OZD | 05/07/2024 | 07/11/2024 |  | 05/21/2024 |  |  | 05/07/2024 | 05/21/2024 |  |  | 03/31/2025 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024330391 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an 81-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was bleeding heavily from the groin. the impella cp pump was explanted, and groin was surgically closed. the patient later died without pump support. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added- h6 impact code 4627, h6 component code 925. date of death is unknown. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26321 | 1220648-2024-11443 | 19358983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19358983&pc=OZD | 11/20/2021 | 05/21/2024 |  | 05/21/2024 |  |  | 11/20/2021 | 05/21/2024 |  |  | 06/30/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108275 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/20/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia. a reperfusion sheath was inserted initially to treat the cold leg, and then an antegrade sheath was placed to help perfuse leg. the patient remained on support, but it is unknown when the patient was weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26322 | 1220648-2024-11444 | 19358994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19358994&pc=OZD | 01/06/2022 | 03/25/2025 |  | 05/21/2024 |  |  | 01/06/2022 | 05/21/2024 |  |  | 09/30/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022143994 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2022 |  |  |  |  | The complainant reported a patent noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hematoma at the impella access site. femostop was placed with light pressure. hematoma was resolved and femostop removed. track ooze was developed and 2 preclose proglide was deployed with manual pressure. the patient developed slightly pink tinged urine. the labs were not hemolyzed. however, heparin was removed from purge and the device was repositioned slightly at bedside and the urine cleared. patient's condition was stable post procedure. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event (access site bleeding) are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system instructions for use for the related event (hemolysis) are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
26323 | 1220648-2024-11446 | 19359032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359032&pc=PYX | 10/01/2022 | 07/18/2024 |  | 05/21/2024 |  |  | 05/02/2024 | 05/21/2024 |  |  | 10/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 2022141215 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Edema | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | On 2024-05-02 complaints team forwarded publication from switzerland entitled: "acute pulmonary edema during impella rp support: too much of a good thing" in which the bipella rp patient had an adverse event; "right ventricular hemodynamic support with impella rp was initiated, shortly after the patient developed pulmonary edema, and intubation was necessary." the rp was explanted after 46 hours and patient placed on ecmo. | [5 empty] | The investigation for the pulmonary edema-vascular damage has been completed. no product was returned for investigation. the root cause of the vascular damage/pulmonary edema was not determined since product was not returned and there is a lack of clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
26324 | 1220648-2024-11445 | 19359033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359033&pc=OZD | 04/13/2022 | 04/15/2025 |  | 05/21/2024 |  |  | 04/13/2022 | 05/21/2024 |  |  | 11/30/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023160934 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2022 |  |  |  |  | The user facility reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient developed bleeding at the peel-way sheath. to resolve the issue, they changed from a 6fr sheath to a 7fr sheath. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26325 | 1220648-2024-11447 | 19359088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359088&pc=PYX | 07/25/2021 | 07/11/2024 |  | 05/21/2024 |  |  | 05/02/2024 | 05/21/2024 | 05/21/2024 |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  | IMPELLA RP | IMPELLA RP |  | PYX | PYX |  | IMPELLA RP | IMPELLA RP |  | ABIOMED, INC. | ABIOMED, INC. |  | US | US | UNKNOWN | UNKNOWN | UNKNOWN | UNKNOWN | No | No |  | Returned to Manufacturer | Returned to Manufacturer | Yes | No | Health Professional | Health Professional |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Literature | Health Professional | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | On 2024-05-02 complaints team forwarded publication out of bonn regarding 60 impella patients, 20 of which were rp, other 40 cp and 5.0. the entitled: "role of acquired von willebrand syndrome in the development of bleeding complications in patients treated with impella rp devices" detailed complications as 9 major bleed, 5 ecmo placements, 0 access site bleed, and 10 avw sydrome. patients were those implanted from january 2019 to december 2020. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added h6 component code 925. added e1 authors. authors: mehmet oezkur1,5*, sara reda2, heiko rühl2, nils theuerkauf3, stefan kreyer3, georg daniel duerr1,5, efstratios charitos1, miriam silaschi1, marta medina1, sebastian zimmer4, christian putensen3 & hendrik treede1,5 f6/f8 should have been left blank in the initial report. d1 and d2 corrected. d4 model number corrected. g4 pma/510k number corrected. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
26326 | 1220648-2024-11448 | 19359143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359143&pc=OZD | 05/07/2024 | 07/12/2024 |  | 05/21/2024 |  |  | 05/07/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485792 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unsuccessful impella cp implant attempt to a patient in acute myocardial infarction/cardiogenic. it was further noted, the staff felt the impella device was not fully primed, and the easy guide lumen was taken off. the pump was pulled off the wire and once there was confirmation of fluid coming from the motor housing, the staff attempted multiple times to place the wire through the pump in the proper location. the physician ultimately decided to open a new pump. there was no report of patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation into the delivery issues and the priming issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the delivery issue was use related since the pump was pulled off the wire due to concern of pump not purging. the cause of the priming issue was use related since the data logs showed purge values within specification indicating that the device was purging as intended.
26327 | 1220648-2024-11449 | 19359164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359164&pc=OZD | 04/01/2022 | 07/15/2024 |  | 05/21/2024 |  |  | 05/02/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Literature | Health Professional | User facility |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Ventricular Fibrillation | Cardiac Perforation | Thrombocytopenia | [6 empty] | Life Threatening | Disability | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | On 2024-05-02 complaints forwarded abstract from journal heart and lung transplantation entitled: "¿long-term¿ use of impella - safe to do?" in which impella supports from 5/2017 to 5/2021 were analyzed. the complications noted during impella cp support were hemolysis, access site bleeding, thrombocytopenia, vascular complications, ventricular arrhythmias, access site infection, sepsis, and cva. | [5 empty] | The investigation for the hemolysis, access site bleeding, thrombocytopenia, vascular complications, ventricular arrhythmias, access site infection, sepsis, and cva has been completed. without the device nor additional clinical details, the root cause of the issues could not be determined. g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
26328 | 1220648-2024-11450 | 19359232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359232&pc=OZD | 05/05/2024 | 06/10/2024 |  | 05/21/2024 |  |  | 05/05/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1204292 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complaint reported that the perfusionist changed the purge system and the purge disc was not recognized after repeating the procedure several times. the perfusionist opened and inserted a new purge system again, without success. therefore, the automated impella controller (aic) was changed, and the problem is solved. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
26329 | 1220648-2024-11451 | 19359238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359238&pc=OZD | 11/12/2021 | 05/21/2024 |  | 05/21/2024 |  |  | 11/12/2021 | 05/21/2024 |  |  | 07/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022120606 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/12/2021 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support due to elevation myocardial infarction. while on support, the patient struggled to keep mean arterial pressure above 60. the patient went into ventricular tachycardia when pressure levels were lowered to 5. the patient was shocked once. the patient¿s labs were hemolyzed, however, urine was still clear and yellow. the decision was made to explant the device due to hemolysis concerns. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular. perforation).¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
26330 | 1220648-2024-11452 | 19359261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359261&pc=PYX | 05/08/2024 | 07/22/2024 |  | 05/21/2024 |  |  | 05/08/2024 | 05/21/2024 |  |  | 04/30/2024 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023228000 |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during implant procedure due to the 74-year-old male patient's dilated right ventricle, the impella rp was unable to be placed. the insertion was therefore aborted. there was no reported patient harm, and the patient was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella rp smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ impella rp smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient condition since insertion was aborted due to a dilated right ventricle.
26331 | 1220648-2024-11454 | 19359266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359266&pc=OZD | 11/24/2021 | 03/25/2025 |  | 05/21/2024 |  |  | 11/24/2021 | 05/21/2024 |  |  | 08/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022124832 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported during impella insertion, the right femoral artery was found to have a dissection after the insertion of the 13 cm sheath with introducer. a covered stent was placed over the dissection and the decision was made to use a longer peel away sheath due to further proximal tortuosity. it was reported that after removal of the impella, the perclose device was unsuccessful in closing the impella access site due to calcification. manual pressure was held to achieve hemostasis. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: warnings & cautions: cautions physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
26332 | 1220648-2024-11453 | 19359268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359268&pc=OZD | 05/05/2024 | 06/11/2024 |  | 05/21/2024 |  |  | 05/05/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025486126 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/05/2024 |  |  |  |  | The user facility reported a 68-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a patient experienced oozing/bleeding from their access site following impella cp implant. to manage the condition, manual pressure was applied, sutures were tightened, and two units of packed red blood cells were transfused. the patient remained on support until successfully weaned and explanted. upon explanting the pump perclose was used and failed so covered stent was deployed and hemostasis was achieved. the patient was reported as stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue since tightening the perclose suture achieved hemostasis. g1- corrected mfr site name and address | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26333 | 1220648-2024-11455 | 19359357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359357&pc=OZD | 11/12/2021 | 05/21/2024 |  | 05/21/2024 |  |  | 11/12/2021 | 05/21/2024 |  |  | 06/30/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094984 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/12/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. while attempting to implant the impella cp, the delivery was unsuccessful due to the vessel size of the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26334 | 1220648-2024-11456 | 19359474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359474&pc=OZD | 05/09/2024 | 07/09/2024 |  | 05/21/2024 |  |  | 05/09/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025464727 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient had an attempted impella 5.5 implant, however it was unsuccessful. the implant was aborted due to the patient's anatomy. | [5 empty] | The investigation into the delivery issues been completed since the original report was submitted. the device was not returned for investigation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. g.1 revised reporting contact fax number from manufacturer device report number 1220648-2024-11456 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿
26335 | 1220648-2024-11457 | 19359486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359486&pc=OZD | 04/29/2024 | 06/07/2024 |  | 05/21/2024 |  |  | 04/29/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1327897 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the purge disc detection of the automated impella controller (aic) was broken. the issue was discovered during prep, and there was no patient harm reported. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. it was only reported that the purge disk detection flag broken. the purge disc flag was confirmed to be broken. the root cause of the issue was a broken purge disk flag. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
26336 | 1220648-2024-11458 | 19359494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359494&pc=OZD | 05/05/2024 | 07/08/2024 |  | 05/21/2024 |  |  | 05/05/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2025456532 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support for 71-year-old female patient, the purge pressure was high (1000mmhg pf <1ml/h). there was no purge line kink observed. tpa was added to the purge. there was no reported patient harm. | [5 empty] | The investigation for the reported high purge pressure has been completed. clinical details states that high purge pressure was reported on day 5 of support. no product was returned. the data logs confirm multiple high purge pressure and purge system blocked alarms. the data logs show a motor current spike followed by a rise in purge pressure before alarm was triggered. there was a simultaneous low purge flow. the condition resolved after about 4 hours. it is unknown if tissue plasminogen activator (tpa) was added to the purge. the root cause of high purge pressure was most likely biomaterial related due to gradual rise in purge pressure following a motor current spike. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist in section: using the automated impella controller with the impella catheter. added h6 impact code 4644 g1-mdr reporting address updated g1-mfr site name and address updated. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
26337 | 1220648-2024-11459 | 19359547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359547&pc=OZD | 05/03/2024 | 07/22/2024 |  | 05/21/2024 |  |  | 05/05/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025480738 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that two hours following successful impella cp weaning and explant, the patient underwent an arterial thrombectomy in the foot. afterwards, the patient had acute shock symptoms and was diagnosed with intestinal ischemia and expired a short time later. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. added- h6 component code 925 g1-corrected mfr site name and address.
26338 | 1220648-2024-11461 | 19359648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359648&pc=OZD | 05/06/2024 | 07/29/2024 |  | 05/21/2024 |  |  | 05/06/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1287215 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation | Distributor |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/06/2024 |  |  |  |  | The distributor reported a patient (unknown race, gender and age) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on support, the automated impella controller (aic) displayed a yellow alarm. the aic was swapped in response to the failure. there was no patient harm. | [5 empty] | The investigation for the reported console (aic) yellow alarm has been completed. the console was returned for evaluation. upon functional inspection, the 4-in-1 gen2 assy was tested and found noise in the fringe pattern. the optical bench was tested against a known good console and found that it produced repetitive controller errors. data logs were reviewed which found it recorded ¿processsampleforopmdatapacketcrcerror: calculator placement signal 1045¿ repeatedly, the console displayed ¿controller error (opm comm crc invalid) [optical pump connected] ¿ alarm¿ event #1045. real time logs confirmed that all signals received from the optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. then, event #1045 occurred 23 times, and every time, it was resolved immediately. this confirms the reported failure mode. the root cause for the controller error and loss of placement signal was an optical bench fw communication protocol anomaly, resulting in the misalignment of data communication packets received by epc. the aic sw operated as designed. f6/f8 should have been left blank in the initial report. h6 med dev prob code 2885 changed to 2880. h8-changed to reuse. | The aic was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26339 | 1220648-2024-11460 | 19359649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359649&pc=OZD | 05/03/2024 | 07/08/2024 |  | 05/21/2024 |  |  | 05/03/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025488344 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after removing the impella cp from the 73-year-old male patient, the manta closure system was inserted with significant bleeding. the decision was made to implant two 9x50mm stents and apply manual compression. the bleeding resolved. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the following sections: section: general patient care considerations. section: entering cardiac output. section: potential adverse events. (united states) added- h6 component code 925 h6 impact code 4626 should be 4641. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26340 | 1220648-2024-11463 | 19359662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359662&pc=OZD | 05/05/2024 | 07/12/2024 |  | 05/21/2024 |  |  | 05/05/2024 | 05/21/2024 |  |  | 11/30/2025 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024427374 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 61-year-old male noted as high-risk coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. during patient support, the impella 5.5 was dislocated. therefore, the physician decided to have the impella 5.5 explanted and replaced with an impella cp. the patient is noted to be stable and no harm was reported. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for investigation. based on clinical details, the root cause of the positioning issues is most likely due to use as the pump was pulled out. e.1 facility name was revised as previously entered incorrectly on the initial manufacturer device report 1220648-2024-11463. g.1 revised reporting contact fax number in accordance with updated procedures since manufacturer device report 1220648-2024-11463 was submitted. g.1 revised manufacturing site name and address section as previously entered incorrectly on manufacturer device report 1220648-2024-11463. h.6 code 3038 was reported incorrectly on manufacturer device report 1220648-2024-11463. a new code has been added to component codes. h.10 additional manufacturer narrative instructions for use (ifus) were reported on manufacturer device report 1220648-2024-11463 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26341 | 1220648-2024-11462 | 19359664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359664&pc=OZD | 05/10/2024 | 07/16/2024 |  | 05/21/2024 |  |  | 05/10/2024 | 05/21/2024 |  |  | 04/30/2025 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330814 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hemorrhagic Stroke | [11 empty] | Life Threatening | Disability | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that care was withdrawn from the patient and the patient subsequently expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate the use of the device with the patient's outcome. | The user facility reported a patient in acute myocardial infarction/cardiogenic shock implanted with an impella 5.5 device for mechanical circulatory support. the device was implanted and over that night, it was noted the patient had been coagulopathic with oozing at the impella access site and developed a small hemorrhagic stroke. the patient received multiple blood and blood products and anticoagulation was withheld due to bleeding. the patient was noted to be in stable condition. |  |  |  |  | B.5 additional information was received, and abiomed's medical safety officer review was completed. d.6b explant date was added. the investigation for the reported access site bleed and stroke has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding issue was not determined since the product was not returned and sufficient information was not provided in the clinical description. the true root cause of the stroke could not be determined as sufficient information was not provided. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-11462 was submitted. h.6 added code 925 in accordance with updated procedures since the initial manufacturer device report 1220648-2024-11462 was submitted. h.10 additional manufacturer narrative instructions for use were revised from manufacturer device report 1220648-2024-11462 in accordance with updated procedures. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
26342 | 1220648-2024-11464 | 19359686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359686&pc=OZD | 04/24/2024 | 07/09/2024 |  | 05/21/2024 |  |  | 05/07/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025455705 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Life Threatening | Death | Disability |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support developed a subarachnoid hemorrhage (sah) while on support and died of multiple organ failure. additional information received noted the physician was unsure of the sah cause and unclear whether there was a connection or not between the event and the impella cp device. the timing was unknown as the computed tomography scan had not been taken and extracorporeal membrane oxygenation was also inserted. there was no treatment for the sah, and the patient had eventually developed multiorgan failure due to the onset of sah. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the ventricle perforation issue has been completed since the original report was submitted. no product was returned for evaluation. data logs were not returned. clinical details noted that cardiac tamponade was noted after impella 5.5 procedure and pericardial effusion was performed. active bleeding was still noted after drainage and blood products were administered. no impairment of impella operations were noted. the root cause of the ventricle perforation could not be determined as no product or logs were returned for evaluation. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ e.1 first name was revised as previously entered incorrectly on the initial manufacturer report number 1220648-2024-10500. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer report number 1220648-2024-10500 was submitted. h.10 additional manufacturer narrative instructions for use - a portion was omitted from the initial manufacturer report number 1220648-2024-10500 and has been added.
26343 | 1220648-2024-11465 | 19359702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359702&pc=OZD | 05/09/2024 | 07/05/2024 |  | 05/21/2024 |  |  | 05/09/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489221 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported there was an inability to deliver an impella cp device for the support of a patient with a high-risk percutaneous coronary intervention procedure. the pump failed due to anatomy despite all efforts and wire exchanges. there was no report of patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the pump failed to deliver due to anatomy despite all efforts and wire exchanges. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ d4-corrected model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26344 | 1220648-2024-11466 | 19359738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359738&pc=OZD | 05/06/2024 | 06/13/2024 |  | 05/21/2024 |  |  | 05/06/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467983 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2024 |  |  |  |  | The complainant reported a patient post cardiotomy cardiogenic shock / low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support after having undergone coronary artery bypass grafting x 3 with repair to left ventricle free wall. the implant was unsuccessful. the wire and pigtail was across the aortic valve and, unfortunately, the impella was unable to cross the valve into the left ventricle. after attempting for numerous hours, the physician decided to abort the impella 5.5 insertion and close the patient. the patient was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the pump was unable to cross the valve into left ventricle.
26345 | 1220648-2024-11467 | 19359746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359746&pc=OZD | 05/10/2024 | 06/28/2024 |  | 05/21/2024 |  |  | 05/10/2024 | 05/21/2024 |  |  | 03/31/2025 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024324228 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella 5.5 pump began to experience high purge pressure during patient support. tissue plasminogen activator was run through the purge to manage the noted issue. support was maintained with the implanted pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. it was reported that the impella 5.5 pump began to experience high purge pressure during patient support. the data logs confirm high purge pressure alarms and show a gradual buildup in biomaterial. the root cause of the high purge pressure was most likely biomaterial buildup.
26346 | 1220648-2024-11468 | 19359781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359781&pc=OZD | 05/10/2024 | 07/10/2024 |  | 05/21/2024 |  |  | 05/10/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440639 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was to be implanted with an impella cp for mechanical circulatory support. the impella cp was unable to be repositioned away from the mitral valve. the impella cp was exchanged. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: delivery issues impella cp with smart assist for use during cardiogenic shock for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related. f6/f8 should have been left blank in the initial report. g1 reporting contact fax number missing. h6 med dev prob code 1158.
26347 | 1220648-2024-11469 | 19359790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359790&pc=OZD | 05/12/2024 | 07/15/2024 |  | 05/21/2024 |  |  | 05/12/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448984 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that hemolysis was suspected in the patient on impella cp support. it was further reported that hemolysis was due to incorrect position of the impella. a decision was made to explant device after patient recovery. the duration of impella support lasted 15 hours. | [5 empty] | The investigation for the hemolysis has been completed. data logs showed pump was started and run without issue. about 3.5 hours into the case, low flow occurred and remained to the rest of the case that triggered auto suction response and suction alarms. logs also showed pump was in aorta in a short time and was resolved during support. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
26348 | 1220648-2024-11470 | 19359798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359798&pc=PYX | 05/09/2024 | 07/12/2024 |  | 05/21/2024 |  |  | 05/09/2024 | 05/21/2024 |  |  | 10/31/2025 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024413929 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The user facility reported a 66-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported after 73 hours of support the pump was explanted as it had come out of position and was causing arrhythmias. the patient was reported as stable. | [5 empty] | The investigation into the positioning issues and the arrhythmia has been completed since the original report was submitted. the device was not returned for investigation. data logs were reviewed and no positioning alarms were noted. the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information provided. as the necessary information was not provided, the root cause of the positioning issue and arrhythmia was not determined. h6 medical device problem code 2993 should be removed from initial report. f6 and f8 should have been blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
26349 | 1220648-2024-11471 | 19359864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359864&pc=OZD | 05/10/2024 | 07/12/2024 |  | 05/21/2024 |  |  | 05/10/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1091702 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that care was withdrawn from the patient and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. | The user facility reported that the automated impella controller (aic) while in use annunciated "controller error" alarms five times between day of implant and the next day. additionally, on the next day, the aic battery level instantaneously drop from 100% to 0% and then back to 100%. there was no report of patient harm. |  |  |  |  | The investigation for the reported console (aic) battery failure (red alarm) has been completed. the console was returned for evaluation. functional testing was performed and found that on initial bootup, the console displayed ¿system self check failed.¿ upon reviewing the logs via telnet, it was confirmed that the reported failure mode symptoms were reproduced (communication issue with powermod_1). the communication issue was persistent. upon visual inspection, a cable from pbm (j13) was not fully seated at the 4-in-1 connector (x45). there was no physical damage found on the connector. further testing of the aic was performed by running the optical test pump for more than 24 hours and the controller error was not reproduced. data logs were reviewed confirmed that upon several manual boot ups, the console displayed ¿controller error (loss of comm (pbm1)) -alarm¿, event #1023. the root cause for the controller error was a loose cable connection at the 4-in-1 connector (x45). added- b5 mso review, d9 device return date. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26350 | 1220648-2024-11472 | 19359874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359874&pc=OZD | 05/13/2024 | 08/08/2024 |  | 05/21/2024 |  |  | 05/13/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485193 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after implantation of the impella cp, the patient developed a hematoma. heparin was stopped and a sandbag was applied to achieve hemostasis. the patient remained on impella support for an additional day. | [5 empty] | The investigation for the reported access site bleed has been completed. the impella was returned for evaluation and there were no issues found in the returned product. without sufficient clinical information, the root cause of the access site bleed could not be determined. b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-11472. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11472 and should not have been. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26351 | 1220648-2024-11473 | 19359900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359900&pc=OZD | 05/14/2024 | 07/15/2024 |  | 05/21/2024 |  |  | 05/14/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025500826 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and exhibited bleeding from the femoral impella access site. consequently, heparin was withheld to obtain hemostasis. the patient was noted to be in stable condition. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was determined to be use issue because the complication was managed by tightening perclose sutures. d4-corrected model number. | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26352 | 1220648-2024-11474 | 19359930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359930&pc=OZD | 11/24/2021 | 03/25/2025 |  | 05/21/2024 |  |  | 11/24/2021 | 05/21/2024 |  |  | 04/30/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022071839 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported that there was a large hematoma at the impella site due to significant bleeding. the complication was successfully managed by administering blood products to the patient, removing the impella and placing a larger non-peel-away sheath. no further information was provided. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26353 | 1220648-2024-11475 | 19359942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359942&pc=OZD | 05/14/2024 | 07/05/2024 |  | 05/21/2024 |  |  | 05/14/2024 | 05/21/2024 |  |  | 11/30/2025 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024423296 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The user facility reported a 64-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported after the implant the, the patient's distal flow was sluggish and external fem-fem bypass was placed. after five days of support patient developed limb ischemia and a thrombectomy was performed that was partially successful. the decision was made to replace the impella cp with an impella 5.5. the patient was reported as stable. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. f6/f8-should have been left blankh6 impact code 4627 changed to 4629. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
26354 | 1220648-2024-11476 | 19359971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359971&pc=OZD | 05/15/2024 | 07/16/2024 |  | 05/21/2024 |  |  | 05/15/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1663002 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported impella automated impella controller (aic) while in use had an unexpected shut down. the screen went black and turned back on in approximately 10 seconds. when the aic turned back on, the screen looked purple. when the aic turned back on, impella flows slowly ramp back up to the p-level p-8 that it was on before the shutdown. the aic was exchanged, and the patient was noted to be in stable condition. | [5 empty] | The investigation for the console (aic) shutdown issue has been completed. the console was returned for evaluation. during functional inspection, the reported complaint could not be reproduced. the batteries were tested and confirmed to be working correctly. a firmware check was performed, and communication between the 4-in-1 and pbm was confirmed to be working correctly. the pbm was temporarily removed from the console and inspected for physical defects, and none were found. data logs were reviewed which revealed an occurrence of an unexpected controller shutdown, and immediately preceding the unexpected controller shutdown, battery 1 was reporting an unknown error via its status bits. battery manager channel 1, for battery 1, was also not accurately detecting the state of the ac power connection, in comparison to battery manager channel 2. the discrepancy in status and charging information between battery 1 and battery 1 indicated a potential malfunction of the pbm preceding the unexpected controller shutdown. the root cause of the reported complaint (unexpected controller shutdown) was determined to most likely have been an intermittent malfunction of the pbm. no other potential causes were identified. added-d9 device return date. d4-corrected model number. | The automated impella controller (aic) was not received from the customer; and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26355 | 1220648-2024-11477 | 19359978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19359978&pc=OZD | 05/13/2024 | 07/15/2024 |  | 05/21/2024 |  |  | 05/15/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501580 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support during pci. during support, there was concern for the patient's mental status. the patient was taken for a ct scan and went into cardiac arrest. return of spontaneous circulation was achieved and the patient was diagnosed with a cardiac tamponade and a large pericardial effusion was noted. the team attempted to drain and the patient arrested again. the ct scan showed apical wall perforation. the etiology of tamponade is unclear and team is considering impella (during pci) versus wire versus coronary perforation. the patient was successfully weaned from support two days later. | [5 empty] | The investigation for the ventricle perforation has been completed. product was not returned for investigation. the cause of the ventricle perforation issue was not determined since product was not returned and sufficient clinical information was not provided. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torqueing the impella, adjustments should be performed under imaging guidance.¿
26356 | 1220648-2024-11478 | 19360047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19360047&pc=OZD | 05/15/2024 | 07/12/2024 |  | 05/21/2024 |  |  | 05/15/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025463315 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Contamination | Use of Incorrect Control/Treatment Settings | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. the team reported that the device was contaminated during preparation for insertion. the team exchanged the device with another impella cp. there was no patient harm. there was no allegation that device was contaminated when it was received. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for investigation. the root cause of the delivery issue was most likely use issue since the pump was contaminated during the preparation. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d.6a revised from manufacturer device report 1220648-2024-11478 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report 1220648-2024-11478 and should not have been. h.6 code 1126 was reported incorrectly on manufacturer device report 1220648-2024-11478. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from manufacturer device report 1220648-2024-11478. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26357 | 1220648-2024-11479 | 19360069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19360069&pc=OZD | 05/13/2024 | 07/21/2024 |  | 05/21/2024 |  |  | 05/13/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494690 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The user facility reported a 69-year-old female with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the patient was on support because they needed an aortic dissection repair, which was unrelated to the impella device. the patient was brought to the operating room for the dissection procedure, and the pump migrated into the aorta. the doctor then decided to remove the impella cp. the patient survived the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. data logs confirmed the pump was in the aortic area (about 45 minutes into the case). the pump then was stopped manually and disconnected from the console. the product was not returned for review. based on clinical description and data analysis, the root cause of positioning issue was most likely patient movement since the issue occurred while the patient was being moved to operating room table. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11479 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-11479 was submitted. h.6 added code 4629 as it was inadvertently left off the previously submitted manufacturer device report number 1220648-2024-11479.
26358 | 1220648-2024-11480 | 19360165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19360165&pc=OZD | 12/15/2023 | 06/03/2024 |  | 05/21/2024 |  |  | 04/25/2024 | 05/21/2024 |  |  | 06/30/2025 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379249 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured ventricular tachycardia with a "possible" relationship to the impella device used: ¿pt went into ventricular tachycardia with a pulse. code blue called. synchronized cardioversion with 200j of energy performed. repeat rhythm was sinus. dopamine gtt discontinued.¿ the patient recovered with no sequelae. "rn observed on monitor that pt was in monomorphic ventricular tachycardia with rates in the 200¿s. pt cardioverted a total of 7 times at 200j. in between shocks pt was given a 150 mg amiodarone bolus and 100 mg lidocaine. pt became minimally responsive and was in sinus bradycardia in the low 50's, rate continued to drop and therefore the decision was made to intubate the patient with versed and etomidate. impella was repositioned x2 with echo at bedside and 2 cardiologists; post repositioning pt had return of bradycardic rhythm and was given 0.25 mg epi to assist with hr and bp while dobutamine was increased to 7.5 and norepi was started at 5. pt remained stable for roughly 1 hour." the patient recovered with no sequelae. "pt went back into monomorphic ventricular tachycardia in the 150's, synchronized shock at 200j and 150 mg amio bolus given, amio gtt resumed at 1 and pt returned to sinus bradycardia w/ pacs and pvcs." the patient recovered with sequelae. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined.
26359 | 1220648-2024-11481 | 19360201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19360201&pc=OZD | 04/15/2022 | 05/21/2024 |  | 05/21/2024 |  |  | 04/15/2022 | 05/21/2024 |  |  | 08/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP | [5 empty] | US |  | 2022134237 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/15/2022 |  |  |  |  | The complainant reported a 66 year old female patient needing an impella cp for mechanical circulatory support due to heart failure issues. the patient coded prior to the procedure, however was stabilized and the impella cp was placed in the right femoral artery. during support, it was reported that the patient was stable, however there was retroperitoneal bleeding (rp) noted. the patient received blood for the rp. in addition, there were multiple strokes noted on ct. the patient was not on any pressers or inotropes, and the impella groin site was within normal limits. the patient had good urine output and good pulses bilaterally. the impella was weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26360 | 1220648-2024-11482 | 19360361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19360361&pc=OZD | 11/24/2021 | 03/25/2025 |  | 05/21/2024 |  |  | 11/24/2021 | 05/21/2024 |  |  | 09/30/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022139441 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2021 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported that the wire perforated the distal circumflex artery. balloon tamponade was attempted for 20 minutes but was unsuccessful. papyrus 2.0, 2.5 and 3.0 micro catheters all trapped. the physician attempted to use a coil and the balloon did not deflate, so the wire stretched out and also got stuck. the balloon was inflated for over two hours with attempts to snare unsuccessful. additionally, it was also reported that the patient had no pulse in the right leg and slight oozing was noted at the groin access site. three units of blood were administered to the patient and systemic heparin was withheld. the device was subsequently weaned successfully and explanted. no further information was provided. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
26361 | 1220648-2024-11483 | 19360501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19360501&pc=OZD | 11/25/2021 | 05/21/2024 |  | 05/21/2024 |  |  | 11/25/2021 | 05/21/2024 |  |  | 09/30/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022135692 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/25/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding. perclose sutures were unsuccessful, manipulating pressure was then applied and hemostasis was achieved. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer . as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26362 | 1220648-2024-11484 | 19362097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19362097&pc=OZD | 04/20/2022 | 05/21/2024 |  | 05/21/2024 |  |  | 04/20/2022 | 05/21/2024 |  |  | 12/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173629 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/20/2022 |  |  |  |  | The complainant reported that a 72 year old male patient was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced oozing of the left femoral artery at the impella site. the patient was weaned from the impella and the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26363 | 1220648-2024-11290 | 19363498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19363498&pc=OZD | 04/16/2024 | 06/07/2024 |  | 05/21/2024 |  |  | 04/16/2024 | 05/21/2024 |  |  | 02/28/2025 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320385 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/16/2024 |  |  |  |  | The user facility reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a left ventricle thrombus was found. the thrombus was confirmed with echocardiogram. the family withdrew care and the patient expired after 72.04 hours of support. there is not have enough information to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
26364 | 1220648-2024-11397 | 19363506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19363506&pc=OZD | 01/04/2022 | 05/20/2024 |  | 05/21/2024 |  |  | 01/04/2022 | 05/21/2024 |  |  | 10/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022150450 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/04/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the patient experienced significant swelling at the impella cp site. manuel pressure was held for 20min. right heart cath was completed. the physician was pleased with hemodynamics, but decided to explant the device. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26365 | 1220648-2024-11348 | 19363509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19363509&pc=OZD | 05/10/2024 | 07/08/2024 |  | 05/21/2024 |  |  | 05/10/2024 | 05/21/2024 |  |  | 07/31/2025 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442834 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility reported a 71-year-old male with cardiomyopathy was to be implanted with an impella cp device for mechanical circulatory support. it was reported that resistance was met during attempted delivery of an impella cp pump. it was noted that despite multiple attempts, the catheter would buckle when trying to cross the aortic valve. it was noted that the aortic valve was a new bioprosthetic valve which resulted in some stiffness. this combined with potential thickening of the valve via endocarditis, prevented the catheter from crossing the valve. during this time the patient was shocked for frequent ventricular tachycardia (vt) and amiodarone and lidocaine boluses were given. the implant was subsequently aborted. the patient was reported stable. | [5 empty] | The investigation for the reported ventricular arrhythmia and delivery issues has been completed. the impella device was not returned for evaluation. however, clinical information was sufficient to determine root cause. the root cause of the delivery issue was most likely patient condition due to newer, stiff aortic valve as well as thickening of the valve due to endocarditis preventing the valve from opening enough to allow the impella to cross. clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system . section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿
26366 | 1220648-2024-11354 | 19363510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19363510&pc=OZD | 04/24/2024 | 09/24/2025 |  | 05/21/2024 |  |  | 04/24/2024 | 05/21/2024 |  |  | 05/31/2025 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024344565 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/24/2024 |  |  |  |  | The user facility reported that a patient undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the physician attempted to reposition the impella and it inadvertently pulled back into the ascending aorta. multiple attempts were made to re-advance the pump back into the left ventricle, but attempts were unsuccessful and the device kinked partially due to the long tunnel and hard pushing. the physician made the decision to implant another impella 5.5. | The complainant also reported that, while on support, the pump experienced a placement signal not reliable alarm that was resolved with alternative means of monitoring. |  |  |  |  | The impella device was received from the customer however the evaluation of the device has not yet been completed. upon evaluation and/or investigation closure, a supplemental mdr will be filed. | The investigation for the reported optical issues and positioning issues has been completed. the impella 5.5 pump was returned. the pump was visually inspected and catheter was found to be torqued/kinked during support. no other physical abnormalities or damages were found. the optical bench 5s parameters were found to be normal and the light continuity test passed without any issues. pump was connected to the console and ran at p-9 in bucket of water. no optical signal issue or alarm reproduced during the testing. data logs were reviewed, and placement signal unreliable alarm observed during support on (b)(6) 2024 with placement signal spiking. the optical 5s parameters snr, contrast was found to be decreasing, while the lamp, gain had a slight increase in trend. the pump was changed on the second console and the snr, contrast was found little variable all over the place and low in range. the before pumps used on first and second console had no similarities of optical 5s failure trend. no other abnormalities were found in the logs. the cause of the positioning issue was use related per return product investigation and clinical review. added- h6 impact code 4629 in accordance with updated procedures, sections d6 have been revised from the initial submission. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the cause of the optical signal issue was not determined since issue could not be reproduced with return product and based on inconclusive evidence per log review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number -2. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email updated.
26367 | 1220648-2024-11349 | 19363512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19363512&pc=OZD | 05/11/2024 | 07/16/2024 |  | 05/21/2024 |  |  | 05/13/2024 | 05/21/2024 |  |  | 11/30/2025 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024428462 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella cp pump experienced positioning issues resulting in a complete atrioventricular block. the pump was removed in response to the noted issue and a new impella cp was implanted for ongoing patient support. there was no patient harm. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for investigation and data logs were not available to review. the cause of the positioning issue most likely use related due to improper technique resulting in a complete atrioventricular block. f6 and f8 should have been blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
26368 | 1220648-2024-11353 | 19363517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19363517&pc=OZD | 05/14/2024 | 07/19/2024 |  | 05/21/2024 |  |  | 05/14/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025452382 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The complainant reported an 80-year-old male patient with acute myocardial infarction / cardiogenic shock and pre-operation placement was implanted with an impella cp for mechanical circulatory support. the complainant reported ventricular septal perforation for the patient. the impella was operating at p8 level, but. due to hemolysis, the level was lowered. medical staff administered haptoglobin, which resolved the issue. the patient survived the event and the cp was eventually explanted. | [5 empty] | The investigation for hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g2 report source added foreign f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
26369 | 1220648-2024-11437 | 19363523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19363523&pc=OZD | 05/01/2024 | 06/25/2024 |  | 05/21/2024 |  |  | 05/01/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025465322 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported that the patient was placed onto impella cp support for acute myocardial infarction / cardiogenic shock. while on support, the pump experienced positioning issues that could not be resolved despite multiple attempts. the pump was left as is, pointed toward the mitral valve. | [5 empty] | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. based on clinical description and data analysis, the root cause of positioning issue was most likely use related. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
26370 | 1220648-2024-11301 | 19363534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19363534&pc=OZD | 02/17/2022 | 03/25/2025 |  | 05/21/2024 |  |  | 02/17/2022 | 05/21/2024 |  |  | 10/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023159824 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, physicians attempted to perform a pre-closure of the access site, but the perclose device failed and a hematoma developed at the site and the procedure was aborted. subsequently, the impella was left in the patient overnight so that the patient could receive surgical repair of the access site during device removal. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26371 | 1220648-2024-11438 | 19363543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19363543&pc=OZD | 12/04/2021 | 03/25/2025 |  | 05/21/2024 |  |  | 12/04/2021 | 05/21/2024 |  |  | 07/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022122815 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2021 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported patient was brought down for bilateral reperfusion sheaths in legs. patient had fasciotomies on all extremities. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26372 | 1220648-2024-11350 | 19363549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19363549&pc=OZD | 09/14/2023 | 06/27/2024 |  | 05/21/2024 |  |  | 05/14/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A publication reviewed indicated a patient was implanted with an impella 5.5 pump and received support for 16 days before being removed. it was indicated there was a 1-centimeter thrombus noted at the impella 5.5 pump's cannula and led to the pump being surgically explanted. | [5 empty] | The investigation into the thrombus issue been completed since the original report was submitted. the device was not returned for investigation. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ note: product-specific information: model/catalog and lot/serial numbers were not made available through the article. respectively, sections d4 fields are unknown.
26373 | 1220648-2024-11351 | 19363559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19363559&pc=OZD | 05/14/2024 | 11/20/2024 |  | 05/21/2024 |  |  | 05/14/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1579730 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The user facility reported a 58-year-old male patient was implanted with an impella 5.5 as a bridge to transplant. it was reported that a loss of placement signal occurred during support. the automated impella controller (aic) was swapped, and the issue resolved. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The investigation into optical signal issue has been completed. the issue was not able to be reproduced in the lab when running a known good pump and purge system. the logs confirmed that a placement signal not reliable alarm was triggered on the reported complaint date. the logs confirmed that immediately prior to and following the placement signal not reliable alarm, the optical bench reported an oscillating contrast value far outside the normal range of contrast. the cause of the automated impella controller (aic) issue was a defective optical bench. d9 date device returned to manufacturer added. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11351. d2 common device name. d3 fax number missing. d6b missing. f6/f8-should have been left blank. h1 correction inadvertently selected. h11 corrected data entered incorrectly.
26374 | 1220648-2024-11396 | 19363573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19363573&pc=OZD | 01/04/2022 | 05/20/2024 |  | 05/21/2024 |  |  | 01/04/2022 | 05/21/2024 |  |  | 10/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151353 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 96 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/04/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced an improving hematoma prior to implant of the impella. it was reported that the appearance of the hematoma started to look worse as it increased to the size of a grapefruit and was hard. the patient was taken for surgical cut down and the impella was explanted. the patient was given two units of packed red blood cells. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26375 | 1220648-2024-11398 | 19363580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19363580&pc=OZD | 03/28/2022 | 05/20/2024 |  | 05/21/2024 |  |  | 03/28/2022 | 05/21/2024 |  |  | 12/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178884 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/28/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella cp device for mechanical circulatory support. the impella cp was unable to advance a 14 x 25 cm sheath through the patient¿s vasculature due to acute calcified angulation. the medical team replaced the 14x25 cm sheath with a short impella sheath and an attempt was made to advance the impella catheter, which was unsuccessful. the impella was removed and a 16f gore dryseal which was able to advance past angulation, anf them impella was placed without issues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
26376 | 1220648-2024-11485 | 19363920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19363920&pc=OZD | 05/14/2024 | 07/12/2024 |  | 05/21/2024 |  |  | 05/14/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481343 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The user facility reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma at their groin site during impella cp support. the decision was made to explant the pump. the patient was stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. h6 added component code f6/f8 should have been left blank. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26377 | 1220648-2024-11486 | 19364919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19364919&pc=OZD | 05/14/2024 | 07/16/2024 |  | 05/21/2024 |  |  | 05/14/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025460516 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The user facility reported a 76-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was placed when the impella 5.5 failed due to anatomy size. the impella cp was placed to allow the patient to have valve surgery. the pump was exhibiting low flows and was removed and replaced after two hours. the low flow pump was seen to have thrombosis. | [5 empty] | The investigation for the reported low pump flow has been completed. the product and data logs were not returned. the cause of the low pump flow was not determined since the complaint device not returned for investigation and no data logs were available for review. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26378 | 1220648-2024-11487 | 19366744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19366744&pc=OZD | 04/21/2024 | 09/12/2024 |  | 05/21/2024 |  |  | 04/21/2024 | 05/21/2024 |  |  |  |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025448014 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/21/2024 |  |  |  |  | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 70-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient on impella cp support turned on impella side in the night for an extended period and impella became lodged behind the mitral valve apparatus. ectopy was noted and attempts were made to dislodge with repositioning but unsuccessful. it was further reported that there was oozing around sheath pot low, and the patient was taken off heparin. the decision was made to withdraw care and the patient expired. the relation between the patient death and adverse events/malfunction remains unknown. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-11487 was provided in sections b2, b5, h1, and h6. | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. based on clinical description, the root cause of positioning issue was most likely use related. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. as the necessary information was not provided, the root cause of the vt/vf (ectopy) was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26379 | 1220648-2024-11488 | 19366834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19366834&pc=OZD | 04/21/2022 | 05/21/2024 |  | 05/21/2024 |  |  | 04/21/2022 | 05/21/2024 |  |  | 01/31/2024 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187802 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/21/2022 |  |  |  |  | The complainant reported a 71 year old female patient was implanted with an impella cp for mechanical circulatory support. the patient was transferred to (b)(6) hospital where the patient coded and return of spontaneous circulation was obtained. the patient was intubated and brought to the cardiovascular lab. while on support, the patient had severe calcified bilateral iliac disease at the bifurcation that precluded placement of the pump. the patient arrested again and return of spontaneous circulation was obtained. due to the instability of the patient and patient anatomy, the procedure was aborted and the pump removed. the patient was escalated to a venoarterial extracorporeal membrane oxygenation (va ecmo). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26380 | 1220648-2024-11489 | 19366959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19366959&pc=OZD | 04/22/2022 | 05/21/2024 |  | 05/21/2024 |  |  | 04/22/2022 | 05/21/2024 |  |  | 09/30/2022 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP | [5 empty] | US |  | 1505023 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/22/2022 |  |  |  |  | The complainant reported that a 91 year old patient with a chronic total occlusion of the left anterior descending artery (lad) decompensated after the initial balloon was implanted with an impella cp for mechanical circulatory support. the patient was presented with severe limb ischemia. the lab was unable to perform antegrade perfusion; therefore, the impella was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
26381 | 1035166-2024-00056 | 19367148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19367148&pc=DYB | 04/22/2024 | 06/11/2024 |  | 05/21/2024 |  |  | 04/22/2024 | 05/21/2024 |  |  | 01/01/2025 |  | 05/21/2024 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR.INC |  |  | US |  | DP-14846 |  | 0052-3021 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 56 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Deformation | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that patient had cardiac arrest at home and brought to the hosp. patient had return of spontaneous circulation(rosc). echo showed ejection fraction of 10%.impella rp was placed in the left femoral and notable kink was seen in the sheath in the femoral-iliac anatomical region. 027 impella guidewire was advanced through dual lumen pulmonary artery (pa) catheter to hold position, then pa catheter advanced successfully through this kink in sheath. pa catheter to left pa and 027 wire was buried very deep in left pa. rp was advanced over wire and to left pa. the impella wire was removed and rp started at p2. impella stopped alarm occurred and motor current high on aic. the impella was restarted at p2 successfully. then slow gradual increase in p level until reaching p9 with flows of 4-4.1 lpm. the case continued with no further similar alarms, only suction alarms. the case was completed. the patient was very sick and was being coded during our procedure. cause of death would be cardiac arrest and physician did not say its related to device. | [5 empty] | No product was provided for analysis. procedure completed successfully with the same device. the patient was very sick and was being coded during our procedure. according to physician cause of death would be cardiac arrest. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. no product was returned to oscor inc. for evaluation; however, a complaint notification was provided. without the return of the actual device in question for evaluation, oscor inc. is unable to determine the exact cause for this incident. at this time, it is not possible to assign a definitive root cause for the event as reported. the complaint is non-verifiable as the product was not returned for evaluation. in conclusion, based on the information available at this time it cannot be confirmed that the product failed to meet requirements. no corrections or corrective actions will be taken. device was not returned for analysis and review of the manufacturing documents did not reveal any discrepancies that would have contributed to this event. this complaint is non-verifiable. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
26382 | 1220648-2024-11490 | 19367162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19367162&pc=OZD | 04/22/2022 | 05/21/2024 |  | 05/21/2024 |  |  | 04/22/2022 | 05/21/2024 |  |  | 12/31/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176480 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/22/2022 |  |  |  |  | The complainant reported that a 57 year old male patient was implanted with an impella cp for mechanical circulatory support. while on support there were suction alarms due to low pump flow. the patient was successfully weaned and the impella device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26383 | 1220648-2024-11491 | 19368019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19368019&pc=OZD | 04/29/2022 | 05/20/2025 |  | 05/21/2024 |  |  | 04/29/2022 | 05/21/2024 |  |  | 09/30/2023 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022137557 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2022 |  |  |  |  | The complainant reported a 68 year old patient was implanted with an impella cp for mechanical circulatory support. during implantation, the physician attempted to advance the impella cp through the iliac artery, there was a loss of wire control as the patient was coding and being emergently intubated. the wire pulled back and dissected the artery. the insertion attempt was aborted, and the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26384 | 1220648-2024-11492 | 19368032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19368032&pc=OZD | 04/30/2022 | 05/21/2024 |  | 05/21/2024 |  |  | 04/30/2022 | 05/21/2024 |  |  | 01/31/2024 |  | 05/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190561 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/30/2022 |  |  |  |  | The complainant reported a 65 year old male patient with a high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. after post-dilation of the stents, a perforation was identified. the patient was taped and coded. return of spontaneous circulation (rosc) was achieved and covered stents were placed. hemolysis was noted and the impella cp position was confirmed. the patient was successfully weaned and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
26385 | 1220648-2024-11372 | 19350213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19350213&pc=OZD | 05/13/2024 | 07/12/2024 |  | 05/20/2024 |  |  | 05/13/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025473868 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The user facility reported a 67-year-old male with high risk of percutaneous coronary intervention was attempted to be implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 passed through the subclavian artery but could not advance into the aorta. after several attempts to advance the impella 5.5, the medical staff aborted the procedure. the medical staff discussed further options, including a decision to implant an impella cp via the right axillary artery via graft. the patient¿s outcome remains unknown. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related. the doctors were unable to place the pump due to vascular anatomy as per the clinical description. b.7 revised as this information was inadvertently omitted from manufacturer device report number 1220648-2024-11372. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11372 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since manufacturer device report number 1220648-2024-11372 was submitted. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use during impella. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
26386 | 1220648-2024-11373 | 19350311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19350311&pc=OZD | 05/14/2024 | 06/28/2024 |  | 05/20/2024 |  |  | 05/14/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103521 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The user facility reported that the automated impella controller (aic) had a battery failure alarm after turning on the console. the console failed the manual reboot. no patient harm or involvement was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported automated impella controller (aic) battery failure alarm has been completed. the console was returned for evaluation. upon inspection, there was a battery failure alarm immediately after the initial boot-up. there was also a "battery level low (50%) - alarm". the console was rebooted four more times and the battery failure alarm was persistent. this confirmed the reported failure mode. when the console was turned on, the user performed a force shutdown by toggling the battery switch to off. upon restart, the log recorded "unexpected controller shutdown - alarm", due to force shutdown. the batttest showed cell imbalance in cell 3 of battery 2 . also, the battery 2 lcd was flickering and partially blank. the front panel optical connector was contaminated as received, and upon cleaning unable to remove the debris, which is unrelated to the reported failure mode. the root cause of the battery failure alarm was cell imbalance in cell 3 of battery 2. a.3 e.2 and e.3were revised from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures.
26387 | 1220648-2024-11374 | 19350540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19350540&pc=OZD | 05/14/2024 | 07/08/2024 |  | 05/20/2024 |  |  | 05/14/2024 | 05/20/2024 |  |  | 11/30/2025 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024428588 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted an impella 5.5 device for support coming off bypass. the impella pump was placed with difficulty. the team was successful with the impella 5.5 via an .018 v18 guidewire. the pump was placed, but the pump had low flows and suction. this prompted the decision to explant and place a new impella 5.5 pump. it was noted upon explant, a kink was visible above the motor of the catheter. there was no report of harm to the patient. | [5 empty] | The investigation for the product damage has been completed. the impella 5.5 pump was returned. upon visual inspection, no biomaterial or damaged impeller blades were observed. cannula kink was found during inspection. data logs were reviewed, and suction alarm were observed with low flows during the entire case. the cause of the low pump flow issue was due to a kinked cannula per return product investigation and clinical review. added-h6 med dev prob code 1248, h6 impact code 4629, h6 component code 925. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
26388 | 1220648-2024-11375 | 19350595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19350595&pc=OZD | 03/25/2022 | 05/20/2024 |  | 05/20/2024 |  |  | 03/25/2022 | 05/20/2024 |  |  | 12/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP | [5 empty] | US |  | 2023172110 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/25/2022 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced decreased urine output and hematuria. the patient was started on dialysis through explant of the impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26389 | 1220648-2024-11376 | 19350901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19350901&pc=OZD | 06/19/2022 | 05/20/2024 |  | 05/20/2024 |  |  | 06/19/2022 | 05/20/2024 |  |  | 08/31/2024 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022128773 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/19/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella 5.5 support due to acute myocardial infarction /cardiogenic shock. while placing the impella, the patient had multiple runs of ventricular tachycardia and was shocked several times to resolve. the patient remained on support and the pump was later explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26390 | 1220648-2024-11377 | 19351000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351000&pc=OZD | 04/15/2022 | 05/20/2024 |  | 05/20/2024 |  |  | 04/15/2022 | 05/20/2024 |  |  | 01/31/2024 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187700 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/15/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, towards the end of the procedure the patient began to cough rather violently causing the previous groin stick to bleed and form a goofball size hematoma. the patient was treated with two perclose sutures and one angioseal, resulting in hemostasis. the patient remained on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).
26391 | 1220648-2024-11378 | 19351052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351052&pc=OZD | 04/15/2022 | 05/19/2025 |  | 05/20/2024 |  |  | 04/15/2022 | 05/20/2024 |  |  | 10/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022138505 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that when changing the body position, the indwelling position became shallower, and only the tip of the pigtail was within the left ventricle. due to the risk of position adjustment and the fact that hemodynamic status was maintained, the impella was pulled out to the aorta (ao) and surgically removed. the patient¿s retaining ring was said to be closed. the patient remained on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26392 | 1220648-2024-11380 | 19351076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351076&pc=OZD | 02/11/2022 | 03/26/2025 |  | 05/20/2024 |  |  | 02/11/2022 | 05/20/2024 |  |  | 07/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110568 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/11/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient showed signs of bleeding from all access sites and mouth/nose. heparin was withheld due to bleeding concern and impella pump was explanted. the patient condition was reported as stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26393 | 1220648-2024-11352 | 19351106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351106&pc=OZD | 04/20/2024 | 06/11/2024 |  | 05/20/2024 |  |  | 04/20/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485089 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Retroperitoneal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2024 |  |  |  |  | The user facility reported a 82-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was taken for a computed tomography (ct) scan due to low hemoglobin (hgb). the team suspected a retroperitoneal bleed. three units of blood were given. the ct scan found no bleeding. hgb drawn shortly after resembled baseline hgb, so the team suspected low value was an error. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the vascular damage - peripheral vessel issue has been completed since the original report was submitted. the device was not returned for investigation according to the clinical details, 3 units of blood were administered after noting retroperitoneal bleeding. the cause of the vascular damage issue was not determined since sufficient clinical information was not provided and unknown relation to impella. in accordance with updated procedures, sections d6 have been revised from the initial submission.
26394 | 1220648-2024-11382 | 19351118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351118&pc=OZD | 03/26/2022 | 03/27/2025 |  | 05/20/2024 |  |  | 03/26/2022 | 05/20/2024 |  |  | 12/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174938 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2022 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced urine was pink tinged, and progressed to black urine. the patient received fluids and echocardiography was performed to reposition the impella in the ventricle. during echocardiography, the medical team found what appeared to be a large clot on the impella inlet. it was also reported that the patient's right foot was cold. the plan of care did not alter, with continuing patient on impella support until successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related events are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26395 | 1220648-2024-11383 | 19351140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351140&pc=OZD | 11/12/2021 | 03/26/2025 |  | 05/20/2024 |  |  | 11/12/2021 | 05/20/2024 |  |  | 08/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022125539 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 96 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, a hematoma was noted. the hematoma was expelled and impella was removed at bedside. hemostasis was achieved. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Additional information provided in g1, h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26396 | 1220648-2024-11385 | 19351163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351163&pc=OZD | 12/15/2021 | 03/26/2025 |  | 05/20/2024 |  |  | 12/15/2021 | 05/20/2024 |  |  | 03/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022069504 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2021 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced an infection at the groin access site and the patient's urine output was bloody. staff performed an echocardiogram and attempted to reposition the device, but this was difficult due to the patient's anatomy and did not resolve the issue. the pump power level could not be increased due to continuous episodes of ventricular tachycardia (vt). it is unknown if good pump position was confirmed, if the impella positioning or hemolysis were resolved, or whether the impella device was removed specifically due to the reported issues. however, it was reported that the patient was later weaned from the device and survived these procedures. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26397 | 1220648-2024-11379 | 19351169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351169&pc=OZD | 01/15/2022 | 05/19/2025 |  | 05/20/2024 |  |  | 01/15/2022 | 05/20/2024 |  |  | 09/30/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022136671 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/14/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was dislocated into the aorta after care maneuver and there were several unsuccessful attempts performed to reposition impella. patient was hemodynamically stable with va echocardiogram support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26398 | 1220648-2024-11381 | 19351175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351175&pc=OZD | 08/02/2022 | 05/20/2024 |  | 05/20/2024 |  |  | 08/02/2022 | 05/20/2024 |  |  | 07/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2022111221 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced thrombocytopenia. the patient also experienced bleeding and a hematoma that continued after impella removal. the patient required a blood transfusion. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26399 | 1220648-2024-11386 | 19351191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351191&pc=OZD | 01/18/2022 | 05/20/2024 |  | 05/20/2024 |  |  | 01/18/2022 | 05/20/2024 |  |  | 10/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151348 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/18/2022 |  |  |  |  | The user facility reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced bleeding from the peel away sheath, this complication was managed by peeling away and repositioning the sheath. blood continued to ooze around the repositioning sheath, the patient received 5 units of red blood cells and controlled by a fem stop. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26400 | 1220648-2024-11388 | 19351376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351376&pc=OZD | 12/16/2021 | 05/20/2024 |  | 05/20/2024 |  |  | 12/16/2021 | 05/20/2024 |  |  | 09/30/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022135372 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/16/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the patient experienced hematuria and hemodynamic instability. the patient went to the operating room device removal by cutdown with repair of the femoral artery. following the procedure, the patient was reportedly hemodynamically stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26401 | 1220648-2024-11390 | 19351452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351452&pc=OZD | 01/18/2022 | 05/20/2024 |  | 05/20/2024 |  |  | 01/18/2022 | 05/20/2024 |  |  | 09/30/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140878 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/18/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced bleeding at the access site, a sheath was placed, and heparin was placed on hold. the nurse noted that the patient is doing well and is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26402 | 1220648-2024-11393 | 19351522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351522&pc=OZD | 01/04/2022 | 05/20/2024 |  | 05/20/2024 |  |  | 01/04/2022 | 05/20/2024 |  |  | 10/30/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022136227 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/04/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was attempted to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella sheath was inserted without any issues; however, the impella wire came back and the physician attempted to advance impella without the wire three times. the impella prolapsed and bent back. the physician was unable to straighten impella back out. it was pulled back to the sheath and then the impella broke off at inlet cage. the case was aborted. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26403 | 1220648-2024-11392 | 19351525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351525&pc=OZD | 01/03/2022 | 05/13/2025 |  | 05/20/2024 |  |  | 01/03/2022 | 05/20/2024 |  |  | 09/30/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144085 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2022 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that volume given as the patients left ventricular end-diastolic pressure was 15. the pump placement was checked and the pump was repositioned. suction alarms occurred at p2. the impella device was removed and flushed with sterile water. no clot was noted. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. correction : section h1: the type of reportable event in the initial report should have been malfunction, which was inadvertently reported incorrectly in the initial report.
26404 | 1220648-2024-11387 | 19351526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351526&pc=OZD | 11/15/2021 | 05/20/2024 |  | 05/20/2024 |  |  | 11/15/2021 | 05/20/2024 |  |  | 06/30/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108627 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Decrease in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/15/2021 |  |  |  |  | The complainant reported the 68 year old male patient was on impella cp for hemodynamic support. while on support, there was bleeding at cannulation/impella groin sites, lost of all native pulsatility, and suction alarms were noted. systemic/purge heparin was put on hold due to bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). this report is being filed as part of a retrospective review of historical records.
26405 | 1220648-2024-11389 | 19351608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351608&pc=OZD | 12/20/2021 | 05/20/2024 |  | 05/20/2024 |  |  | 12/20/2021 | 05/20/2024 |  |  | 10/31/2021 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022134823 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/20/2021 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the impella was repositioned to address hematuria and suction events. additionally, it was reported that the patient experienced staph bacteremia which caused postponement of ventricular assist device. all indwelling lines were removed/replaced and the impella cp was explanted. no further information was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
26406 | 1220648-2024-11391 | 19351625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351625&pc=OZD | 11/26/2021 | 05/20/2024 |  | 05/20/2024 |  |  | 11/26/2021 | 05/20/2024 |  |  | 08/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022130403 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/26/2021 |  |  |  |  | The complainant reported that a 70 year old male patient needing mechanical circulatory support. impella cp was placed without incidence in right femoral artery. during support, it was reported that the patient lost distal pulses in the right lower extremity. an external bypass was created by 6f retrograde access in left femoral artery and 5f antegrade access in the right superficial femoral artery. distal right extremity pulses present with doppler. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26407 | 1220648-2024-11394 | 19351632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351632&pc=OZD | 01/19/2022 | 05/20/2024 |  | 05/20/2024 |  |  | 01/19/2022 | 05/20/2024 |  |  | 09/30/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142513 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/19/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had a huge hematoma and an enlarged scrotum. the impella was removed, and ultrasound of the leg confirmed some peripheral flow. patient received 4 units of blood. physician felt that the bleeding had stopped and did not feel that further intervention was necessary. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26408 | 1220648-2024-11395 | 19351685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351685&pc=OZD | 11/30/2021 | 05/20/2024 |  | 05/20/2024 |  |  | 11/30/2021 | 05/20/2024 |  |  | 07/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022111912 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/30/2021 |  |  |  |  | The complainant reported that while the 73 year old male patient was on impella cp support, there was oozing and a hematoma. forward sutures were placed, manual pressure was held, perclose tightened and additional suture was added. site dressed using stat seal of entry noted, reinforced with folded gauze. impella site with slow oozing, worsened after turning patient. reinforced to limit patient movement. knee immobilizer ordered. manual pressure held in icu x 40 min. angle of entry reinforced with additional gauze. re-evaluation by ic fellow and oozing had slowed. plan to closely monitor and hold pressure if bleeding re-occurs. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). this report is being filed as part of a retrospective review of historical records.
26409 | 1220648-2024-11400 | 19351951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351951&pc=OZD | 01/04/2022 | 05/20/2024 |  | 05/20/2024 |  |  | 01/04/2022 | 05/20/2024 |  |  | 02/28/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022047036 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/04/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced an episode of atrial fibrillation after the device was implanted. the physician ordered electrophysiology consultation. weaning of the device was successful. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26410 | 1220648-2024-11399 | 19351972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19351972&pc=OZD | 11/30/2021 | 05/20/2024 |  | 05/20/2024 |  |  | 11/30/2021 | 05/20/2024 |  |  | 09/30/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022135437 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/30/2021 |  |  |  |  | The complainant reported that the 62 year old female patient was on impella cp. while on support, it was reported known pad (peripheral artery disease) bilateral lower extremity and no palpable pedal pulses. physician aware and consulting surgery for impella and lower extremity management. patient remains on ventilation. diminished distal pedal pulses. the patient also had dark urine output, concerns for hemolysis. echo was performed at bedside. impella measured 4.7 from av annulus to inlet cage. to prevent further damage, medical intervention repositioning the impella to pull back the device 8cm. echo confirmed appropriate placement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26411 | 1220648-2024-11402 | 19352052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19352052&pc=OZD | 01/06/2022 | 05/20/2024 |  | 05/20/2024 |  |  | 01/06/2022 | 05/20/2024 |  |  | 09/30/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142121 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/06/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis and black colored urine. the impella position was verified. the pump speed level was decreased to p4 in an effort to resolve the issue. after the speed reduction, the patient was reportedly stable and urine output was abundant. weaning was successful. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26412 | 1220648-2024-11401 | 19352085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19352085&pc=OZD | 01/21/2022 | 05/20/2024 |  | 05/20/2024 |  |  | 01/21/2022 | 05/20/2024 |  |  | 10/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152434 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 01/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had some blood loss dropping hemoglobin to 6.5. it was suspected that the bleeding was from groin site oozing. no suction alarms, just retrograde flow alarms occurred. manual pressure was held, and angle stick matched with extra gauze. bleeding from site was controlled and two units packed red blood cells were given. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26413 | 1220648-2024-11405 | 19352098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19352098&pc=OZD | 12/09/2021 | 03/26/2025 |  | 05/20/2024 |  |  | 12/09/2021 | 05/20/2024 |  |  | 09/30/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022138819 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/09/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient had high international normalized ratio and low to non-existent clotting factors. the surgeon removed the impella through sheath and then isolated and removed the sheath with limited blood loss. a clot existed in the peel-away sheath after 24 hours in the patient. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26414 | 1220648-2024-11403 | 19352262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19352262&pc=OZD | 01/14/2022 | 05/20/2024 |  | 05/20/2024 |  |  | 01/14/2022 | 05/20/2024 |  |  | 09/30/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022136745 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/14/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced steady oozing at the impella site. gentle traction on percloses seemed to mitigate oozing. oozing returned after the patient was transferred to the unit. a figure 8 suture was placed above the access site. the site was reported to be dry and soft after suture placement. additionally, during support the patient's hemoglobin was 9.6. the patient was administered packed red blood cells. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26415 | 1220648-2024-11404 | 19352272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19352272&pc=OZD | 11/30/2021 | 05/20/2024 |  | 05/20/2024 |  |  | 11/30/2021 | 05/20/2024 |  |  | 09/30/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140361 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/30/2021 |  |  |  |  | The complainant reported the 51 year old male patient was on impella cp support due to experiencing an arrest at gym while running on treadmill. while on support, suction events, and hematoma were noted. all heparin was held due to hematoma. it was also noted that patient pulled and explanted pump themselves. nurse applied manual pressure for an hour with hemostasis. patient received 3 units of blood and fluids. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) :¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). this report is being filed as part of a retrospective review of historical records.
26416 | 1220648-2024-11406 | 19352309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19352309&pc=OZD | 02/08/2022 | 05/20/2024 |  | 05/20/2024 |  |  | 02/08/2022 | 05/20/2024 |  |  | 07/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022118898 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/08/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support after experiencing ventricular tachycardia. while on support, the patient¿s urine became red and plasma free hemoglobin was ordered as a result. patient was on continuous renal replacement therapy and volume was being removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
26417 | 1220648-2024-11407 | 19352523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19352523&pc=OZD | 02/09/2022 | 05/22/2025 |  | 05/20/2024 |  |  | 02/09/2022 | 05/20/2024 |  |  | 11/30/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166524 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced dissection at groin access site post sheath removal. complication was managed with dissection at access site-ballooned and angiojet. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26418 | 1220648-2024-11408 | 19352742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19352742&pc=OZD | 03/28/2022 | 04/17/2025 |  | 05/20/2024 |  |  | 03/28/2022 | 05/20/2024 |  |  | 12/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163374 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was being implanted with the impella cp, a 100 degree kink formed in the 14f x 25 cm abiomed sheath placed at the right femoral artery. the 14fx25 cm sheath was removed and a 7f x 25 cm terumo sheath was used instead with a successful outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26419 | 1220648-2024-11409 | 19352813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19352813&pc=OZD | 02/16/2022 | 03/26/2025 |  | 05/20/2024 |  |  | 02/16/2022 | 05/20/2024 |  |  | 10/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157024 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a minor hematoma was noticed on impella site and was able to be almost resolved. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event (access site bleeding) are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26420 | 1220648-2024-11411 | 19353137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19353137&pc=OZD | 04/07/2022 | 03/26/2025 |  | 05/20/2024 |  |  | 04/07/2022 | 05/20/2024 |  |  | 01/31/2024 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187157 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella access site and two figure-eight sutures were placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26421 | 1220648-2024-11412 | 19353181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19353181&pc=OZD | 05/02/2024 | 07/09/2024 |  | 05/20/2024 |  |  | 05/02/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025455032 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the female patient experienced bleeding from the puncture site. a ct scan was performed and the patient was administered 2 packs of red blood cells and volume. the bleeding stopped after the puncture angle was adjusted. the patient is noted to be stable, and hemostasis was achieved with manual pressure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was determined to be use issue because the issue was resolved and the bleeding was stopped by adjusting the puncture angle according to best practice. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the following sections: section: general patient care considerations. section: entering cardiac output. section: potential adverse events (united states). added h6 component code 925 g1-corrected mfr site name and address.
26422 | 1220648-2024-11414 | 19353298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19353298&pc=OZD | 03/20/2022 | 05/20/2024 |  | 05/20/2024 |  |  | 03/20/2022 | 05/20/2024 |  |  | 06/30/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022134226 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/20/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was some bleeding which was managed with quick clot heparin was dropped, and light pressure dressing in addition to our standard dressing best practice, angle match with forward tension sutures, which resolved the ooze immediately. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26423 | 1220648-2024-11410 | 19353303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19353303&pc=OZD | 04/06/2022 | 03/25/2025 |  | 05/20/2024 |  |  | 04/06/2022 | 05/20/2024 |  |  | 11/30/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162664 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2022 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that upon explantation of the impella cp, the patient experienced a bleed back from left femoral artery. two perclose devices were used to achieve hemostasis, but a post procedural angiogram showed no distal flow. the common femoral artery was ballooned and the popliteal was ballooned and stented. the physician was unsure if the perclose was cinched down too tightly, if there was a clot, or if the lack of distal flow was due to the patient's chronic distal disease. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26424 | 1220648-2024-11415 | 19353319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19353319&pc=OZD | 03/21/2022 | 05/20/2024 |  | 05/20/2024 |  |  | 03/21/2022 | 05/20/2024 |  |  | 08/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022125255 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/21/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there was oozing at the insertion site. a supplemental stitch placed and angle matching. femstop with belt pressure only. heparin was also held, and blood was given due to bleeding continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26425 | 1220648-2024-11413 | 19353337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19353337&pc=PYX | 06/16/2016 | 07/15/2024 |  | 05/20/2024 |  |  | 05/02/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | PHYSICIAN ASSISTANT | Foreign | Literature | Health Professional | Company representation |  |  | 37 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | On 2024-05-02 complaints team forwarded publication entitled "percutaneous right ventricular assist device in sepsis due to right ventricular failure and pulmonary hypertension" out of odense that details rp support on 37yo female who on day 2 of support had hemolysis. "on the second day after initiation of the impella rp, jaundice appeared with serological evidence of hemolysis. this is a recognized event related to the impella drive, and the highest level of bilirubin, 285 lmol/l, occurred on day four. continuous renal replacement therapy had been started before deployment of the impella rp, to correct volume overload, and dialysis was continued for 13 days. | [5 empty] | The investigation for the hemolysis has been completed. neither the impella nor the data logs were returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
26426 | 1220648-2024-11416 | 19353802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19353802&pc=OZD | 04/14/2022 | 03/25/2025 |  | 05/20/2024 |  |  | 04/14/2022 | 05/20/2024 |  |  | 01/31/2024 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184600 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2022 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced limb ischemia due poor femoral caliber with disease. as the patient was stable and already weaned to p2, the decision was made to explant the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26427 | 1220648-2024-11418 | 19353893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19353893&pc=OZD | 05/03/2024 | 07/15/2024 |  | 05/20/2024 |  |  | 05/03/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025478960 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Insert | Failure to Deliver | Patient-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during implant procedure for 80-year-old male patient, the impella cp was unable to be advanced after 10cm in the vessel due to excessive calcification. the vascular situation on the other side was noted to be even worse. therefore, the procedure was aborted. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue is determined to be patient condition due to the calcification of the vessels that resulted in unsuccessful implantation of pump. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high risk pci in the following sections: section: potential adverse events (united states) section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
26428 | 1220648-2024-11419 | 19353955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19353955&pc=OZD | 04/14/2022 | 05/19/2025 |  | 05/20/2024 |  |  | 04/14/2022 | 05/20/2024 |  |  | 01/31/2024 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190683 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction and cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. it was reported that during repositioning of the impella after insertion, the impella popped across the aortic valve and wouldn't re-advance. the medical team removed the impella cp and replaced it with a new impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26429 | 1220648-2024-11420 | 19354011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19354011&pc=OZD | 05/09/2024 | 06/19/2024 |  | 05/20/2024 |  |  | 05/09/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1274262 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Female |  |  | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella aic didn't recognize the purge disk during the third day of support. the purge cassette was changed but the problem persisted. at this time the patient was ready to be weaned and impella was removed without any adverse event to the patient. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
26430 | 1220648-2024-11424 | 19354041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19354041&pc=OZD | 05/14/2024 | 07/17/2024 |  | 05/20/2024 |  |  | 05/14/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1233943 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | No patient involvement: the complainant reported that the aic had a battery failure. the aic was removed from service. | [5 empty] | The investigation for the reported console (aic) battery failure issue has been completed. data logs were reviewed and found that battery failure issues were confirmed. there were multiple instances of battery failure alarms (transport connection blown/missing) (event set #360) and battery failure alarms (low voltage) [v < 12v] (event set #350) found on the reported occurrence date. results of running batttest on telnet revealed that ¿fu¿ was missing on battery 2 and that vcell3 in battery 2 had a voltage of 0. the console was returned for evaluation and during testing, the battery failure issue was able to be reproduced. therefore, the root cause of the battery failure alarm was due to an internal battery cell imbalance (found in vcell3 of battery 2). added- d9 device return date d4-corrected model number. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26431 | 1220648-2024-11425 | 19354089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19354089&pc=OZD | 04/14/2022 | 03/25/2025 |  | 05/20/2024 |  |  | 04/14/2022 | 05/20/2024 |  |  | 12/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178611 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2022 |  |  |  |  | The complainant reported a patient in non-st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced a large hematoma that the medical team used pressure to treat. while the patient was being treated for the hematoma, the impella was accidentally pulled out of the ventricle into the aorta. the impella was removed, flushed, and re-inserted without difficulty. the patient remained on impella support until successfully weaned and the impella explanted. after the impella was explanted, a small thrombus was found at the tip of 14fr sheath. the impella 14fr was removed, manual pressure was held on the wires with an 8fr sheath over 1 wire. the common femoral artery (cfa) was observed to have no flow, and a thrombus was found with an angiogram at the vessel distal to the insertion site. the 8fr sheath was removed and replaced with a 14fr short sheath. a penumbra 7fr aspiration device was used for thrombus removal but was unsuccessful. the physician switched to a 12fr penumbra and was able to successfully remove thrombus and restore flow down cfa. manual compression held for 45 mins with hemostasis achieved, and good pulses in leg. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26432 | 1220648-2024-11423 | 19354110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19354110&pc=OZD | 05/12/2024 | 07/16/2024 |  | 05/20/2024 |  |  | 05/12/2024 | 05/20/2024 |  |  | 09/30/2025 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397111 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. when the peel away was being removed, the team saw the impella cp lost its left ventricle position. the cp was unable to be re-delivered and the patient decompensated and was being given chest compressions. the team then placed a new pump. plan is to send for emergent coronary artery bypass grafting. | [5 empty] | The investigation for the positioning issue has been completed. data logs were reviewed which confirmed pump was in aortic area (about 20 minutes into the case). pump then was stopped manually and disconnected from the console. product was not returned for review. based on clinical description and data analysis, the root cause of positioning issue was most likely use related. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: delivery issues impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
26433 | 1220648-2024-11422 | 19354130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19354130&pc=OZD | 05/11/2024 | 07/15/2024 |  | 05/20/2024 |  |  | 05/11/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025494698 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male | 204 KG | Non Hispanic | American Indian Or Alaskan Native | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | Death |  |  |  |  | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2024 |  |  |  |  | The user facility reported a 23-year-old male patient was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced access site bleeding requiring four units of packed red blood cells. the patient also developed sepsis and expired while on support. | [5 empty] | The investigation for the access site bleed and infection has been completed. the impella device was not returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the access site bleeding was determined to be patient condition because has patient does have an escalating white count of 83,000 currently and a ck of 16,000. and patient continues to have persistently high lactic, acidosis and increasing vasoactive medication requirements. the root cause of the infection could not be determined without additional clinical details. added-h6 component code 925 f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: table 3.2 impella catheter components: the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
26434 | 1220648-2024-11426 | 19354189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19354189&pc=OZD | 05/14/2024 | 07/08/2024 |  | 05/20/2024 |  |  | 05/14/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440976 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Malposition of Device | Infusion or Flow Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, there were positioning issues. the pump was found too deep and hung on the papillary muscle. staff attempted to reposition the pump back to readjust the position and the pump was pulled back past the aortic valve. the physician was unable to advance the impella past the valve. when attempting to advance, the impella stopped due to high motor current. the impella would not restart after multiple attempts were made. the impella was removed. the physician would manage patient with inotropes and vaso-pressors if needed. | [5 empty] | The investigation for the pump stop, saturated flow, and positioning issues has been completed. the impella cp pump was not returned. data logs were reviewed and saturated flows were observed at p-9 and p-8 since the start of the case. high motor current pump stop was observed after 52 hours of support. saturated flows were observed since case start with no motor current spike and diastolic waveforms indicating likely gradual ingestion of biomaterial. there was no heparin or sodium bicarbonate purge usage for first 2.5 hours during support and the last 2.5 hours before pump stop. pump stop was found shortly after the position in aorta alarm events. as per clinical pump was unable to be restarted after multiple attempts. biomaterial was observed in pump outflow after explant. no other abnormalities were seen. the cause of the positioning issue was most likely use related per clinical review. the cause of the pump stop issue was most likely biomaterial per log and clinical review. f6/f8 should have been blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
26435 | 1220648-2024-11427 | 19354261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19354261&pc=OZD | 05/15/2024 | 06/19/2024 |  | 05/20/2024 |  |  | 05/15/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1320564 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | Device Fell | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2024 |  |  |  |  | The complainant reported that an automated impella controller (aic) was dropped, resulting in potential physical damage. the aic was removed from service, and no patient harm has been reported. | [5 empty] | The investigation for the reported console (aic) damage has been completed. the console was returned for evaluation. upon evaluation, the console's front optical connector was found broken off. data logs were not relevant to this case. based upon clinical details, the cause of the broken optical connector was the console being dropped. added- d9 device return date. added- h6 med dev prob code 2588 h8-usage of device corrected to reuse. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
26436 | 1220648-2024-11428 | 19354283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19354283&pc=OZD | 05/15/2024 | 05/27/2025 |  | 05/20/2024 |  |  | 05/15/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438148 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Hispanic |  | Malfunction | Yes | Yes | Failure to Deliver | Device Sensing Problem | Difficult to Advance | Material Twisted/Bent |  |  |  | Hemolysis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | 05/15/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump. upon the initial attempt, the pump was to pass an obstruction at the level of the innominate artery and the guidewire subsequently kinked. the wire was replaced, but the pump once again was unable to pass the presumed obstruction. once removed, a kink was noted in the cannula and at the catheter distal to the motor housing. the pump was replaced and a new pump was utilized for implant. there was no patient harm. | Additional information was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that the patient experienced hemolysis. the event was stated to have a probably relationship to the impella device. two units of blood were ultimately given to treat the condition. | It was later reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that the sensor on the device was also not working after the failed attempts to place the pump. this factor also led to the decision to replace the device. |  |  |  | The investigation for the product damage and delivery issues has been completed. the product was not returned for investigation. a boston scientific 0.018 guide wire was utilized after a kink was noted on the initially placed guide wire. the doctor was not able to advance the pump due to concern about the innominate artery stenosis. direct insertion was attempted using a new pump after the reported kink on the cannula and catheter. the cause of the delivery issue and subsequent product damage was most likely patient condition related to disease vessels, based on given clinical details. d4-corrected model number f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. should additional information become available, a supplemental mdr will be filed. the following updates were made to manufacturer device report 1220648-2024-11428. b1 updated to both adverse event and product malfunction. b2 updated to reflect new information received. b5 updated with new information received. h1 type of report updated. h6 updated to reflect new failure mode added. instructions for use hemolysis: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ d4 model number was inadvertently reported incorrectly on manufacturer device report 1220648-2024-11428. f6/f8-should have been left blank on manufacturer device report 1220648-2024-11428. g1 reporting contact email revised on manufacturer device report 1220648-2024-11428 in accordance with updated procedures. g2 report source; study was missing from manufacturer device report 1220648-2024-11428. | After further review the impella pump was not implanted, thus the following corrections were done to manufacturer device report number 1220648-2024-11428-2, as the patient would not experience hemolysis from pump that was not implanted. b1 revised to reflect product malfunction only. b2 updated to blank. h1 type of report updated to malfunction. h6 health effect - clinical code and health effect - impact code were erroneously entered on manufacturer device report number 1220648-2024-11428-2. instructions for use hemolysis were erroneously entered on manufacturer device report number 1220648-2024-11428-2. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. the investigation for the product damage and delivery issues has been completed. the product was not returned for investigation. a boston scientific 0.018 guide wire was utilized after a kink was noted on the initially placed guide wire. the doctor was not able to advance the pump due to concern about the innominate artery stenosis. direct insertion was attempted using a new pump after the reported kink on the cannula and catheter. the cause of the delivery issue and subsequent product damage was most likely patient condition related to disease vessels, based on given clinical details.
26437 | 1220648-2024-11421 | 19354334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19354334&pc=OZD | 05/01/2024 | 03/07/2025 |  | 05/20/2024 |  |  | 05/10/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1417541 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility reported a 80-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in japan reported the automated impella controller (aic) had impella connect issues. the hospital's wi-fi in the icu was weak and periodically disconnected. the case may have been offline for more than 4 hours and then the wi-fi reconnected and impella connect judged the case to be new. it was at this point that a data streaming failure occurred, where the case was displayed on the portal site, but there was no digital data populating the case tile and no prospective logging of alerts / alarms. the data streaming failure was confirmed by admin diagnostic page. no known patient harm or injury. | [5 empty] | The impella device was received from the customer on (b)(6)2024. and therefore,an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) ¿ other has been completed. based on clinical details and data logs, the root cause of aic not shutting down correctly was identified as a software anomaly. d2 common device name revised on manufacturer device report 1220648-2024-11421 in accordance with updated procedures. d4 model number corrected and revised on manufacturer device report 1220648-2024-11421. f6/f8-should have been left blank on manufacturer device report 1220648-2024-11421. g1 reporting contact email revised on manufacturer device report 1220648-2024-11421 in accordance with updated procedures.
26438 | 1220648-2024-11431 | 19354616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19354616&pc=OZD | 04/29/2024 | 09/22/2025 |  | 05/20/2024 |  |  | 04/29/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025463762 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/29/2024 |  |  |  |  | It was reported that a 5.5 pump was placed to support a patient needing mechanical circulatory support. the patient experienced access site bleeding that was resolved by pausing heparin. the patient also experienced thrombocytopenia that was treated by removing heparin and adding one unit packed red blood cells. the patient also experienced ventricular tachycardia that was resolved with shock. the patient continued on support until the device was successfully weaned. | The complainant also reported that the ongoing pump support after 7 days lost optical signal. the pump remained on for support despite this without any intervention. |  |  |  |  | The investigation for the reported access site bleed, thrombocytopenia, and ventricular tachycardia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed, thrombocytopenia, and ventricular tachycardia could not be determined. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. as noted in the impella ifu: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned. data logs were reviewed and placement signal not reliable observed on 04/30/24. the optical 5s parameters signal-to-noise ratio (snr) and contrast was found to be decreasing while the gain and lamp increased. the optical signal issue resolved after 31 hours. no other abnormalities were found. the root cause of the optical signal failure was most likely an air gap between the automated impella controller and the patient cable plug. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-11431. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should not have had entries in the initial report. g1 reporting contact email updated.
26439 | 1220648-2024-11429 | 19354861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19354861&pc=OZD | 05/06/2024 | 05/29/2024 |  | 05/20/2024 |  |  | 05/06/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453243 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2024 |  |  |  |  | The complainant reported a patient in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced gastrointestinal bleeding that was resolved by withholding heparin, several units of packed red blood cells, and embolization. the patient continued on impella support as a bridge to transplant. subsequently, the patient had the heart transplant and the impella was removed without issues. | [5 empty] | The investigation for the vascular damage has been completed. the impella was not returned for evaluation. additionally, clinical details was insufficient to determine the root cause. the cause of vascular damage - peripheral vessel cannot be determined since the complaint device not returned for investigation, insufficient clinical data provided and the relation to impella unknown. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).
26440 | 1220648-2024-11430 | 19354941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19354941&pc=OZD | 04/13/2024 | 06/04/2024 |  | 05/20/2024 |  |  | 04/13/2024 | 05/20/2024 |  |  |  |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451040 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2024 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the patient experienced ventricular tachycardia, while on impella support, which was resolved by reducing impella flow. the patient continued on support until the device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the access site bleeding ¿ minor and vf/vt/af/at (ventricular fibrillation) issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was most likely patient movement, upon transfer to bed catheter bumped by blanket and subsequent bleeding noted to site as per the clinical description. as the necessary information was not provided, the root cause of the vt/vf was not determined. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
26441 | 1220648-2024-11432 | 19355589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19355589&pc=OZD | 02/13/2022 | 10/29/2025 |  | 05/20/2024 |  |  | 02/13/2022 | 05/20/2024 |  |  | 03/31/2022 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1478518 |  | 0048-3075-AU |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient's right leg was noted to be cooler with weakened pulses. the impella cp was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | F6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-11432 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-11432 in accordance with updated procedures. h6 component code and investigation conclusions revised from the initial submission in accordance with updated procedures.
26442 | 1220648-2024-11433 | 19356272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19356272&pc=OZD | 04/14/2022 | 03/25/2025 |  | 05/20/2024 |  |  | 04/14/2022 | 05/20/2024 |  |  | 12/31/2023 |  | 05/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180775 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. during delivery of the device, the 14 f sheath was advanced but could not advance all of the way up. it was recommended to use a stiff wire; however, the motor could not exit the sheath to attempt stiff wire when the cp was attempted. multiple attempts and techniques were used with no success. the procedure was aborted due to patient anatomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26443 | 1220648-2024-11365 | 19348704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19348704&pc=OZD | 12/09/2021 | 03/26/2025 |  | 05/19/2024 |  |  | 12/09/2021 | 05/19/2024 |  |  | 09/30/2023 |  | 05/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140912 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/09/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed bleeding at the impella access site. an extra suture was placed around the blue hub which stopped the oozing. additionally, it was noted that the pump had been explanted due to issues with hemolysis, hemoglobin was down to 7. the patient was successfully weaned from support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26444 | 1220648-2024-11366 | 19348715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19348715&pc=OZD | 11/12/2021 | 03/26/2025 |  | 05/19/2024 |  |  | 11/12/2021 | 05/19/2024 |  |  | 07/31/2023 |  | 05/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114052 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed site oozing and a large hematoma. fem stop was placed to control bleeding. blood products were also given. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26445 | 1220648-2024-11367 | 19348721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19348721&pc=OZD | 11/20/2021 | 03/26/2025 |  | 05/19/2024 |  |  | 11/20/2021 | 05/19/2024 |  |  | 07/31/2023 |  | 05/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121778 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/20/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed oozing from the impella site at the time of insertion. systemic heparin at 600u/hr which was up from 400 from the morning with activated clotting time (act) 131. the physician reported the hemoglobin started to drop and as a result, the patient was successfully weaned off support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26446 | 1220648-2024-11368 | 19348722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19348722&pc=OZD | 11/21/2021 | 03/26/2025 |  | 05/19/2024 |  |  | 11/21/2021 | 05/19/2024 |  |  | 06/30/2023 |  | 05/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022098184 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 11/21/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported the patient had bleeding from the access site. a mattress suture was placed to resolve the bleeding. the patient was successfully weaned off of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26447 | 1220648-2024-11369 | 19348726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19348726&pc=OZD | 11/21/2021 | 03/26/2025 |  | 05/19/2024 |  |  | 11/21/2021 | 05/19/2024 |  |  | 09/30/2023 |  | 05/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022135399 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a rapidly growing hematoma at groin site was noted. manual pressure held by staff for 15 minutes. heparin and integrilin were stopped, and protamine given. site bleeding was controlled by femstop. it was decided to do a quick wean of the impella and explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26448 | 1220648-2024-11370 | 19348736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19348736&pc=OZD | 11/22/2021 | 03/26/2025 |  | 05/19/2024 |  |  | 11/22/2021 | 05/19/2024 |  |  | 06/30/2023 |  | 05/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022098179 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced pink/red urine. it was suggested that p levels could be decreased in order to reduce hemolysis. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: hemolysis impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26449 | 1220648-2024-11371 | 19348742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19348742&pc=OZD | 11/30/2021 | 03/26/2025 |  | 05/19/2024 |  |  | 11/30/2021 | 05/19/2024 |  |  | 08/31/2023 |  | 05/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022131988 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/30/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, hemolysis was noted. the device was repositioned and possible causes of hemolysis were addressed. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: hemolysis: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26450 | 1220648-2024-11300 | 19339714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19339714&pc=OZD | 02/17/2022 | 03/27/2025 |  | 05/17/2024 |  |  | 02/17/2022 | 05/17/2024 |  |  | 10/31/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152141 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, a duplex doppler of the impella-side leg showed no pulse on the right side, though the foot was warm to the touch. no intervention was reported, and the patient remained on impella cp support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26451 | 1220648-2024-11302 | 19340083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340083&pc=OZD | 05/08/2024 | 01/30/2025 |  | 05/17/2024 |  |  | 05/08/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1132985 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The complainant reported a patient undergoing coronary artery bypass was implanted with an impella 5.5 device for mechanical circulatory support. after priming the device, the automated impella controller (aic) displayed "impella defective" alarm. the aic may have an eeprom writing/reading issue and was swapped with a new aic. the aic was sent back for investigation. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the issue with the console (aic) has been completed. console imc logs were reviewed which confirm that; " impella defective (error reading eeprom) - alarm was issued. eeprom file of the pump 4982153 was reviewed but no problem found with it. the aic was returned for evaluation. the failure was not reproducible on boot up, inspection of the internal circuitry of the console revealed no observable anomaly, the console was able to run a known good pump and purge line for a few hours without a problem. the root cause for the impella defective alarm could not be determine due to inability to reproduce. corrections to the initial mfr report: f6/f8 should have been left blank. g1 mdr reporting contact email
26452 | 1220648-2024-11303 | 19340143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340143&pc=OZD | 05/09/2024 | 06/28/2024 |  | 05/17/2024 |  |  | 05/09/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025458070 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella cp was prepped and primed for insertion, but the patient's anatomy would not allow placement. both iliac arteries were poor targets for impella and the bilateral axilla were diseased and torturous. the patient was prepped with the impella introducer however the procedure was aborted. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related. the impella implant was unsuccessful as both iliac arteries were poor targets for impella and the bilateral axilla were diseased and torturous as per the clinical description. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
26453 | 1220648-2024-11304 | 19340176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340176&pc=OZD | 04/17/2024 | 05/17/2024 |  | 05/17/2024 |  |  | 04/17/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025449037 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/17/2024 |  |  |  |  | The user facility reported a 56-year old male with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the patient experienced access site bleeding and a hematoma while on impella cp support. the patient was administered one unit of packed red blood cells. the doctor placed a mattress suture and confirmed the groin was stable. the patient eventually expired. there is not have enough information to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not returned, and the investigation has been completed. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved by adding an additional mattress suture. it was also noted that activated clotting time at the time of bleed in the intensive care unit was 200, which was above the range limit. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26454 | 1220648-2024-11305 | 19340208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340208&pc=OZD | 11/17/2023 | 05/17/2024 |  | 05/17/2024 |  |  | 11/17/2023 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1079314 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 05/10/2024 |  |  |  |  | The us hospital had an aic (automated impella controller) that failed to purge. this was noted by the biomed department. there was no noted patient use at the time. no harm or injury noted. the investigation into the aic determined the purge disc/flag to be the issue. | [5 empty] | The impella device was received from the customer on (b)(6) 2023. data analysis imc logs revealed that the aic (automated impella controller) repeatedly issued a purge disc not detected alarm (event set #402) after the purge cassette was inserted on multiple occasions. during operation, the purge disc is inserted along with the purge cassette. the console¿s reporting that the purge disc was not detected indicated an issue with the purge disc detection flag. see attachment ¿ic1079_imc.png¿ for more information. device analysis console ic1079 was tested after arriving in fs. the purge disc flag was confirmed to be broken. see attachment ¿ic1079_broken_flag.png¿. before returning this console back to the customer, fs was instructed to replace the purge pressure sensor/retainer ((b)(6) or 5500-0222), validate sensor accuracy via section 12 of the pm procedure (0042-7309), and perform a full functional check on the console and optical system (0042-7316).
26455 | 1220648-2024-11306 | 19340213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340213&pc=OZD | 02/17/2022 | 03/27/2025 |  | 05/17/2024 |  |  | 02/17/2022 | 05/17/2024 |  |  | 03/31/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022058378 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had excessive bleeding at the impella access site due to high levels of posterior calcification which resulted in two perclose device failures. pressure was held and 1 perclose device was successfully placed, followed by 8fr angioseal. femstop was placed to hold pressure until total hemostasis was achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26456 | 1220648-2024-11307 | 19340248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340248&pc=OZD | 05/10/2024 | 07/10/2024 |  | 05/17/2024 |  |  | 05/10/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025469316 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility reported a 51-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that there was no resolution of suction alarm as the physician tried to reposition the impella cp several times unsuccessfully. it was further reported that the cp p-level would not go over 2. the patient hemodynamics remained unchanged. mds determined that in the absence of critical alarms the explant would be the next morning. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation for the positioning issue has been completed. data logs were confirmed the failure mode and clinical narrative. logs showed after the pump was started, the pump position was in the aorta and low flow occurred that triggered impella position in aorta, auto suction response and suction alarms. the pump position occurred in 1 minute and resolved but low flow still remained to the rest of the case. no motor current spikes observed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. the instructions for use referenced in the initial report is for the impella cp with smartassist system in following sections. section: use of echocardiography for positioning of the impella catheter. section: understanding and managing impella catheter position alarms d4-corrected . model number. f6/f8-should have been left blank in the initial report.
26457 | 1220648-2024-11308 | 19340313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340313&pc=OZD | 05/10/2024 | 07/10/2024 |  | 05/17/2024 |  |  | 05/10/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488406 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having triple vessel disease. while on support, the pump started with success and the impella plug got caught under the prepping drape and was pulled by the physician. a pump failure alarm was present, and the staff attempted to restart the pump three times with no success. the pump was removed by the physician. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported pump stop has been completed. the cp pump was not returned. data logs were reviewed and no pump stop, defective alarm observed. no other abnormalities were found at the end of the case. the cause of the pump stop issue was not determined since product was not returned and due to inconclusive evidence per data log review. d4-corrected model number f6/f8 should have been left blank in the initial report. g4 pma/510(k)number missing.
26458 | 1220648-2024-11309 | 19340348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340348&pc=OZD | 05/11/2024 | 06/24/2024 |  | 05/17/2024 |  |  | 05/11/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025473883 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2024 |  |  |  |  | The user facility reported a 52-year-old male was implanted with an impella 5.5 as a bridge to transplant. it was reported that on day 8 of patient support, the pump came out of position and was unable to be redelivered across valve under echocardiogram guidance. the pump was explanted, and the pump was replaced by an intra-aortic balloon pump (iabp). the patient was reported as stable. | [5 empty] | The investigation into positioning issues has been completed since the original report was submitted. the product was not returned for investigation. based on clinical description and data analysis, the root cause of positioning issue was most likely patient movement. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26459 | 1220648-2024-11310 | 19340356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340356&pc=OZD | 04/20/2024 | 06/05/2024 |  | 05/17/2024 |  |  | 04/20/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448747 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2024 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the team placed an impella cp and extracorporeal membrane oxygenation (ecmo) when the patient went into pulseless electrical activity arrest and coded in the cardiac catheterization lab. the patient was transferred from community to a tertiary center for continued care. when at the second center, the patient was seen to have a left ventricle perforation, presumed from the impella cp, and was taken to the operating room for patch repair. additionally, the team planned for escalation to the impella 5.5 pump. the team also observed the left sided ecmo access site to have a fistula and aneurysm. the patient survived for four hours on the impella 5.5 pump when the patient expired with care withdrawn. there is not have enough information to exclude the impella as an associated factor in the patient¿s outcome. | [5 empty] | The investigation for the ventricle perforation has been completed. no product was returned for investigation. the root cause of the ventricle perforation was not determined since product was not returned and there is a lack of clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿
26460 | 1220648-2024-11311 | 19340402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340402&pc=OZD | 02/18/2022 | 03/27/2025 |  | 05/17/2024 |  |  | 02/18/2022 | 05/17/2024 |  |  | 10/31/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156227 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced limb ischemia. an external femoral to femoral bypass was performed and a plan was made to escalate the patient to venoarterial extracorporeal membrane oxygenation (va ecmo). a clot was found in the 6f-6f bypass system, and the sheath was exchanged for 8f arrowflex and retrograde to 7f x 25 cm pinnacle. the impella was explanted and exchanged for va ecmo. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26461 | 1220648-2024-11312 | 19340516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340516&pc=OZD | 12/08/2021 | 03/27/2025 |  | 05/17/2024 |  |  | 12/08/2021 | 05/17/2024 |  |  | 09/30/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022143207 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2021 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella cp device for mechanical circulatory support. while placing the impella cp in the patient, the sheath was unable to pass either groin after multiple attempts due to the patient's tortuous anatomy. the case was then aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26462 | 1220648-2024-11313 | 19340521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340521&pc=OZD | 12/15/2021 | 04/08/2025 |  | 05/17/2024 |  |  | 12/15/2021 | 05/17/2024 |  |  | 07/31/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121666 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was noted that there was oozing from the blue suture hub/touhy connection, the staff redressed and no further oozing noted. it was also noted that the patient had a large amount of blood loss, and the nurse noted that it sounds like it is caused by loss of appropriate angle and resolves with dressing change to angle match. the patient¿s hemoglobin was 7.7 so two (2) packed red blood cells were given. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26463 | 1220648-2024-11314 | 19340691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340691&pc=OZD | 12/16/2021 | 05/17/2024 |  | 05/17/2024 |  |  | 12/16/2021 | 05/17/2024 |  |  | 03/31/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022059598 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/16/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. while attempting to advance the impella through the sheath, the patient¿s extremely tortuous iliac and aorta anatomy made it very difficult to advance the sheath and caused the 25-centimeter sheath to kink in two locations. the physician was able to successfully advance the sheath across the aortic valve, the left anterior descending stent was placed without complications and the hrpci procedure was successfully completed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: warnings & cautions: cautions as noted in the impella ifu ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, eval this report is being filed as part of a retrospective review of historical records.
26464 | 1220648-2024-11315 | 19340867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340867&pc=OZD | 04/20/2024 | 05/23/2024 |  | 05/17/2024 |  |  | 04/20/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466321 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2024 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had no flow distal to the impella sheath following impella cp implant. a femoral-femoral bypass was performed to treat the condition and support with the implanted pump was maintained. the decision was made to withdraw care and the patient expired. per medical safety officer review, there is not have enough information to exclude the impella as an associated factor in the patient¿s outcome. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the pump was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26465 | 1220648-2024-11317 | 19340975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340975&pc=OZD | 02/19/2022 | 03/27/2025 |  | 05/17/2024 |  |  | 02/19/2022 | 05/17/2024 |  |  | 09/30/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022147397 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/19/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced hematuria and continuous suction waveforms from the impella after volume loss from vomiting. it was recommended to the medical team to confirm position of impella with echocardiography. the patient was transfused with one unit of blood and the p level of the impella was decreased. patient was stable and the medical team decided to wean and explant the impella. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26466 | 1220648-2024-11316 | 19340984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19340984&pc=OZD | 12/18/2021 | 05/17/2024 |  | 05/17/2024 |  |  | 12/16/2021 | 05/17/2024 |  |  | 10/31/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2022148972 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/16/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced atrial fibrillation, a bolus of amiodarone antiarrhythmic was given and an iv drip started, the patient was then reported to be experiencing normal sinus rhythm. it was also reported the patient had signs of hemolysis in the urine, plans were made to start weaning the patient. the patient was weaned down to a pressure level of p4, then the physicians increased the pressure back to p8 to optimize for upcoming coronary artery bypass graft (cabg). the patient remained on impella support pending the upcoming cabg procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26467 | 1220648-2024-11318 | 19341071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341071&pc=OZD | 04/20/2024 | 06/05/2024 |  | 05/17/2024 |  |  | 04/20/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454484 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Unspecified Infection | Tachycardia | [10 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support for escalation of care from an impella cp device. during support, the patient had an embolic stroke and the patient's mental status was monitored. additionally, the patient also had a low-grade fever. antibiotics were started and systemic heparin was withheld. furthermore, the patient had runs of ventricular tachycardia. the patient was shocked eight times and an echocardiogram showed the pump was seven centimeters from the aortic valve. the pump was repositioned. the decision was made to withdraw care. support was withdrawn and the patient passed shortly after. there is no association between impella use and the patient's outcome. patient's peri-procedural condition was critical. | [5 empty] | The investigation into the reported issues has been completed since the original report was submitted. the pump was not returned for investigation. because evidence has not been provided, the root cause of the infection could not be determined. as this clinical information was not provided, the true root cause of the stroke/cerebrovascular accident and the ventricular tachycardia/ventricular fibrillation could not be determined. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ h.6 added code 925 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26468 | 1220648-2024-11319 | 19341140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341140&pc=OZD | 12/09/2021 | 03/27/2025 |  | 05/17/2024 |  |  | 12/09/2021 | 05/17/2024 |  |  | 09/30/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022138077 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/09/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during placement, the patient experienced ventricular fibrillation (vf) and advanced cardiac life support protocol was initiated. the patient also experienced bleeding from the mouth and oozing around groin sites. no information was provided reporting medical or surgical intervention. the patient continued on support until successful weaning. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26469 | 1220648-2024-11320 | 19341220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341220&pc=OZD | 12/29/2021 | 03/27/2025 |  | 05/17/2024 |  |  | 12/29/2021 | 05/17/2024 |  |  | 08/31/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022128029 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/29/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during removal of the impella cp, perclose was attempted two (2) times but failed, requiring surgical cutdown of femoral artery for closure. additionally, upon a successful explant and switching to impella 5.5, hemolysis immediately ceased, and urine was clear. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26470 | 1220648-2024-11321 | 19341229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341229&pc=OZD | 12/10/2021 | 03/27/2025 |  | 05/17/2024 |  |  | 12/10/2021 | 05/17/2024 |  |  | 09/30/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022132743 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Female |  |  |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped due to a thrombus formation. the pump was exchanged. no additional information was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event (thrombus) are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smart assist system instructions for use for the related event (pump stop) are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions. warnings, cautions & precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. g1 manufacturing site facility name.
26471 | 1220648-2024-11322 | 19341287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341287&pc=OZD | 12/13/2021 | 05/17/2024 |  | 05/17/2024 |  |  | 12/13/2021 | 05/17/2024 |  |  | 08/31/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022130189 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/13/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported the patient experienced oozing at the access site. after 45 minutes of manual pressure, the site was still bleeding. the staff re-dressed the access site and sutures were placed. additionally, the patient experienced rapid atrial fibrillation with two failed cardioversions. the staff administered dobutamine and vasopressin. the patient also spiked a low-grade temp at 101.3. the staff administered antibiotics. information was provided confirming that the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
26472 | 1220648-2024-11323 | 19341305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341305&pc=OZD | 05/12/2024 | 11/22/2024 |  | 05/17/2024 |  |  | 05/13/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1102032 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The user facility reported a 49-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported that while providing support to a patient implanted with an impella 5.5 the aic (automated impella controller) generated multiple alarms regarding purge flow increasing and decreasing. the cassette was changed with no resolution. the aic was replaced to resolve the issue. | [5 empty] | The impella device was received from the customer on 15-may-2024 and therefore,¿an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into automated impella controller (aic) - purge issue ¿ other has been completed. the device was returned but the failure mode was not reproduced. a purge and pump were run overnight but had no purge flow alarms. the was no device issue found, root cause not determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11323 d4 model number f6/f8-should have been left blank.
26473 | 1220648-2024-11324 | 19341372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341372&pc=OZD | 05/13/2024 | 07/18/2024 |  | 05/17/2024 |  |  | 05/13/2024 | 05/17/2024 |  |  | 08/31/2025 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025443544 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The user facility reported a 54-year-old female patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the cp was placed but the pump had low flows and was explanted and replaced after just 22 minutes of support. the 2nd pump was successfully placed. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. pump was not returned for investigation. data logs were reviewed and low flow was observed since the case start. suction alarms were observed with auto suction response immediately after start and low flows during the whole case at all p-levels. no motor current spike was observed. also, placement signal was low, declining likely due to deteriorating patient condition. per data log review and clinical information provided, the root cause of the low pump flow issue was most likely patient condition related. f6/f8 should have been left blank in the initial report.
26474 | 1220648-2024-11325 | 19341405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341405&pc=OZD | 05/13/2024 | 11/19/2024 |  | 05/17/2024 |  |  | 05/14/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1791478 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/14/2024 |  |  |  |  | No patient involvement: complainant in japan reported the automated impella controller (aic) was newly purchased and the staff was setting up the aic when there was a system self check failed change console immediately error. even after rebooting, the problem did not improve and a repair request was made. | [5 empty] | The investigation for the console (aic) system self check issue has been completed. the data logs from the console confirm that there was a system self check failure event that was issued on the date of occurrence. from the logs, its was observed that the failure occurred when the console rebooted itself after initially booting up with numerous epm communication errors. the console was received for investigation. the failure was reproducible upon start up of the console. the console's internal circuitry inspected visually for any non conformance with no observable anomaly. there was nominal power being supplied from the pbm to the impellatronics board as well as the epc. isolative testing by swapping with a known good impellatronics board confirmed the original to be defective. the root cause of the system self check failure was a defective impellatronics board. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26475 | 1220648-2024-11326 | 19341442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341442&pc=OZD | 03/01/2022 | 03/27/2025 |  | 05/17/2024 |  |  | 03/01/2022 | 05/17/2024 |  |  | 09/30/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022151101 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/01/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on impella support, the patient experienced a gastrointestinal bleed, and purge related heparin administration was halted in an attempt to control the issue. it was not specified how this issue was resolved, but the patient was kept on impella support for several days afterward and survived. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26476 | 1220648-2024-11327 | 19341445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341445&pc=OZD | 05/14/2024 | 01/16/2025 |  | 05/17/2024 |  |  | 05/14/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474740 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | Hemodynamic instability | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/14/2024 |  |  |  |  | The complainant reported a 57-year-old male patient with elective placement was implanted with an impella 5.5 for mechanical circulatory support. while on 5.5 support, the patient became hypotensive and tachypneic. the patient's chest was reopened, revealing cardiac tamponade. medical staff identified a bleeding source at the anastomosis of the 10mm platinum graft, where it appeared the suture line gave way. a partial clamp was placed, and the suture line was reinforced and tested. a large clot burden was also removed, and the patient¿s chest was left open with a sump drain in place. medical staff administered three units of packed red blood cells as well as platelets. the patient remains on support. | It was further reported that care was withdrawn from the patient and the patient subsequently expired. the patient's outcome and it was determined that there is not enough information to associate use of the device with the outcome. | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured mediastinum and right chest blood clot with a "probable" relationship to the impella device used: the bleeding in chest requiring blood clot evacuation from mediastinum and right chest¿. it was reported that there was bleeding in chest post cabg. had to reopen chest to evacuate clot for the second time. the also endured subarachnoid hemorrhage with a "possible " relationship to the impella device used: ¿patient had cerebral and cerebellar edema and subarachnoid hemorrhage and the patient ultimately expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the reported access site bleed and thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and thrombus could not be determined. added- b5-mso review, d6b f6/f8 should have been left blank in the initial report. | B5 additional information added h1 health effect - clinical code added to reflect additional information b2 was updated to death in accordance with updated procedures on manufacturer device report 1220648-2024-11327. b2 date of death was added in accordance with updated procedures on manufacturer device report 1220648-2024-11327. g1 reporting contact email revised on manufacturer device report 1220648-2024-11327 in accordance with updated procedures. g2 report source; study was inadvertently omitted from manufacturer device report 1220648-2024-11327. h1 type of reportable event updated to death in accordance with updated procedures on manufacturer device report 1220648-2024-11327 h6 health impact code 1802 added in accordance with updated procedures on manufacturer device report 1220648-2024-11327.
26477 | 1220648-2024-11328 | 19341478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341478&pc=OZD | 05/14/2024 | 06/27/2024 |  | 05/17/2024 |  |  | 05/14/2024 | 05/17/2024 |  |  | 10/31/2025 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024422641 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient indicated as a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support and experienced an access site complication. at the end of the impella assisted pci, a hematoma formed at the access site. the decision was made to wean the patient from the impella cp pump. the patient was hemodynamically stable during weaning, and the next decision was made to remove the pump. hemostasis was achieved with one perclose and one angioseal with no blood products were administered. | [5 empty] | The investigation into the access site bleeding ¿ major issue been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26478 | 1220648-2024-11329 | 19341607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341607&pc=PYX | 01/02/2022 | 03/27/2025 |  | 05/17/2024 |  |  | 01/02/2022 | 05/17/2024 |  |  | 11/30/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022151972 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/02/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella rp device for mechanical circulatory support. during the attempt to insert the first impella rp, the device kinked in the main body so a new impella rp was used. the second impella rp was unable to navigate into position, so the device was removed and reinserted after rewire. the device was still unable to pass so there were several attempts that included a buddy wire as well as switching wire to rpa before ultimately deciding to abort the impella rp procedure. | [5 empty] | The impella rp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26479 | 1220648-2024-11330 | 19341757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341757&pc=OZD | 12/14/2021 | 05/27/2025 |  | 05/17/2024 |  |  | 12/14/2021 | 05/17/2024 |  |  | 06/30/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022102880 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Ischemia | Perforation of Vessels | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the peel-away was removed and then hematoma was noted at impella access site. manual pressure improved but the physician opted to remove the impella. the physician reinserted peel-away sheath, and reinserted impella. hematoma resolved. systemic heparin was removed due to oozing from the right femoral artery arterial sheath (removed) as well as from the impella site. it was also noted that the patient had suspected gastrointestinal bleeding (gib). patient was supposed to have a ct scan to identify source of bleeding, however, results are unknown. it was recommended to cut heparin dose in purge or eliminating it if concern over heparin-induced thrombocytopenia (hit) /bleeding continued. patient received one (1) unit of packed red blood cells. the medical staff was unable to doppler the lower left extremity (lle) and an ultrasound was ordered. able to doppler patient pulses but not dp. patient weaning of levophed. (lle) (impella) pulses were dopplerable. patient off all pressors. additionally, it was reported that there was concern over anoxic brain injury but the patient, prior to transfer to u of m, was responding to commands, nodding. patient successfully weaned off support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26480 | 1220648-2024-11331 | 19341762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341762&pc=OZD | 01/03/2022 | 03/27/2025 |  | 05/17/2024 |  |  | 01/03/2022 | 05/17/2024 |  |  | 09/30/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140949 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, patient experienced moderate oozing on both groins. pressure was held by sheath pullers to resolve the oozing. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26481 | 1220648-2024-11332 | 19341866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341866&pc=OZD | 12/14/2021 | 05/17/2024 |  | 05/17/2024 |  |  | 12/14/2021 | 05/17/2024 |  |  | 09/30/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022141689 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/14/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support thrombus was identified when the cannula clotted off in the unit. the patient went to lab and a trans-septal percutaneous ventricular assist device (pvad) was placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26482 | 1220648-2024-11333 | 19341916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341916&pc=OZD | 03/02/2022 | 03/27/2025 |  | 05/17/2024 |  |  | 03/02/2022 | 05/17/2024 |  |  | 10/31/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157920 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced atrial flutter with a pulmonary artery diastolic pressure in the 130's mmhg and went into ventricular tachycardia (vt) and decompensated. the impella p level was increased from p2 to p8 and the patient was taken to the operating room and placed on venoarterial extracorporeal membrane. the patient progressed with no further indications of heart block or ectopy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26483 | 1220648-2024-11334 | 19341955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19341955&pc=OZD | 12/15/2021 | 05/17/2024 |  | 05/17/2024 |  |  | 12/15/2021 | 05/17/2024 |  |  | 09/30/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022143010 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/15/2021 |  |  |  |  | The user facility reported a 68-year-old patient needing mechanical circulatory support. it was reported that the patient was on impella cp and while on support the pump was minimally out of position. the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26484 | 1220648-2024-11335 | 19342044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19342044&pc=OZD | 05/10/2024 | 06/10/2024 |  | 05/17/2024 |  |  | 05/10/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1153140 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility's biomedical engineer reported that the automated impella controller (aic) had the selection wheel ripped off and the protective cover was snapped off and the inside was loose. there was no patient involvement. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console (aic) knob issues has been completed. data logs were reviewed and found that after the pump was used for 24 hours, it was moved to the spare console (b)(6) due to the front panel optical connector might have been broken. the aic was returned for evaluation and was visually inspected which confirmed that the front panel optical connector was broken (the blue receptacle was loose inside the aic) and the rotating knob was missing. no issue was found in the rotatory knob switch (both press and rotate functions working as intended). the root cause of the broken front panel optical connector was most likely due to a use issue. the root cause of the missing rotating knob was most likely due to a use issue.
26485 | 1220648-2024-11336 | 19342132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19342132&pc=OZD | 01/03/2022 | 03/27/2025 |  | 05/17/2024 |  |  | 01/03/2022 | 05/17/2024 |  |  | 09/30/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022147291 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, a clot was found in the patient's left ventricle. the impella cp was weaned and explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26486 | 1220648-2024-11337 | 19342207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19342207&pc=OZD | 12/14/2021 | 03/27/2025 |  | 05/17/2024 |  |  | 12/14/2021 | 05/17/2024 |  |  | 08/31/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127378 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2021 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, patient developed hematoma at impella site, therefore blood transfusion was done. also, the patient had hemolysis and thrombocytopenia. no additional information was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related events are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26487 | 1220648-2024-11338 | 19342421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19342421&pc=OZD | 01/11/2022 | 03/27/2025 |  | 05/17/2024 |  |  | 01/11/2022 | 05/17/2024 |  |  | 10/31/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151806 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/12/2022 |  |  |  |  | The complainant reported a patient with heart failure/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was found to have heparin-induced thrombocytopenia, so heparin was removed from the purge bag. it was also reported that the following day, the patient became febrile with a systemic vascular resistance of 500, raising concerns of sepsis. patient notes did not confirm whether the impella was implicated as the source of the sepsis. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26488 | 1220648-2024-11339 | 19342547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19342547&pc=OZD | 05/09/2024 | 06/27/2024 |  | 05/17/2024 |  |  | 05/09/2024 | 05/17/2024 |  |  | 07/31/2025 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442202 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The user facility reported an 80-year-old male with a high risk of percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. it was reported that the patient also had a 14 french sheath placed. the medical staff planned for a transfer to a tertiary center for a coronary artery bypass graft but discovered oozing at the groin access site. the staff then placed a femoral stop and later sutured to resolve the issue. the patient remained on support despite the access site issue and drop in hemoglobin. the medical staff eventually explanted the impella cp for the coronary artery bypass graft procedure. the patient survived the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved by sutures.
26489 | 1220648-2024-11340 | 19342655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19342655&pc=OZD | 04/19/2024 | 06/10/2024 |  | 05/17/2024 |  |  | 04/19/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454983 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported a 43-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported the patient was on impella 5.5 support when there were low pump flows noted. upon starting the pump, flows were low and 0 diastolic flows with a very negative diastolic left ventricle value were noted. the doctor decided to replace the pump due to fear of having damaged the pump by not removing the wire when the pump was started. a new impella 5.5 was used successfully. | [5 empty] | The investigation for the low pump flow has been completed. according to the clinical, immediately after beginning impella 5.5 support flows were low and the diastolic lv value was negative. an attempt was made to reposition the pump successfully with no change to the flow. the decision was then made to explant the pump and replace it with another impella 5.5 device. no product was returned for a visual inspection. data log analysis confirmed the presence of suction alarms and a high motor current pump stop shortly after beginning support. initially, the second pump serial number was given for log analysis. the correct pump serial number is (b)(6). the most likely cause of the low pump flow was biomaterial. d4-updated serial number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26490 | 1220648-2024-11341 | 19343025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19343025&pc=OZD | 10/26/2022 | 07/23/2024 |  | 05/17/2024 |  |  | 04/09/2024 | 05/17/2024 |  |  | 04/30/2024 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023226867 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 59 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | Information obtained from literature: literature number: 24003656, p22-1: title: a case of restricted mitral valve opening due to impella insertion after aortic valve replacement surgery: impella 5.5 insertion was also scheduled before weaning from cardiopulmonary bypass. on the day of surgery, the surgical procedure was completed without any problems. the physician first tried to insert the lmpella 5.5 through the right axillary artery, but it did not pass, so a 9mm graft was sewn to the ascending aorta graft. after insertion, an attempt was made to withdraw from cardiopulmonary bypass, but a decrease in systemic blood pressure and an increase in pulmonary artery pressure were observed. transesophageal ultrasound (tee) confirmed that the impella had been inserted too deeply, which interfered with the anterior leaflet of the mitral valve and restricted its release. by shallowing the impella 5.5, mitral valve mobility was achieved, and hemodynamics were temporarily improved. however, it was difficult to stop the bleeding, and the same phenomenon repeated several times, so the position was adjusted each time. lmpella 5.5 is usually inserted through the axillary artery, but in this case, it was inserted through a graft that was attached to the replacement blood vessel for the ascending aortic artery, so the angle at which impella was inserted into the aortic valve was different from normal. it is possible that interference with the anterior leaflet of the mitral valve was likely to occur. when using impella in the perioperative period, it is considered desirable to use tee, in clinical practice, keeping in mind the possibility of restricting mitral valve release. the patient is stable. | [5 empty] | The investigation into the positioning issues has been completed. the abiomed products were not returned for evaluation and data logs were not available. the root cause of the positioning issue was patient condition related due to patient's aortic valve replacement surgery that changed the insertion process. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
26491 | 1220648-2024-11342 | 19343032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19343032&pc=OZD | 05/10/2024 | 06/11/2024 |  | 05/17/2024 |  |  | 05/10/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025478359 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | Ventricular Fibrillation | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. after obtaining left ventricle access for implantation, the patient went into ventricular tachycardia/ventricular fibrillation which required three shocks to convert. return of spontaneous circulation was obtained and the impella was successfully implanted. three days after implant, the patient exhibited hemolysis and flow issues. the pump was explanted to resolve the hemolysis and a visible clot was evident. the patient was placed on venoarterial extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ hemolysis impella 5.5 with smart assist for use during cardiogenic shock with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation into the reported issues has been completed since the original report was filed. the device was returned. the pump and data logs were returned. upon product inspection, biomaterial was found on the impeller. the data logs confirmed positioning alarms and a motor current spike with saturated flows. as the necessary information was not provided, the root cause of the vt/vf and the thrombus was not determined. the root cause of hemolysis and the saturated flow was determined to be biomaterial as a clot was found on the impeller during the product inspection and a motor current spike was seen in the data logs. in accordance with updated procedures, sections d6 and f9 have been revised from the initial submission.
26492 | 1220648-2024-11343 | 19343070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19343070&pc=OZD | 05/10/2024 | 07/12/2024 |  | 05/17/2024 |  |  | 05/10/2024 | 05/17/2024 |  |  | 11/30/2025 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446059 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and after over 10 hours of support, the patient self-explanted the impella pump. the patient was returned to the cardiac catheterization lab for a new impella cp pump placement, which is now supporting. the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. data logs were not made available. based on clinical description, the root cause of positioning issue was use related as the patient self-removed the pump. d4 corrected model number.
26493 | 1220648-2024-11344 | 19343132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19343132&pc=OZD | 05/10/2024 | 07/12/2024 |  | 05/17/2024 |  |  | 05/10/2024 | 05/17/2024 |  |  | 11/30/2025 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434301 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The complainant reported a 72-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed a hematoma at the access site. medical staff attempted manual compression and further troubleshooting without success. medical staff then explanted the cp. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since the product was not returned and sufficient information was not provided in the clinical description. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11344 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since manufacturer device report number 1220648-2024-11344 was submitted. h.6 codes 1888 and 925 were reported incorrectly on manufacturer device report number 1220648-2024-11344. h.10 additional manufacturer narrative instructions for use were revised from manufacturer device report number 1220648-2024-11344 in accordance with updated procedures.
26494 | 1220648-2024-11346 | 19343196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19343196&pc=OZD | 05/13/2024 | 06/19/2024 |  | 05/17/2024 |  |  | 05/13/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1377267 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The complainant reported that the automated impella controller (aic) on startup, two messages displayed - system self check failed and change console immediately. the aic was removed from service. there was no patient involvement. | [5 empty] | The investigation into the aic - system self-check error has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. a firmware check was performed, which revealed that there was a lack of communication between the 4-in-1 and the power battery manager channel 1. the cause of the aic issue was contamination of the pbm pca due to leakage of the supercapacitors. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
26495 | 1220648-2024-11345 | 19343197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19343197&pc=OZD | 05/12/2024 | 06/27/2024 |  | 05/17/2024 |  |  | 05/12/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025492072 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2024 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and experienced multiple runs of ventricular tachycardia (vt). the physician made the decision to remove the cp and escalate to an impella 5.5. reportable due to patients vt and necessary intervention. | [5 empty] | The investigation into the vf/vt/af/at (multiple runs of vtach) issue been completed since the original report was submitted. the device was not returned for investigation. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
26496 | 1220648-2024-11347 | 19343270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19343270&pc=OZD | 04/09/2024 | 09/17/2025 |  | 05/17/2024 |  |  | 04/09/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025466986 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced ventricular tachycardia (vt) that was treated with lidocaine, atrial fibrillation that was treated with lopressor, and additional runs of vt that subsided with no specified resolution. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation for the optical signal issue has been completed. the impella 5.5 pump was returned. visually inspected and observed no kink in catheter or other damages to the pump. there was no burr or rough edges found in the return pump. optical bench 5s parameters were verified and all were found to be in a normal range. the pump passed the light continuity test and no placement signal issue alarms reproduced. no other abnormalities were observed. data logs were reviewed, contrast observed to decrease for 98 hours and recover back to normal range. all other light, lamp, gain were stable during the case. no placement signal issue alarms found. few impella position in aorta alarms found but did not look like false alarm due to placement signal monitoring not being disabled. the intermittent waveform abnormality did not reproduce with the pump and console during return product investigation. the root cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. e1 middle name was incorrectly reported on the initial manufacturer device report number 1220648-2024-11347. g1 updated with correct mdr reporting contact email. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-11347.
26497 | MW5155088 | 19343454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19343454&pc=OZD | 05/13/2024 | 05/15/2024 |  | 05/17/2024 |  |  |  | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED INC. | [6 empty] | 0048-0003 |  | No | [6 empty] | Health Professional |  |  | Voluntary report | Yes | Physician | [6 empty] | 74 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | [7 empty] | Cardiac Arrest | Exsanguination | Hypovolemic Shock | Cardiac Tamponade | Cardiac Perforation | [8 empty] | Death | [28 empty] | Initial submission | [9 empty] | An impella-assisted percutaneous coronary intervention was performed. the impella was removed after the case. two hours later, patient was found to be in tamponade with hemopericardium. pericardiocentesis was performed but she was in hemorrhagic shock. chest was opened by cardiothoracic surgery and she was found to have perforation of the posterior lateral wall which was sutured closed. unfortunately patient sustained multiple cardiac arrests during this time and we were unable to achieve spontaneous circulation.
26498 | 1220648-2024-11355 | 19343934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19343934&pc=OZD | 05/11/2024 | 07/15/2024 |  | 05/17/2024 |  |  | 05/11/2024 | 05/17/2024 |  |  | 07/31/2025 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388014 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Unknown |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient indicated as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support and experienced an access site complication and positioning issue. the impella cp pump was successfully implanted. the right femoral artery was oozing. no hematoma was noted. manual pressure was applied, angle of insertion was maintained, and two purse string sutures were used. additionally, it was noted that the pigtail of the pump was against the apex. the physician attempted to pull the impella cp pump back but could not as the inlet was in outflow track right under the valve. the physician acknowledged the recommendation that the pigtail be in a free space of the ventricle. however, the decision was made to maintain the position of the pigtail and closely monitor. | [5 empty] | The investigation for the positioning issue and access site bleed has been completed. data logs showed no positioning alarms were triggered during support. there were intermittent suction alarms on but were resolved quickly. product was not returned for review. based on clinical description and data analysis, the root cause of positioning issue was most likely use related. the root cause of access site bleeding was most likely due to anticoagulation management related as the act was 340 (normal range 160 -180) at the time the bleeding occurred. added-h6 impact code 4628. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26499 | 1220648-2024-11356 | 19344452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19344452&pc=OZD | 05/10/2024 | 06/21/2024 |  | 05/17/2024 |  |  | 05/13/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025449981 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient was on impella 5.5 support and the patient had bowel ischemia possibly secondary to an embolic phenomenon. the cause is unknown. the patient went for an exploratory lap, purulent drainage, small bowel resection and ileostomy creation. four units of blood products were provided to the patient. the patient was on systemic heparin. | [5 empty] | The investigation into the vascular damage - peripheral vessel issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the bowel ischemia was not determined since product was not returned and there is limited clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26500 | 1220648-2024-11357 | 19344933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19344933&pc=OZD | 05/13/2024 | 06/20/2024 |  | 05/17/2024 |  |  | 05/13/2024 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025463207 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced no distal pulses. the impella cp pump was placed via the 14 french sheath, but the pump was explanted due to an ischemic leg. the leg had no pulses, was removed via cutdown, and the patient was placed instead on venous-arterial extra-corporeal membrane oxygenation. the impella cp pump had only supported for five hours before it was explanted. the patient was noted to be in stable condition. | [5 empty] | The investigation into the limb ischemia - reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26501 | 1220648-2024-11358 | 19345123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19345123&pc=OZD | 04/19/2024 | 06/28/2024 |  | 05/17/2024 |  |  | 04/19/2024 | 05/17/2024 |  |  | 10/31/2025 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024409085 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 support for a 71-year-old male, the physician requested lysis therapy to the patient due to purge pressure greater than 1100 mmhg and flow values of 1 milliliter per hour. there was no reported patient harm. | [5 empty] | The investigation into the high purge pressure issue has been completed since the original report was filed. the pump was returned for investigation. biomaterial was present on the pump around the impeller blade and purge gap. clinical details noted that soon after implantation high purge pressure occurred and was resolved with tpa adding to purge. the root cause of the high purge pressure is most likely due to biomaterial as issue was resolved with tpa. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
26502 | 1220648-2024-11359 | 19345871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19345871&pc=OZD | 01/21/2022 | 03/27/2025 |  | 05/17/2024 |  |  | 01/21/2022 | 05/17/2024 |  |  | 10/31/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142136 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella access site and the amount of heparin in the purge was being reduced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26503 | 1220648-2024-11360 | 19345962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19345962&pc=OZD | 02/09/2022 | 03/27/2025 |  | 05/17/2024 |  |  | 02/09/2022 | 05/17/2024 |  |  | 11/30/2022 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021013081 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2022 |  |  |  |  | The complainant reported a patient, ethnicity unknown, noted as a very complex high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported prior to the procedure that the patient was in cardiac arrest in the intensive care unit. the patient was reportedly stabilized and the impella implanted, and a successful revascularization of the arteries was performed which was reported to have improved the patient¿s condition. however, the impella was explanted at the end of the revascularization due to the risk of leg ischemia. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. d1 brand name.
26504 | 1220648-2024-11361 | 19346104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19346104&pc=OZD | 02/09/2022 | 03/27/2025 |  | 05/17/2024 |  |  | 02/09/2022 | 05/17/2024 |  |  | 09/30/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022145838 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced no pulses in the right foot and was weaned of levophed as appropriate. it was reported that an echocardiogram revealed very poor left ventricle and ejection fraction function and good right ventricle function. additionally, suction alarms were occurring at pressure 7. a scan was scheduled to observe the flow in the right leg with consideration for femorofemoral bypass if signs of modeling occurred. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26505 | 1220648-2024-11362 | 19346170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19346170&pc=OZD | 12/04/2021 | 03/26/2025 |  | 05/17/2024 |  |  | 12/04/2021 | 05/17/2024 |  |  | 09/30/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022139741 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was reported that the patient experienced a hematoma. manual pressure was applied in addition to femostop and a sandbag for hematoma resolution. the hematoma slightly improved but was still present, with bruising in the area and bleeding from the site. two units of packed red blood cells (prbcs) and one unit of platelets were given. the sheath was secured and dressing changed. the decision was made to remove the impella due to the hematoma at the site and difficulty of removing at bedside or the cath lab. the patient was taken to the operating room for removal via cutdown. the impella was explanted without issue. per the vascular surgeon, there was no major damage to the artery that caused the bleeding, and no damage caused by the impella sheath. the artery was repaired and closed without issue. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26506 | 1220648-2024-11363 | 19346483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19346483&pc=OZD | 01/06/2022 | 03/26/2025 |  | 05/17/2024 |  |  | 01/06/2022 | 05/17/2024 |  |  | 08/31/2023 |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022130372 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, an echocardiogram revealed a left ventricle thrombus, so the physician decided to explant the device. a surgical cutdown to the right femoral artery was performed, and the impella was removed. distal pulses were found to be absent after cp removal. an angiogram showed occlusion to the right iliac artery (ia). a large thrombus was removed from the right ia and was believed to have been attached to the end of the pigtail catheter. a successful surgical closure to the right femoral artery was performed, and distal pulses returned to baseline. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26507 | 1220648-2024-11364 | 19346819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19346819&pc=OZD | 01/18/2022 | 05/17/2024 |  | 05/17/2024 |  |  | 01/18/2022 | 05/17/2024 |  |  |  |  | 05/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient was found to have a hematoma. an angiogram was performed though the re-access port and it was identified that the path of the impella sheath was through a branch off of the profunda and caused the hematoma. the physician placed a balloon for dry closure. once the accepted time had passed, the balloon was deflated, and the groin was still bleeding. surgical repair was performed, and the closure was successful. | [5 empty] | D4: the serial number, lot number, expiration date, and udi are not available at this time. h4: the manufacturing date is not available at this time. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26508 | 1220648-2024-11236 | 19329624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19329624&pc=OZD | 04/01/2022 | 05/14/2025 |  | 05/16/2024 |  |  | 04/01/2022 | 05/16/2024 |  |  | 01/31/2024 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184596 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patients activated partial thrombin time was higher than expected and systemic heparin was withheld. the 14 french sheath was found to be clotted which was supplying distal perfusion catheter. the physician placed a sheath single access style to perfuse, but this was also to found to be clotted. the patient's leg was warm and post tibial pulses were present, so the medical team monitored the patient and did not intervene further. the patient was successfully weaned and the impella cp was explanted with no further complications reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: use of impella with transcatheter aortic valves in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26509 | 1220648-2024-11235 | 19329864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19329864&pc=OZD | 04/25/2022 | 05/16/2024 |  | 05/16/2024 |  |  | 04/25/2022 | 05/16/2024 |  |  | 09/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022135323 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/25/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved with removal of the swan and site repair by the physician. the patient remained on impella support for five days and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations as noted in the impella ifu ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26510 | 1220648-2024-11237 | 19329894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19329894&pc=OZD | 12/25/2021 | 05/14/2025 |  | 05/16/2024 |  |  | 12/25/2021 | 05/16/2024 |  |  | 04/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079329 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/25/2021 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient went into ventricular tachycardia (vt) and was shocked into a normal sinus rhythm (nsr) with amiodarone and epinephrine administered. the patient also experienced access site bleeding and received multiple blood products. the bleeding was unable to be controlled, and the impella was explanted from the right femoral artery and replaced with a new impella cp that was implanted into the left femoral artery. a 14fr x 25cm sheath placed to tamponade bleeding, and the patient was moved to the intensive care unit for rest. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26511 | 1220648-2024-11238 | 19330184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330184&pc=OZD | 04/25/2024 | 06/26/2024 |  | 05/16/2024 |  |  | 04/25/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1462815 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The complainant reported that during training, the battery was unable to be loaded and both automated impella controllers (aic) displayed an error message. there was no patient involvement. | [5 empty] | The device has been returned for evaluation. upon completion of the investigation, a supplemental mdr will be filed. | The investigation into the aic - battery failure (red alarm) has been completed since the initial report was submitted. after reviewing the logs, the battery failure issue was confirmed. there were numerous instances of battery failure alarm (transport connection blown/missing) and battery comm. failure alarms found from the reported occurrence date. results of running batttest on telnet revealed that ¿fu¿ was missing on both batteries. the console was tested and battery failure alarm issue was able to be reproduced. the issue was determined to be caused by depleted batteries. the console was last disconnected from ac power on (b)(6) and was never plugged back to alternating current power since. these depleted batteries could not be charged back up because they were locked out due to low cell voltage (less than 2.3v). the root cause of this issue was the automated impella controller not being routinely charged. d.1 brand name, d.2 common device name and d.4 model number were revised from the initial submission in accordance with updated procedures. d.4 serial number was revised as previously entered incorrectly. g.1 manufacturing site facility name and address line 2 were revised as previously entered incorrectly. g.6 type of report - 30-day was inadvertently omitted from the previously submitted report. h.6 medical device problem code was revised from the initial submission in accordance with updated procedures.
26512 | 1220648-2024-11239 | 19330219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330219&pc=OZD | 05/08/2024 | 06/28/2024 |  | 05/16/2024 |  |  | 05/08/2024 | 05/16/2024 |  |  | 06/30/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024361306 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced ischemia to the lower extremities. the patient had mottled lower extremities with decreased pulse cap refill. additionally, the patient had hemolysis with frank blood in their urine. the staff tried to wean the pump and the patient was not stable. the pump was in the proper position. the team discussed possible impella rp implant. the impella cp was explanted after two days of support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation into the hemolysis and the limb ischemia ¿ reversible issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the hemolysis issue was most likely patient condition related, based on data log review. limb ischemia can occur during impella support. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
26513 | 1220648-2024-11240 | 19330223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330223&pc=OZD | 12/14/2021 | 05/14/2025 |  | 05/16/2024 |  |  | 12/14/2021 | 05/16/2024 |  |  | 09/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022146436 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2021 |  |  |  |  | The complainant reported that a patient was placed on impella cp support for high risk percutaneous coronary intervention. while on support, it was reported that the perclose failed so an angioseal was used. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: difficult to insert/remove pump through introducer. impella cp with smartassist for use during cardiogenic shock and high risk cpi section: potential adverse events (united states); ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26514 | 1220648-2024-11242 | 19330289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330289&pc=OZD | 04/30/2024 | 06/26/2024 |  | 05/16/2024 |  |  | 04/30/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2025487860 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | The complainant reported that during impella cp support, the 44-year-old male patient had puncture site that continued to bleed. the vascular surgeon "narrowed" the puncture site slightly and the issue resolved. no blood products were given and the patient is noted to be stable. | [5 empty] | The device has not been received for evaluation. upon completion of the investigation, a supplemental mdr will be submitted. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since the vascular surgeon narrowed the puncture site slightly which resolved the issue.
26515 | 1220648-2024-11243 | 19330342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330342&pc=OZD | 05/08/2024 | 07/09/2024 |  | 05/16/2024 |  |  | 05/08/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025450028 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported right before priming, the impella 5.5 device was unplugged from the automated impella controller (aic) to remove torque from catheter. when the impella 5.5. device was plugged back in, aic alarm ¿impella defective¿ displayed. troubleshooting was attempted but was unsuccessful as the alarm condition did not clear. a new impella 5.5 device was opened for use and successfully primed. there was no report of harm to the patient. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the pump did not start issue has been completed since the original report was filed. the pump was returned for investigation. the device was visually inspected, and no physical damages or abnormalities were found. data logs were reviewed and impella defective alarm was observed after pump plug connect. however, the cause of the impella defective issue was not determined since failure was unable to be reproduced with return product and due to inconclusive evidence based on data log review. d9 date device returned to manufacturer added d4 model number corrected. d6a d6b added.
26516 | 1220648-2024-11244 | 19330349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330349&pc=OZD | 01/04/2022 | 05/16/2024 |  | 05/16/2024 |  |  | 01/04/2022 | 05/16/2024 |  |  | 10/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022150776 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hematoma to the right femoral artery. the femostop was placed and the physician was waiting for partial thromboplastin time (ptt) results and hematoma control before adding heparin to the purge. the ptt was subtherapeutic and heparin was added to the purge. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26517 | 1220648-2024-11245 | 19330382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330382&pc=OZD | 12/25/2021 | 10/13/2025 |  | 05/16/2024 |  |  | 12/25/2021 | 05/16/2024 |  |  | 10/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152121 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/25/2021 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced a cold foot and lack of pulses in the lower extremities. the impella sheath was found to be occlusive to the lower leg and the physician placed a tee grad flow sheath. this did not resolve the limb ischemia, and the medical team decided the patient will be managed and potentially escalated to 5.5 or extracorporeal membrane oxygenation. it was also reported the patient experienced bleeding at the impella access site. the patient received several units of blood as a result of the bleed. the patient was transferred to another hospital, where no further information was provided. | [5 empty] | D6b unknown date of explant. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26518 | 1220648-2024-11246 | 19330384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330384&pc=OZD | 05/11/2024 | 07/12/2024 |  | 05/16/2024 |  |  | 05/11/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025455083 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 05/11/2024 |  |  |  |  | The complainant reported a 63-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with a impella cp for mechanical circulatory support. the complainant reported the patient had cardiogenic shock and had a 14fr and the cp inserted via the femoral artery. the patient developed a hematoma, which medical staff treated with pressure, blood products, sutures, and purge fluid exchange / heparin adjustment. the patient had coagulopathy due to a liver injury. the patient remains on support. | It was further reported that care was withdrawn from the patient and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. |  |  |  |  | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be anticoagulation management because the hematoma became stable after the act values were down to the therapeutic range. added-b5 mso review, d6b, h6 component code 925 f6/f8 should have been left blank in the initial report.
26519 | 1220648-2024-11248 | 19330436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330436&pc=OZD | 01/04/2022 | 05/16/2024 |  | 05/16/2024 |  |  | 01/04/2022 | 05/16/2024 |  |  | 08/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022130168 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia. the decision was made to remove the impella, and once the impella was removed hemostasis was achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26520 | 1220648-2024-11250 | 19330525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330525&pc=OZD | 12/14/2021 | 03/17/2025 |  | 05/16/2024 |  |  | 12/14/2021 | 05/16/2024 |  |  | 05/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089746 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, a hematoma was noted and a drop in blood pressure. packed red blood cells were given and vascular surgery consulted. the patient was to go to the operating room for repair vs covered stent. closure was achieved through balloon tamponade and or cutdown for vessel repair. small hematoma was also noted after the 8 french sheath inserted. pressure held and sheath upgraded to 9 french and site improved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26521 | 1220648-2024-11253 | 19330600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330600&pc=OZD | 12/25/2021 | 05/14/2025 |  | 05/16/2024 |  |  | 12/25/2021 | 05/16/2024 |  |  | 09/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140379 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/25/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during insertion of an impella cp through a 14 french 30 cm cook sheath, the impella was unable to advance. the medical team aborted the insertion. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26522 | 1220648-2024-11257 | 19330786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330786&pc=OZD | 07/18/2023 | 05/30/2024 |  | 05/16/2024 |  |  | 04/24/2024 | 05/16/2024 |  |  | 05/31/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024350381 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured access site hematoma (right axillary) with a "definite" relationship to the impella device used: ¿post impella removal, golf ball sized hematoma noted to r axillary impella site. manual pressure applied. upon return to post area, hematoma increased in size requiring repeat cath showing no evidence.¿ the patient recovered with sequelae. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
26523 | 1220648-2024-11258 | 19330787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330787&pc=PYX | 05/08/2024 | 07/01/2024 |  | 05/16/2024 |  |  | 05/08/2024 | 05/16/2024 |  |  | 08/31/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024392234 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella rp device for mechanical circulatory support who was also supported by the impella 5.5 device. the team tried all troubleshooting and wiring methods but the impella rp placement was aborted. there was no harm to the patient who remained stable. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation the root cause of the delivery issue was most likely patient condition related. the impella was unable to pass through after multiple attempts as per the clinical description. g1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h6 medical device problem code was revised in accordance with updated procedures. a new code has been added to medical device problem codes. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: rp smartassist system with automated impella controller. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
26524 | 1220648-2024-11259 | 19330830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330830&pc=OZD | 12/14/2021 | 05/14/2025 |  | 05/16/2024 |  |  | 12/14/2021 | 05/16/2024 |  |  | 09/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022136140 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | Yes | Failure to Deliver | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2021 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the physician attempted to intervene on the second obtuse marginal artery, however was unable to cross due to calcium load and incredibly acute angle of takeoff. impella weaned and removed. the physician pulled peel-away out over wire but was unable to get control of artery to tamper down perclose sutures due to very large pannus. additional staff members held pressure while physician obtained contralateral access. one (1) additional perclose suture placed. balloon tamponade performed for five (5) minutes x 3 inflations with external pressure. this resulted in appropriate hemostasis with a small hematoma present. medical intervention was not mentioned to resolve the hematoma. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding. impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26525 | 1220648-2024-11260 | 19330847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330847&pc=OZD | 01/05/2022 | 05/16/2024 |  | 05/16/2024 |  |  | 01/05/2022 | 05/16/2024 |  |  | 10/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154263 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/05/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. during support the right femoral artery (rfa) 6fr sheath kinked making it difficult to advance. it was removed and pressure was held for fifteen minutes. the device was explanted after multiple failed percloses attempts. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26526 | 1220648-2024-11247 | 19330868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330868&pc=PYX | 05/12/2024 | 07/17/2024 |  | 05/16/2024 |  |  | 05/12/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX WITH SMART ASSIST |  |  | ABIOMED, INC. |  |  | US |  | 2025479474 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2024 |  |  |  |  | The complainant reported a 51-year-old male patient with cardio myopathy was implanted with an impella rp for mechanical circulatory support. the complainant reported the rp had low pump flow and did not provide information about how the issue was resolved. no patient harm has been reported, and the patient remains on support. | The investigation identified additional failures not reported by the user. temporary pump stop and optical issues were added to the findings in the investigation. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | B.5 additional information was added to reflect new findings from the investigation. h.6 added medical device problem codes to reflect new information from the investigation. the investigation has been completed. data logs show that 88 hours into support, there was a sudden spike in motor current (mc) with a pump stop alarm. the pump was able to be restarted. at the same time, spikes to 300 mmhg, signal-to-noise ratio (snr) dropped to 0, lamp maxes at 100, contrast decreases, and gain increases, and there was a peak signal-to-noise ratio (psnr) (op) alarm. the optical sensor never recovered. several days later, there was a few ingestion signatures (mc spike, speed deviation, upwards placement signal shift) with drops in flows while at p-6. the pump passed the laser continuity test. psnr was reproduced when connected to the lab console. numerous small amounts of biomaterial deposits were observed to be circulating in the bottle when the pump was run in the lab, the borescope view of the impeller looked clean. field imaging confirmed biomaterial on the pump which was flicked off during explant. on the returned 5s parameters, the lamp was high, contrast was low, snr was very low, light was acceptable, and the cavity length was maxed out. the temporary pump stop was not reproduced when running in the lab. the low pump flow was not reproduced. the optical sensor was unable to hold pressure in the uson and was reading -3276 mmhg. imaging the optical sensor showed sensor damage. the root cause of the low pump flow issue was most likely biomaterial. heparin had been withheld prior due to a colonoscopy. the root cause of the temporary pump stop issue was not determined since the pump stop was not reproduced in the lab during testing. the root cause of the optical signal issue was most likely a damaged sensor but the exact cause of the damage was not determined. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ d.4 revised model number from manufacturer report number 1220648-2024-11247 in accordance with updated procedures. d.6b revised from manufacturer report number 1220648-2024-11247 in accordance with updated procedures. d.9 revised as this information was inadvertently omitted from manufacturer report number 1220648-2024-11247. e.1 facility name and address line 2 was previously entered incorrect on manufacturer report number 1220648-2024-11247 and were revised. f.6 and f.8 dates were reported on manufacturer report number 1220648-2024-11247 and should not have been. g.1 revised manufacturing site name and address section as previously entered incorrectly on manufacturer report number 1220648-2024-11247. h.3 device not returned to manufacturer for evaluation should of have been selected as the device was returned at the time of the report being submitted. if the device was not evaluated selection on the previous report should of been device evaluation anticipated, but not yet begun.
26527 | 1220648-2024-11249 | 19330883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330883&pc=OZD | 05/13/2024 | 07/03/2024 |  | 05/16/2024 |  |  | 05/13/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488423 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Hemorrhagic Stroke | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | Additional information states that when explanting the cp, 8mm thrombus was noted on 9fr catheter in descending aorta. clot removed completely. a ct scan was performed which revealed an area of cerebral infarction suggesting early hemorrhage. anticoagulation stopped. it was further reported that the family made a decision to withdraw care after the patient had a stroke. the patient subsequently expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. | The complainant reported a 53-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient had the cp placed via the femoral artery. on the third day of support, a limb started to become ischemic. the limb was externally bypassed, and medical staff discussed escalating the patient to an impella 5.5. the patient remains on support. |  |  |  |  | The investigation for the limb ischemia, thrombus, and stroke has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. added- b5 mso review, d6b, h6 component code 925 h6-clinical codes 4440, 4417 h6-impact code 4641. f6/f8-should have been left blank in the initial report. h5-changed to yes. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
26528 | 1220648-2024-11261 | 19330902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19330902&pc=OZD | 08/16/2023 | 06/13/2024 |  | 05/16/2024 |  |  | 04/24/2024 | 05/16/2024 |  |  | 01/31/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024308308 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured left femoral artery pseudoaneurysm (impella side) with a "definite" relationship to the impella device used: ¿following impella sheath removal from l groin, hematoma noted, believed it was from track ooze as angio shows hemostasis of vessel. arterial us pod1 l groin positive for two-lobbed psa measuring 1.1cm x 1cm x 1.3 cm x 1.4cm. vascular surgery consult. pod2- successful thrombin injection to l femoral artery pseudoaneurysm. le arterial us negative for psa. hematoma remains. significant ecchymosis in l groin extending into proximal thigh that is soft, palpable dp/pt pulses b/l. baseline hgb 15.6. nadir 12.6 on pod1. d/c home in stable condition on pod3 w/ hgb trending up at 13.6.¿ the patient recovered with sequelae. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
26529 | 1220648-2024-11254 | 19331023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331023&pc=OZD | 07/10/2023 | 06/06/2024 |  | 05/16/2024 |  |  | 04/24/2024 | 05/16/2024 |  |  | 05/31/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024355412 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured worsening supraventricular tachycardia with a "possible" relationship to the impella device used: ¿patient had pmh of svt (supraventricular tachycardia) . patient had episode of svt on (b)(6) 2023. patient was placed back on amiodarone drip. impella placement was reviewed. it was founded impella was positioned too deep in lv (causing this arrhythmia). impella was retracted. patient had another episode7/26/23 night. patient was monitored for svt throughout hospital stay until discharge.¿ the patient recovered with no sequelae. | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. d4-updated model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
26530 | 1220648-2024-11255 | 19331031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331031&pc=OZD | 12/14/2021 | 05/22/2025 |  | 05/16/2024 |  |  | 12/14/2021 | 05/16/2024 |  |  | 09/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022141931 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2021 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support an impella cp stopped alarm occurred and multiple attempts to restart were unsuccessful. the physician removed and reapplied the cross clamp a little higher on the aorta with resolution of the alarm and successful restarting of the impella. additionally, it was reported that the patient went into sustained ventricular tachycardia requiring multiple defibrillations. the chest was reopened, and rhythm returned to normal sinus rhythm. no complications were noted after the case was resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26531 | 1220648-2024-11262 | 19331065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331065&pc=OZD | 12/14/2023 | 06/26/2024 |  | 05/16/2024 |  |  | 04/24/2024 | 05/16/2024 |  |  | 09/30/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398830 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating an allegation of ventricular tachycardia: ¿the patient developed ventricular tachycardia during an impella-supported procedure for intervention on acute nstemi. during the procedure, the patient required multiple rounds of cpr before and after the impella cp implant. the surgeon discussed the situation with the patient's spouse, and it was decided to discontinue heroic efforts. the patient developed minimal pulses and ventricular tachycardia before time of death was called.¿ | Additional information: abiomed medical safety office review determined there is not enough information to associate use of device with outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation into the vf/vt/af/at (ventricular tachycardia) issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. b5 additional information added corrected d4 model number. removed f6 and f8. dates included in the initial and should not have been.
26532 | 1220648-2024-11256 | 19331080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331080&pc=OZD | 12/15/2021 | 05/16/2024 |  | 05/16/2024 |  |  | 12/15/2021 | 05/16/2024 |  |  | 09/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022141356 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/15/2021 |  |  |  |  | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. during support the patient experienced hemolysis, the device was turned from p-9 to p-6 and the hemolysis was significantly reduced. additionally, the patient has had intermittent incidents of ventricular tachycardia (vt) and went on ventricular fibrillation (vf). he was defibrillated 7 times. the physician discontinued ketamine. the patient was monitored, and staff confirmed that he did not experience another incident of ventricular tachycardia and was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: hemolysis: cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26533 | 1220648-2024-11263 | 19331216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331216&pc=OZD | 05/07/2024 | 07/12/2024 |  | 05/16/2024 |  |  | 05/09/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481406 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during cardiac surgery. during support, the patient displayed hematuria and hemolyzed labs. the impella was repositioned and the p level was reduced, but the physician ultimately decided to explant the impella. | [5 empty] | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. per the clinical information the pump was repositioned, but due to no change in hemolysis, it was explanted. the root cause of the hemolysis issue was not determined since product was not returned, insufficient clinical information provided, and data logs were inconclusive. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
26534 | 1220648-2024-11264 | 19331337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331337&pc=OZD | 01/06/2022 | 05/16/2024 |  | 05/16/2024 |  |  | 01/06/2022 | 05/16/2024 |  |  | 07/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121804 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/06/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient went into ventricular t fibrillation and the catheter migrated, and urine was pink. the physician stated that echo obtain in ccl and position ¿great¿. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ this report is being filed as part of a retrospective review of historical records.
26535 | 1220648-2024-11266 | 19331389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331389&pc=OZD | 12/26/2021 | 05/14/2025 |  | 05/16/2024 |  |  | 12/26/2021 | 05/16/2024 |  |  | 09/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140844 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/26/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient had oozing at the site. standard care was presented to the patient in the form of manual compression. additionally, patient was having some hemolysis concerns. it was advised to reposition impella to optimal position. doctor was going to be consulted. no additional information related to the resolution of the hemolysis was provided. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event (access site bleeding) are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system instructions for use for the related event (hemolysis) are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26536 | 1220648-2024-11265 | 19331421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331421&pc=OZD | 12/15/2021 | 05/16/2024 |  | 05/16/2024 |  |  | 12/15/2021 | 05/16/2024 |  |  | 05/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080524 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/15/2021 |  |  |  |  | The user facility reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. per the physician, they incorrectly assessed the patient's vasculature and the patient's anatomy was too narrow and calcified to tolerate pump insertion. the complication was managed by ceasing attempt to insert pump and no additional support was necessary. the procedure was successfully completed with complete revascularization. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
26537 | 1220648-2024-11267 | 19331461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331461&pc=OZD | 12/16/2021 | 05/16/2024 |  | 05/16/2024 |  |  | 12/16/2021 | 05/16/2024 |  |  | 06/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022102891 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/16/2021 |  |  |  |  | The user facility reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced perfusion issues in his extremity. the medical professional pried openings into the fourteen french (14 fr) sheath to achieve better perfusion in the patient¿s extremity. it is noted that the medical staff utilized a technique that is not indicated in the impella cp instructions for use (ifu). no additional information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation." this report is being filed as part of a retrospective review of historical records.
26538 | 1220648-2024-11268 | 19331476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331476&pc=OZD | 10/26/2023 | 06/10/2024 |  | 05/16/2024 |  |  | 04/25/2024 | 05/16/2024 |  |  | 08/31/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394277 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating an allegation of groin site bleeding (unknown side): ¿s/p index groin site bleeding noted, no side documented. sheath removed & manual pressure held. hemostasis obtained. baseline hgb 8.1 w/ nadir 6.2. transfused w/ 2 units prbcs. hgb increased to 8.2. d/c to sar in stable condition". the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
26539 | 1220648-2024-11269 | 19331549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331549&pc=OZD | 12/16/2021 | 05/14/2025 |  | 05/16/2024 |  |  | 12/16/2021 | 05/16/2024 |  |  | 08/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022124751 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2021 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support multiple attempts were made to reposition the impella that were unsuccessful. to resolve the positioning issue, the physician decided to remove the impella cp and replace it with a new one. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26540 | 1220648-2024-11270 | 19331605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331605&pc=OZD | 05/09/2024 | 05/15/2025 |  | 05/16/2024 |  |  | 05/09/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025473978 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  | MW5169721 | Yes | Initial submission | Followup | Followup |  |  | 05/09/2024 |  |  |  |  | The user facility reported a patient undergoing coronary artery bypass was to be implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 failed to deliver due to patient anatomy. the physician implanted an impella cp with no patient harm. | The user facility reported a patient undergoing coronary artery bypass was to be implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 failed to deliver due to patient anatomy. following the failed impella 5.5 placement the patient was implanted with impella cp and experienced complications and expired after 2 days of support. it is unknown if the use of the impella 5.5 was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. the complications with the impella cp are documented in manufacturer device report number 1220648-2024-24138. |  |  |  |  | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient condition since resistance was met when advancing the device from axillary insertion site causing the white cable to kink. f6/f8-should have been left blank. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use during impella section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The root cause of the delivery issue was most likely due to patient condition since resistance was met when advancing the device from axillary insertion site causing the white cable to kink. the following fields were updated manufacturer device report on 1220648-2024-11270 based on medwatch (mw) 5169721 that was received. b1 revised to reflect both adverse event and product problem b2 revised to reflect death and date of death. b5 revised with additional information. e4 revised to yes. h1 revised to reflect death. h6 health effect - clinical code and health effect - impact code revised to reflect patient outcome. h9 revised to reflect mw report number. medwatch report. the companion case manufacturer device report number 1220648-2024-24138 has been updated. g1 reporting contact email revised on manufacturer device report 1220648-2024-11270 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 220648-2024-11270.
26541 | 1220648-2024-11271 | 19331629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331629&pc=OZD | 12/16/2021 | 05/16/2024 |  | 05/16/2024 |  |  | 12/16/2021 | 05/16/2024 |  |  | 09/30/2022 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1504946 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 12/16/2021 |  |  |  |  | The user facility reported that a patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support there was continuous oozing noted around the peel away sheath and the patient experienced hematuria. the physician elected to explant the pump. the peel away was removed with 3 perclose closures and contra lateral dry closure. hemostasis was achieved and the patient's distal pulses were palpable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26542 | 1220648-2024-11272 | 19331662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331662&pc=OZD | 01/04/2022 | 04/07/2025 |  | 05/16/2024 |  |  | 01/04/2022 | 05/16/2024 |  |  | 09/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022131554 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed sepsis and was reported to be in poor condition. the decision was made to discontinue therapy. three days later the patient¿s left ventricular ejection fraction improved, catecholamines were no longer required, and the patient was hemodynamically stable. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: cp with smart assist section: table 3.2 impella catheter components as noted in the impella ifu ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings as noted in the impella ifu ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26543 | 1220648-2024-11273 | 19331695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331695&pc=OZD | 05/10/2024 | 06/24/2024 |  | 05/16/2024 |  |  | 05/10/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1167416 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) was not sensing the purge cassette. there was no patient involvement. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the issue with properly detecting the purge pressure disc was reproduced. the purge disc retainer was confirmed to be broken. the cause of the issue was a defective component. d9.
26544 | 1220648-2024-09842 | 19331820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331820&pc=OZD | 04/09/2024 | 05/28/2024 |  | 05/16/2024 |  |  | 04/09/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025445454 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a small hematoma was noted on impella 5.5 access site. systemic heparin was stopped. the patient is stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was determined to be anticoagulation management because heparin was stopped to control the issue. the instructions for use referenced in the initial report is from ifu for the impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states), which now includes the statement "(including ventricular perforation).¿, section: general patient care considerations, and sections: entering cardiac output. d4-updated model number added-d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
26545 | 1220648-2024-09987 | 19331821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331821&pc=OZD | 05/23/2022 | 04/17/2024 |  | 05/16/2024 |  |  | 05/23/2022 | 05/16/2024 |  |  | 01/31/2024 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184799 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/23/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported while on support, the patient's pulses were unable to be dopplered in the right leg. the physician decided to place a right sided antegrade access to help with right leg perfusion. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
26546 | 1220648-2024-09983 | 19331822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331822&pc=OZD | 04/11/2024 | 05/29/2024 |  | 05/16/2024 |  |  | 04/11/2024 | 05/16/2024 |  |  | 07/31/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024370511 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2024 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 implant attempt was unsuccessful. the impella was unable to pass. the doctor decided to remove the impella 5.5 and implant an impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the delivery issue is determined to be patient condition because of the small vessel size that resisted the impella 5.5 implantation. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
26547 | 1220648-2024-09984 | 19331824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331824&pc=OZD | 04/12/2024 | 05/29/2024 |  | 05/16/2024 |  |  | 04/12/2024 | 05/16/2024 |  |  | 06/30/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024431034 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 68 year-old female patient with high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient was transferred to another facility with a peel away sheath sutured in the leg. the medical staff was unable to find a pulse in the leg and foot which had been present earlier according to the patient¿s chart. the patient was taken back to the catheterization lab for percutaneous coronary intervention of the left anterior descending artery. at the end of the procedure, the physicians wanted to remove the peel away sheath that had been in for 24 hours and advance the repositioning sheath. once the repositioning sheath was advanced, bleeding occurred, and one unit of blood product was given to the patient. the patient survived the event and the groin site remained intact, without bleeding or hematoma, after several days. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed and limb ischemia has been completed. the device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause of the access site bleed. the root cause of the access site bleeding was most likely due to use issue since the peel away sheath was left inside the vessel for 24 hours causing the vessel to be dilated to 14fr for 24 hours. as the necessary clinical information was not provided, the root cause of the limb ischemia was not determined. the instructions for use referenced in the initial report is from ifu for impella cp with smart assist system in section: potential adverse events (united states), which now includes a statement "cardiac or vascular injury (including ventricular perforation).¿, section: general patient care considerations, and section: entering cardiac output. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
26548 | 1220648-2024-09840 | 19331825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331825&pc=OZD | 04/09/2024 | 05/23/2024 |  | 05/16/2024 |  |  | 04/09/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025450299 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 57-year-old male of an unknown race in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the cannula was noted to be kinked after the insertion. the cp was removed and replaced without complication. no patient harm was reported. | [5 empty] | The investigation for the reported cannula kink has been completed. product was not returned for evaluation. data logs were reviewed which confirmed the low flow when pump started and remained to the rest of the case that triggered auto suction response and suction alarms. based on clinical description and data analysis, the root cause of low flow was most likely kinked cannula. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26549 | 1220648-2024-09985 | 19331826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331826&pc=OZD | 05/16/2022 | 05/14/2025 |  | 05/16/2024 |  |  | 05/16/2022 | 05/16/2024 |  |  | 03/31/2024 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191950 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during insertion, it was determined that the patient had high grade aortic valve stenosis. as a result of the high-grade aortic valve stenosis, impella insertion was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26550 | 1220648-2024-09986 | 19331827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331827&pc=OZD | 04/12/2024 | 09/23/2024 |  | 05/16/2024 |  |  | 04/12/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445186 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Damaged by Another Device | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | SHOCKWAVE | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 78-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the placement signal of an impella cp pump disappeared immediately after implant. as a result of the noted issue, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue has been completed since the original report was filed. the pump was returned for investigation. no physical abnormalities were observed. however, the data log showed that the issue had occurred. placement signal was lost immediately after the shockwave device used. the cause of the optical signal issue was a damaged sensor due to the shockwave device interaction. | Correction to mdr mfr report 1220648-2024-09986-01: added shockwave as a concomitant medical product due to the investigation findings that confirm the shockwave device interaction was the root cause of the optical signal issue.
26551 | 1220648-2024-09869 | 19331828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331828&pc=OZD | 04/09/2024 | 05/20/2024 |  | 05/16/2024 |  |  | 04/09/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451668 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that bleeding was noted and manual pressure was held. the patient continued to bleed despite pressure. the hematoma grew despite all efforts. cutdown was performed to the left femoral artery (lfa) by the surgeon while a new pump was placed. hematoma evacuation and repair was performed to the lfa. additionally, the team infused blood products. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and limited information is provided in the clinical description. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
26552 | 1220648-2024-09988 | 19331829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331829&pc=OZD | 04/12/2024 | 05/20/2024 |  | 05/16/2024 |  |  | 04/12/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1088164 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility's biomedical engineer reported that the automated impella controller (aic) had a battery failure (red alarm). the aic was replaced. no patient was involved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - battery failure (red alarm) has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. battery failure alarm issue was able to be reproduced. issue was determined to be caused by internal battery cell imbalance.
26553 | 1220648-2024-10040 | 19331830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331830&pc=OZD | 05/09/2022 | 04/18/2024 |  | 05/16/2024 |  |  | 05/09/2022 | 05/16/2024 |  |  | 02/29/2024 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195178 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/09/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced preperitoneal bleeding at the impella insertion site. the right groin site was reportedly hard, a hematoma and bleeding were noted. a computer tomography scan was ordered to rule out retroperitoneal bleeding. the patient received two units of packed red blood cells, also heparin was removed from the purge solution and switched with bicarbonate. the patient was hemodynamically stable and remained on impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
26554 | 1220648-2024-10328 | 19331831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331831&pc=OZD | 05/19/2022 | 05/14/2025 |  | 05/16/2024 |  |  | 05/19/2022 | 05/16/2024 |  |  | 02/29/2024 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023188875 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Injury | Yes | No | Improper Flow or Infusion | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was high purge pressure. moreover, the patient was heparin-induced thrombocytopenia positive, and heparin was removed from the purge. it was reported that purge pressure returned to normal limits. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26555 | 1220648-2024-10039 | 19331832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331832&pc=OZD | 07/08/2022 | 05/14/2025 |  | 05/16/2024 |  |  | 07/08/2022 | 05/16/2024 |  |  | 02/29/2024 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202563 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2022 |  |  |  |  | The user facility reported a 68-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the leg on the impella side was swollen and went from warm to cold to the touch. the patient was successfully weaned from the device and survived explant. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26556 | 1220648-2024-10495 | 19331833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331833&pc=OZD | 04/08/2022 | 05/14/2025 |  | 05/16/2024 |  |  | 04/08/2022 | 05/16/2024 |  |  | 08/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022133011 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. post placement, bleeding and hematoma developed in the right femoral artery. the medical professional attempted numerous times to access the groin to control the bleeding which was initially thought to be coming from a tissue track. utilization of fem-stop and holding manual pressure were unsuccessful as the groin and hematoma continued to increase in size. one (1) unit of packed red blood cells were provided to the patient. the medical staff consulted and determined that a new impella device would be placed through the left femoral artery. the patient was taken to the operating room (or) for groin site exploration and removal of the implanted impella device and sheath. the right femoral (rf) impella was pulled back into the descending aorta prior to insertion of the left femoral impella. the rf device and sheath were removed without issue. surgical examination and device explant identified the right femoral vein was damaged and the probable cause of the bleed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26557 | 1220648-2024-10772 | 19331834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331834&pc=OZD | 04/27/2024 | 06/25/2024 |  | 05/16/2024 |  |  | 04/27/2024 | 05/16/2024 |  |  | 06/30/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024368727 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2024 |  |  |  |  | The user facility reported a 59-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that after impella cp placement, there was noted limb ischemia due to the small nature of the femoral artery. the ischemia was treated by placing external bypass. the distal limb was then perfused effectively. the patient remained on support. | There is not enough information to associate use of the device with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the limb ischemia - reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
26558 | 1220648-2024-10194 | 19331835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331835&pc=OZD | 04/05/2024 | 05/28/2024 |  | 05/16/2024 |  |  | 04/05/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446492 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 74-year-old female patient with acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed leg pain on the impella side. the ultrasound showed poor distal flow. a decision was made to explant the impella cp and evaluate further. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the pump was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
26559 | 1220648-2024-10330 | 19331836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331836&pc=OZD | 05/10/2022 | 04/08/2025 |  | 05/16/2024 |  |  | 05/10/2022 | 05/16/2024 |  |  | 02/29/2024 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023194840 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2022 |  |  |  |  | The user facility reported a 54 year old male patient was implanted with an impella cp for mechanical circulatory support due to heart dysfunction. while on support, the left ventricle metrics improved with reposition. additionally, there was continuous oozing at all sites despite normal coagulations. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26560 | 1220648-2024-10677 | 19331838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331838&pc=OZD | 03/13/2022 | 05/02/2024 |  | 05/16/2024 |  |  | 03/13/2022 | 05/16/2024 |  |  | 12/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174765 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/13/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support there was access site bleeding (oozing) that was treated by giving the patient one unit of blood. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26561 | 1220648-2024-10709 | 19331839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331839&pc=OZD | 03/20/2022 | 05/14/2025 |  | 05/16/2024 |  |  | 03/20/2022 | 05/16/2024 |  |  | 11/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162823 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/20/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, patient experienced bleeding from the mouth and nose due to high activated clotting time (act) and possible hemolysis. bleeding was monitored and patient was then placed on heparinized purge fluid and also given one unit of packed red blood cells. the patient reportedly became hemodynamically stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26562 | 1220648-2024-10415 | 19331841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331841&pc=OZD | 03/26/2024 | 06/03/2024 |  | 05/16/2024 |  |  | 03/26/2024 | 05/16/2024 |  |  | 11/30/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024424530 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during impella cp support, the female patient had constant oozing of blood from the puncture site. the site was treated locally with tranexamic acid and suprarenin compress. two units of packed red blood cells were also administered. the patient was noted to have expired after being on support for 2 days. however, the access site oozing had already been successfully managed. the patient expired from progressive multi-organ failure. the oozing had no effect on the patient's death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- impact code 1802 in accordance with updated procedures, sections h6 and h10 have been revised from the initial submission.
26563 | 1220648-2024-10496 | 19331842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331842&pc=OZD | 05/12/2022 | 05/13/2025 |  | 05/16/2024 |  |  | 05/12/2022 | 05/16/2024 |  |  | 07/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022118499 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2022 |  |  |  |  | The user facility reported a 92-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis that was resolved by removal of heparin and administration of packed red blood cells. patient¿s impella was successfully weaned. patient has good pulses and the groin wound was closed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26564 | 1220648-2024-10853 | 19331843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331843&pc=OZD | 04/28/2024 | 06/03/2024 |  | 05/16/2024 |  |  | 04/28/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025458116 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/28/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the team placed the impella cp due to respiratory distress. the cp limb experienced signs of ischemia and so the team placed external fem-fem bypass. the patient remains on impella cp support on the date of this report. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿ | The investigation for the reported limb ischemia has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia could not be determined as the device was not returned nor sufficient clinical details. the instructions for use referenced in the initial report is for the impella cp with smart assist system.
26565 | 1220648-2024-10473 | 19331845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331845&pc=OZD | 04/08/2024 | 05/23/2024 |  | 05/16/2024 |  |  | 04/08/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025457935 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2024 |  |  |  |  | The user facility reported a 80-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that while preparing for the impella cp implant, the purge solution failed to prime properly through the pump's outlet. the setup was disassembled and reassembled, but the purge fluid once again stopped flowing when the purge line was connected to the purge sidearm. as a result of the persistent issue, the decision was made to replace the pump. there was no patient harm resulting from the failure. | [5 empty] | The investigation for the reported priming issues has been completed. the impella was returned for evaluation. during setup process for impella cp, it was noted that the purge solution was not priming through the pump¿s outlet. troubleshooting processes were run and setup process was repeated and purge fluid confirmed priming prior to connection to purge sidearm. purge fluid dripped once through outlet and stopped. no alarms noted on console, purge priming process continued for another 5 minutes, with no purge solution noted through the outlet. physician elected to setup a new impella cp on a different console. second impella primed appropriately and inserted into patient for duration of case without issues. the data logs show no priming issue and insufficient time for priming process. upon evaluation of the product, there were no abnormalities found and the priming issue was not reproduced. there was no problem found as the root cause of the priming issue. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26566 | 1220648-2024-10706 | 19331846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331846&pc=OZD | 03/18/2022 | 05/03/2024 |  | 05/16/2024 |  |  | 03/18/2022 | 05/16/2024 |  |  | 12/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176474 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 03/18/2022 |  |  |  |  | The user facility reported a 68-year-old patient needing mechanical circulatory support. it was reported that the patient was placed on impella cp support for high-risk percutaneous coronary intervention. while on support, the patient experienced limb ischemia that was resolved by removing the impella. additionally, the patient experienced access site bleeding that was resolved with pump removal and addition of blood products. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records.
26567 | 1220648-2024-10852 | 19331848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331848&pc=OZD | 04/26/2024 | 06/19/2024 |  | 05/16/2024 |  |  | 04/26/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025457196 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2024 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient on impella cp support developed bleeding from the access site two units of blood and sutures were given. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, sutures were given to attain hemostasis as per the clinical description.
26568 | 1220648-2024-10276 | 19331849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331849&pc=OZD | 05/10/2022 | 05/27/2025 |  | 05/16/2024 |  |  | 05/10/2022 | 05/16/2024 |  |  | 12/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176222 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2022 |  |  |  |  | The complainant reported the 22-year-old male patient was implanted with impella cp support due to pulseless electrical activity arrest. while on support, the patient experienced access site bleeding and limb ischemia. no medical or surgical intervention was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26569 | 1220648-2024-11252 | 19331851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331851&pc=OZD | 12/14/2021 | 03/28/2025 |  | 05/16/2024 |  |  | 12/14/2021 | 05/16/2024 |  |  | 09/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022146017 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the peel away sheath was cracked inside of the femoral artery. pressure was held and the sheath was removed, peeled the rest of the way off the catheter and the repo was inserted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26570 | 1220648-2024-11092 | 19331853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331853&pc=OZD | 05/07/2024 | 03/19/2025 |  | 05/16/2024 |  |  | 05/07/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025459446 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Electromagnetic Interference | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | PACEMAKER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/07/2024 |  |  |  |  | The user facility reported a male patient of unknown in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that an interrogation was performed to change the settings of a leadless pacemaker implanted in a patient being managed by impella cp, but it could not be read. the possible causes were noise from electrical outlets around the bed and electromagnetic interference from the impella. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the possible causes were noise from electrical outlets around the bed and electromagnetic interference from the impella. no product was returned for evaluation. the root cause of the positioning issue was unable to be determined as insufficient information was provided to determine if there was a relationship between the mapping device and the use of impella. d.6b explant date was added. b.7 information was added that was inadvertently omitted on the initial manufacturer device report number 1220648-2024-11092. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11092 and should not have been. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-11092 was submitted in accordance with updated procedures.
26571 | 1220648-2024-11251 | 19331854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331854&pc=OZD | 05/12/2024 | 07/10/2024 |  | 05/16/2024 |  |  | 05/13/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  |  |  | Malfunction | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The complainant reported a 74-year-old male patient with a high risk of percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. during the case there were frequent purge pressure alarms as well as at one point air in purge alarm. the complainant reported the patient had st elevated myocardial infarction of an in-stent thrombosis. the pump alarmed for purge, and at the time of explant, the thrombosis was observed. the patient experienced angina at the time of the explant. the patient survived the event. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported thrombus and high purge pressure has been completed. no product was returned for investigation. the pump serial number and console are unknown thus data logs could not be obtained. the cause of the high purge pressure was not determined since product was not returned. the root cause of the thrombus could not be determined without additional clinical details. f6/f8 should have been left blank in the initial report. g1 reporting contact fax number missing.
26572 | 1220648-2024-10708 | 19331855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331855&pc=OZD | 03/12/2022 | 03/28/2025 |  | 05/16/2024 |  |  | 03/12/2022 | 05/16/2024 |  |  | 12/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173053 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2022 |  |  |  |  | The complainant reported a patient in non-st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, patient experienced mottled skin from legs to trunk of body, and a distal aortic clot was observed on the tee (transesophageal echocardiography). upon explant, clot was visualized on impella motor and was prolapsed back over motor and cannula. the medical team determined the entire clot was extracted on explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26573 | 1220648-2024-10787 | 19331856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331856&pc=OZD | 04/22/2022 | 04/17/2025 |  | 05/16/2024 |  |  | 04/22/2022 | 05/16/2024 |  |  | 09/30/2022 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021008120 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/22/2022 |  |  |  |  | It was reported that while the patient was on impella cp support, the patient had bleeding at the access site. systemic heparin was stopped and the reposition sheath was inserted. bleeding was controlled and the patient's left groin site was redressed and sutured. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26574 | 1220648-2024-10707 | 19331857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331857&pc=OZD | 03/18/2022 | 03/28/2025 |  | 05/16/2024 |  |  | 03/18/2022 | 05/16/2024 |  |  | 10/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144084 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis from insertion to withdrawal. no medical or surgical intervention was reported. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26575 | 1220648-2024-10198 | 19331859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331859&pc=OZD | 04/30/2022 | 04/22/2024 |  | 05/16/2024 |  |  | 04/30/2022 | 05/16/2024 |  |  | 06/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMPTEMPORARY NON-ROLLER TYPE |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022104220 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/30/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support it was reported that the patients' labs were hemolyzing and patient's urine output was dropping. the impella was repositioned. the staff was advised to check the placement again or decrease the p level if the hemolysis does not resolve. no further information was provided if hemolysis resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
26576 | 1220648-2024-10995 | 19331860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331860&pc=OZD | 05/03/2024 | 06/13/2024 |  | 05/16/2024 |  |  | 05/03/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492791 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2024 |  |  |  |  | The user facility reported an unspecified female patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that patient on impella cp support developed bleeding in the impella access site. the cp was removed. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella cp with smart assist system states section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d4-corrected model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
26577 | 1220648-2024-11241 | 19331864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331864&pc=OZD | 04/25/2024 | 06/27/2024 |  | 05/16/2024 |  |  | 04/25/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1462815 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The complainant reported that during training, the battery was unable to be loaded and both automated impella controllers (aic) displayed an error message. there was no patient involvement. | [5 empty] | The aic has not been received from the customer for evaluation. upon completion of the investigation, a supplemental mdr will be filed. | The investigation for the reported console (aic) battery failure has been completed. the console was tested and the battery failure alarm issue was able to be reproduced. after reviewing the data logs, the battery error issue was confirmed on this console. there was one instance of battery failure alarm (transport connection blown/missing) found from the reported occurrence date. results of running batttest on telnet revealed that ¿fu¿ was missing on both batteries. issue was determined to be caused by depleted batteries. the console was last disconnected from alternating current (ac) power on 2023.06.23 and was never plugged back to ac power since. these depleted batteries could not be charged back up because they were locked out due to low cell voltage (less than 2.3v). the root cause of this issue was the automated impella controller was not being routinely charged. a.1 revised was not previously entered on the initial submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. d.4 serial number was revised as previously entered incorrectly. g.1 manufacturing site facility name was revised as previously entered incorrectly. g.1 manufacturing site address line 2 was entered on the previous submitted report and should of been left blank. g.1 manufacturing site country was revised as it was inadvertently omitted from the previously submitted report. h.6 medical device problem code was revised from the initial submission in accordance with updated procedures.
26578 | 1220648-2024-10414 | 19331866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331866&pc=PYX | 10/20/2020 | 06/11/2024 |  | 05/16/2024 |  |  | 04/23/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | While conducting a literature review of the impella rp the tct document was located inclusive of the recover right enrolled patients. the impella rp was used in 22 patients and adverse events were noted. they are as follows: four major bleeds, two access site bleeds, two hemolysis, and three device malfunctions. the malfunctions were not classified. the document did conclude the 22 impella rp supported patients that had higher survival than those in the salvage arm, which were 44 patients. both arms had no pulmonary embolism. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. d.4 catalog number, serial number, lot number, expiration date, and udi number are unknown. d.6a/d.6b implant and explant dates are unknown. h.4 device manufacturer date is unknown. impella rp with smart assist system with automated impella controller section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound." ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | The investigation for the reported access site bleed and hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and hemolysis could not be determined.
26579 | 1220648-2024-11274 | 19331869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331869&pc=OZD | 12/26/2021 | 03/28/2025 |  | 05/16/2024 |  |  | 12/26/2021 | 05/16/2024 |  |  | 08/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022126961 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/26/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient experienced nonsustained runs of ventricular tachycardia (vt) and significant amount of ectopy. the medical team closely monitored patient for ectopy. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26580 | 1220648-2024-10661 | 19331873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331873&pc=OZD | 04/18/2024 | 07/10/2024 |  | 05/16/2024 |  |  | 04/18/2024 | 05/16/2024 |  |  | 10/31/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024409055 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Pressure Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella 5.5 support for 65-year-old male patient, the purge system was obstructed. there were no kinks noted in the purge line. the patient's acts (activated clotting time) were between 160/180. the pump's p-levels were reduced but the issue was not resolved. when the device was explanted, part of the heart muscle was around the motor house. the patient was stable post explant. | [5 empty] | The investigation into the high purge pressure issue has been completed. the pump was returned for investigation. the purge system was obstructed. no data logs were returned. the product was inspected, and biomaterial was found blocking the purge gap. the pump was run in the lab after soaking and high purge pressure was not reproduced. the root cause of the high purge pressure was most likely the biomaterial that blocked the purge gap during support. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ e.1 address line 1 and 2 was revised as it was entered previously incorrectly on manufacturer device report number 1220648-2024-10661. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-10661 and should not have been. g.1 revised manufacturing site name and address section as previously entered incorrectly on manufacturer device report number 1220648-2024-10661. h.6 code 3012 was reported incorrectly on manufacturer device report number 1220648-2024-10661. a new code has been added to medical device problem codes. h.10 additional manufacturer narrative instructions for use were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-10661 was submitted. | The impella device was received from the customer on 01-may-2024 and therefore, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
26581 | 1220648-2024-10773 | 19331877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331877&pc=OZD | 04/25/2022 | 03/28/2025 |  | 05/16/2024 |  |  | 04/25/2022 | 05/16/2024 |  |  | 01/31/2024 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185877 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2022 |  |  |  |  | The complainant reported a patient presenting with an unspecified issue was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced access site bleeding. 3 mitral clips were placed, and the physician decided to wean the patient and explant the device. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26582 | 1220648-2024-11275 | 19331894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331894&pc=OZD | 12/18/2021 | 05/16/2024 |  | 05/16/2024 |  |  | 12/18/2021 | 05/16/2024 |  |  | 06/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022099870 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 12/18/2021 |  |  |  |  | The user facility reported a 43-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during explant, it was observed that the patient experienced a hematoma at the impella access site. the patient was administered one unit of packed red blood cells. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records.
26583 | 1220648-2024-11276 | 19331918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331918&pc=OZD | 05/10/2024 | 07/10/2024 |  | 05/16/2024 |  |  | 05/10/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025462019 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support due to needing a myocardial perfusion. as the doctor was pulling away the peel away sheath, the clear plastic ring never broke in half. the doctor had to use hemostats to break off plastic ring. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the difficulty insertion/removing introducer has been completed. no product was returned for evaluation. clinical details noted that the clear plastic ring on the introducer sheath did not break in half as intended and required forceps to remove. the root cause of the difficulty removing the introducer cannot be determined as no product was returned for evaluation. f6/f8 should have been left blank in the initial report. g1 reporting contact fax number missing.
26584 | 1220648-2024-11277 | 19331962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331962&pc=OZD | 12/17/2021 | 05/16/2024 |  | 05/16/2024 |  |  | 12/17/2021 | 05/16/2024 |  |  | 06/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095292 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 12/17/2021 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced hematuria; hemolysis was diagnosed with the observation of tinged urine in the foly catheter. additionally, it was reported that dobutamine was decreased after the patient experienced an episode of ventricular tachycardia. the impella cp was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26585 | 1220648-2024-11278 | 19331986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19331986&pc=OZD | 01/17/2022 | 03/28/2025 |  | 05/16/2024 |  |  | 01/19/2022 | 05/16/2024 |  |  | 09/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022151073 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/19/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient had small occipital embolic stroke and gi bleed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26586 | 1220648-2024-11280 | 19332183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19332183&pc=OZD | 05/12/2024 | 06/27/2024 |  | 05/16/2024 |  |  | 05/12/2024 | 05/16/2024 |  |  | 07/31/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024371685 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2024 |  |  |  |  | The complainant reported a 72-year-old female patient with elective placement was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported the patient had a coronary artery bypass graft. after three days of support, the pump had saturated flows and stopped. medical staff removed the pump, and the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into the saturated flow and the pump stop issues has been completed since the original report was submitted. the device was not returned for investigation. per data log, high purge pressure and purge system blocked alarms were resolved on their own. the root cause of the saturated flow and the pump stop was not determined since product was not returned.
26587 | 1220648-2024-11281 | 19332197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19332197&pc=OZD | 12/26/2021 | 03/28/2025 |  | 05/16/2024 |  |  | 12/26/2021 | 05/16/2024 |  |  | 10/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142127 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/26/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while priming the impella cp the patient experienced ventricular fibrillation (vf). the patient reverted to normal sinus rhythm (nsr) after impella cp support was started. impella was successfully weaned and explanted. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26588 | 1220648-2024-11282 | 19332215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19332215&pc=OZD | 05/12/2024 | 07/29/2024 |  | 05/16/2024 |  |  | 05/13/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025470138 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The user facility reported a 72-year-old male with cardiomyopathy was to be implanted with an impella cp device for mechanical circulatory support. it was that the case was aborted due to the patient¿s tortuous vascular system. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings & cautions: to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the blood vessels in the lower extremities were tortuous, it was not possible to proceed beyond the cia. f6/f8 should have been left blank in the initial report.
26589 | 1220648-2024-11283 | 19332254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19332254&pc=OZD | 05/13/2024 | 06/27/2024 |  | 05/16/2024 |  |  | 05/13/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025450918 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The user facility reported a 73-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient on impella cp support arrived at another facility with femstop in position, due to a significant hematoma extending to patient¿s scrotum. vascular surgery and urology have been consulted to address complication. the patient was reported to be in critical condition. | [5 empty] | The investigation into the access site bleeding ¿ major issue been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26590 | 1220648-2024-11279 | 19332478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19332478&pc=OZD | 05/09/2024 | 06/12/2024 |  | 05/16/2024 |  |  | 05/11/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1084502 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/11/2024 |  |  |  |  | The user facility reported that during an automated impella controller (aic) to aic transfer, the grey and blue portion of the catheter plug port came out with the white cable. the patient was transferred to another aic without further issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) damage has been completed. the console was returned for evaluation. upon visual inspection, the unit was found with the front panel optical connector broken. data logs were reviewed but were unable to confirm the reported failure mode from the logs. the root cause of the broken front panel optical connector was most likely due to a use issue. added- d9 device return date d2-corrected common device name.
26591 | 1220648-2024-11284 | 19332920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19332920&pc=PYX | 05/10/2024 | 09/17/2024 |  | 05/16/2024 |  |  | 05/10/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025479864 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella rp flex device for mechanical circulatory support. it was reported that while advancing the pulmonary artery catheter, the .027 guidewire advanced into the right pulmonary artery and caused a suspected perforation. the impella rp flex pump was still implanted successfully following the event for patient support. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the vascular damage has been completed. the impella was returned for evaluation. during visual inspection, biomaterial was found wrapped around the underside of the impeller. no cannula kinks were noted. the pump passed electrical testing and all 5s parameters were within normal ranges. the guidewire was not returned. when the pump was tested in the pressurized flow loop, it drifted upwards from ~40 to 75 mmhg over a span of 2 hours while cvp remains stable. the dp sensor drift was reproduced in the flow loop. data logs were reviewed which showed that during support, the placement signal (ps) begins to drift upwards with flows drifting downwards. there are also ingestion events noted on the data logs including motor current spikes, ps shifts, speed deviation, drop in flows. the clinical states that due to the perforation, the patient was unable to receive adequate anticoagulation. the root cause of the placement signal issues was dp sensor drift. the root cause of the low pump flow issue was most likely biomaterial. the root cause of the vascular damage great vessel pulmonary artery (pa) perforation issue was most likely use related as the guidewire got caught on inter devices and jumped forward. the instructions for use for the impella rp system with automated impella controller in section: potential adverse events (united states) states the following: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
26592 | 1220648-2024-11285 | 19333051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19333051&pc=OZD | 05/12/2024 | 07/03/2024 |  | 05/16/2024 |  |  | 05/12/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025464733 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2024 |  |  |  |  | The user facility reported a 69-year-old male of unknown race undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 had a high purge pressure alarm. all troubleshooting measures were taken. the impella was removed, and the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure (hpp) has been completed. no product was returned for investigation. the clinical mentions that tpa was run but the timing is unknown, however the hpp is eventually resolved. data logs show that less than 9 hours into the case the purge pressure begins to gradually rise from 600 to 800 mmhg. it remains around 800 mmhg for 27 hours before suddenly increasing again and triggering hpp and purge system blocked alarms. it remains high for 12 hours with continuous alarms and then is resolved over a period of about 13 minutes just prior to being explanted. motor current (mc) also drops with variations in motor speed indicating that something is impacting the motor. the root cause of the high purge pressure issue was most likely a buildup of biomaterial. f6/f8-should have been left blank. g1a reporting contact fax number was missing
26593 | 1220648-2024-11286 | 19333251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19333251&pc=OZD | 05/12/2024 | 07/10/2024 |  | 05/16/2024 |  |  | 05/12/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474197 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during insertion of an impella 5.5 device for support to a patient with cardiomyopathy, the peel away sheath cracked. it was noted there was too much leakage to salvage, and therefore, another peel away sheath was successfully used. there was no report of harm to the patient. | It was further reported that support was withdrawn from the patient and the patient subsequently expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate the use of the device with the patient's outcome. |  |  |  |  | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | B.5 additional information was received and abiomed's medical safety officer review was completed. d.6b explant date was received. the investigation for the introducer damage-mechanical has been completed. no product was returned for investigation. the root cause of the introducer damage mechanical issue was not determined since the product was not returned. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-11286 was submitted.
26594 | 1220648-2024-11287 | 19333456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19333456&pc=OZD | 05/13/2024 | 06/27/2024 |  | 05/16/2024 |  |  | 05/13/2024 | 05/16/2024 |  |  | 11/30/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430098 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2024 |  |  |  |  | The user facility reported a 74-year-old male with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the patient had poor flow down a leg. the medical staff repositioned the impella cp and ultimately ended up removing it. the medical staff consulted vascular surgery to perform a thrombectomy. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation into the limb ischemia ¿ reversible and the thrombus issues been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and the thrombus was not determined.
26595 | 1220648-2024-11288 | 19333571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19333571&pc=OZD | 04/16/2024 | 10/01/2025 |  | 05/16/2024 |  |  | 04/16/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025468724 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/16/2024 |  |  |  |  | The user facility reported a 78-year-old female with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced some oozing around their access site after the patient was moved. the site was re-dressed and lidocaine with epinephrine was administered to manage the condition. in addition, the patient was given two units of blood to assist in treating the blood loss. the patient decompensated and the family decided to withdraw care and the patient expired. there is no association between impella use and outcome. the patient's peri-procedure condition was critical. | Patient continued to decompensate overnight. family withdrew care. | The complainant also reported that the implanted pump was experiencing persistent positioning alarms during support. troubleshooting steps were provided for the staff. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major and the optical signal issue issues has been completed since the original report was submitted. the device was not returned for investigation. it was reported that oozing at the access site was also observed, and the patient was known to have dic (clotting disorder); patient continued to decompensate overnight, and family decided to withdraw care. the root cause of the access site bleeding major issue was most likely patient condition related due to the patient's dic disorder. the root cause of the positioning issue was most likely patient movement as the pump was not in proper position after the patient moved. | B1 product problem updated. b2 updated to death. b5 updated to include placement signal issue description. h1 updated to death. h6 updated to include placement signal coding and death. corrections that were made from the initial manufacturer device report number 1220648-2024-11288. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email reporting contact fax number updated. f6 and f8 should not have had entries in the initial report.
26596 | 1220648-2024-11289 | 19333760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19333760&pc=OZD | 05/13/2024 | 06/24/2024 |  | 05/16/2024 |  |  | 05/13/2024 | 05/16/2024 |  |  | 11/30/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432841 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient planned for high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support and experienced bleeding and hematoma at the access site. manual pressure, sutures and angle matching were performed and successfully resolved the bleeding. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations .¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the product was not returned for investigation. the root cause of access site bleeding issue is determined to be use issue because the issue was resolved with manual pressure, sutures, and angle match.
26597 | 1220648-2024-11291 | 19336718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19336718&pc=OZD | 04/16/2024 | 06/11/2024 |  | 05/16/2024 |  |  | 04/16/2024 | 05/16/2024 |  |  | 09/30/2025 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444123 |  | 1000080 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/16/2024 |  |  |  |  | It was reported that the patient was placed onto impella cp support for acute myocardial infarction / cardiogenic shock. while on support the patient experienced hemolysis that was resolved with 1l normal saline bolus. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation, for the reported issue of hemolysis, a supplemental report will be submitted. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ minor and the hemolysis issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to anticoagulation management since the act was 250 which is above the range limit. based on the clinical details, the root cause of the hemolysis issue is most likely due to the patient¿s condition since it resolved after the patient received saline bolus. impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
26598 | 1220648-2024-11292 | 19336859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19336859&pc=OZD | 04/22/2024 | 09/22/2025 |  | 05/16/2024 |  |  | 04/22/2024 | 05/16/2024 |  |  |  |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025464843 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/22/2024 |  |  |  |  | It was reported that the patient was placed onto impella 5.5 support for acute myocardial infarction / cardiogenic shock. while on support, the patient experienced ventricular tachycardia that was resolved by shock and providing antiarrhythmic medication as well as ventricular fibrillation that was resolved by cardioversion. the pump also experienced suction that was resolved with a 250ml bolus. | The complainant also reported that, while on support, the pump experienced optical signal issues that were resolved by disabling the positioning alarm and utilizing alternate means for monitoring position. the loss of placement signal did not prompt any intervention or pump replacement. |  |  |  |  | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | The low pump flow failure mode was deleted as there was no reports of low pump flow from the complaint intake or case updates. the data logs also show no occurrence of low pump flows. the logs show impella alarms for position in aorta and impella position unknown. positioning alarm was disabled. the data logs confirm multiple position related alarms after implant. without additional clinical details, the root cause of vf/vt/af/at could not be determined. the following have been reported incorrectly or missing from manufacturer device report. b1 adverse event only f6/f8-should have been left blank. g1 reporting contact email revised in accordance with updated procedures. g1 reporting contact facility name missing g1 reporting contact fax number missing. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the data logs show no occurrence of low pump flows. the logs show impella alarms for position in aorta and impella position unknown. positioning alarm was disabled. the data logs confirm multiple position related alarms after implant. no problem was found as the root cause of the optical signal issue, and it is apparent the console was the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-11292. d3 fax number updated. d10 concomitant medical products and therapy dates updated.
26599 | 1220648-2024-11293 | 19337411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19337411&pc=OZD | 02/17/2022 | 03/28/2025 |  | 05/16/2024 |  |  | 02/17/2022 | 05/16/2024 |  |  | 11/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023164327 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation | Distributor |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved with suture. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26600 | 1220648-2024-11294 | 19337477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19337477&pc=OZD | 02/17/2022 | 03/28/2025 |  | 05/16/2024 |  |  | 02/17/2022 | 05/16/2024 |  |  | 11/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023160202 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2022 |  |  |  |  | The user facility reported a patient was on impella cp for mechanical circulatory support. it was reported the pump got kinked in the aorta after placing across atrioventricular (av) and the physician was unable to unkink the pump. per the nurse, the kink was caused by the short habitus of patient. the impella cp device was pulled out and turned off. an intra-aortic balloon pump (iabp) was placed in lieu of impella cp. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26601 | 1220648-2024-11295 | 19337524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19337524&pc=OZD | 02/18/2022 | 03/28/2025 |  | 05/16/2024 |  |  | 02/18/2022 | 05/16/2024 |  |  | 11/30/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165811 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/18/2022 |  |  |  |  | The user facility reported a 63-year-old female on impella cp for mechanical circulatory support. it was reported while the patient was on impella cp support the patient experienced bleeding at the access site. the bleeding was managed by placing a femstop and manual pressure. blood products have been given. the patient will continue to be managed. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26602 | 1220648-2024-11297 | 19337913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19337913&pc=OZD | 03/01/2022 | 05/22/2025 |  | 05/16/2024 |  |  | 03/01/2022 | 05/16/2024 |  |  | 12/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170767 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/01/2022 |  |  |  |  | The complainant reported a 62 year old male patient was implanted with an impella cp for mechanical circulatory support due to a non-st-elevation myocardial infarction. while on support, nasal and oral bleeding was noted. heparin was put on hold due to bleeding issues. patient was successfully weaned from the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26603 | 1220648-2024-11298 | 19337943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19337943&pc=OZD | 08/26/2022 | 03/27/2025 |  | 05/16/2024 |  |  | 08/26/2022 | 05/16/2024 |  |  | 04/30/2024 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224440 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a 76-year-old male on impella cp for mechanical circulatory support. it was reported that while the patient was on impella cp support and developed access site oozing. the heparin was not in purge as a result of the bleeding and the activated clotting time (act) was 231. no additional information. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26604 | 1220648-2024-11299 | 19337961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19337961&pc=OZD | 12/25/2022 | 03/27/2025 |  | 05/16/2024 |  |  | 12/25/2022 | 05/16/2024 |  |  | 07/31/2023 |  | 05/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022119554 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/25/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced delivery issues that were resolved by removal and insertion of a different pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26605 | 1220648-2024-11179 | 19318901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19318901&pc=OZD | 02/19/2022 | 05/30/2025 |  | 05/15/2024 |  |  | 02/19/2022 | 05/15/2024 |  |  | 12/31/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168914 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 33 YR |  | Female |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/19/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being placed on an impella cp for mechanical circulatory support. the complainant reported that an impella cp pump implant attempt had to be aborted due to the patient's anatomy. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26606 | 1220648-2024-11180 | 19319026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319026&pc=OZD | 02/19/2022 | 05/15/2024 |  | 05/15/2024 |  |  | 02/19/2022 | 05/15/2024 |  |  | 11/30/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167595 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/19/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced a clot in the left ventricle located near the inferior-apical region in a 4-chamber view where the impella pigtail was lying just adjacent to clot. the medical team continued to monitor the clot and echocardiography was performed where the clot appeared smaller. it was also reported that the patient was bleeding from multiple sites and heparin was withheld. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
26607 | 1220648-2024-11181 | 19319125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319125&pc=PYX | 05/02/2024 | 07/09/2024 |  | 05/15/2024 |  |  | 05/02/2024 | 05/15/2024 |  |  | 06/30/2024 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2023226931 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The user facility reported a 78-year-old male patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella rp device for mechanical circulatory support. it was reported that access to the patient's vasculature was noted to be challenging and after multiple unsuccessful attempts, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related. the impella was unable to be placed as accessing the patient's vasculature was very challenging as per the clinical description provided. as noted in the impella instructions for use: impella rp smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ b.7 revised as this information was inadvertently omitted from manufacturer device report number 1220648-2024-11181. h.6 code 4627 was reported incorrectly on manufacturer device report number 1220648-2024-11181. a new code has been added to health effect - impact codes. h.10 additional manufacturer narrative instructions for use were revised from manufacturer device report number 1220648-2024-11181 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section warning and cautions: to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels
26608 | 1220648-2024-11182 | 19319186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319186&pc=OZD | 05/05/2024 | 09/21/2024 |  | 05/15/2024 |  |  | 05/07/2024 | 05/15/2024 |  |  | 11/30/2025 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024429095 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/07/2024 |  |  |  |  | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported that while on impella cp support the patient experienced cerebral hemorrhage, ventricular tachycardia, and death. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The investigation into the vf/vt/af/at issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary information was not provided, the root cause of the vt/vf was not determined.
26609 | 1220648-2024-11183 | 19319216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319216&pc=OZD | 05/07/2024 | 07/01/2024 |  | 05/15/2024 |  |  | 05/07/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474205 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was planned for an impella 5.5 device implant for mechanical circulatory support. however, the impella 5.5 pump was unable to cross the medtronic corvevalve transcather aortic valve replacement after multiple attempts, and an impella cp was placed instead. there was no report of patient harm. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. caution: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the pump could not be placed across the medtronic core tavr valve. impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d.4 revised model number from the initial submission in accordance with updated procedures. h.6 health effect - impact code and medical device problem code were revised in accordance with updated procedures. new codes have been added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
26610 | 1220648-2024-11184 | 19319270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319270&pc=OZD | 05/08/2024 | 06/28/2024 |  | 05/15/2024 |  |  | 05/08/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025444316 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient experienced sustained ventricular tachycardia during impella 5.5 support. the patient was cardioverted to treat the condition. an echocardiogram was utilized following the event and it was noted that the pump was too deep. the pump was subsequently repositioned and support with the pump continued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation into the vf/vt/af/at (sustained ventricular tachycardia) issue has been completed since the original report was filed. the pump was returned for investigation. to determine the true root cause of the vt/vf, relevant clinical information related to cardiac arrythmia would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. d6b added.
26611 | 1220648-2024-11185 | 19319302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319302&pc=OZD | 05/09/2024 | 06/24/2024 |  | 05/15/2024 |  |  | 05/09/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025464844 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The user facility reported a 75-year-old male patient of unknown race and ethnicity noted as high risk percutaneous coronary intervention was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that an impella 5.5 pump could not advance through the patient's anatomy due to the small axillary artery. the implant was subsequently aborted and an impella cp pump was placed for patient support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section warning and cautions: to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels | The investigation into the delivery issues has been completed since the original report was submitted. the product was not returned for investigation. the root cause of the delivery issue was most likely patient condition related, impella 5.5 aborted due to small axillary artery.
26612 | 1220648-2024-11186 | 19319359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319359&pc=OZD | 02/19/2022 | 05/19/2025 |  | 05/15/2024 |  |  | 02/19/2022 | 05/15/2024 |  |  | 11/30/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169308 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2022 |  |  |  |  | The user facility reported a 68-year-old patient was presented in the emergency room with shortness of breath (sob) and hypertension. the patient was taken to the cardiac catheterization laboratory (ccl) with non-st elevation myocardial infarction (nstemi). the impella was implanted without issue. left heart catheterization (lhc) identified lesions in the left anterior descending (lad), circumflex (circ) and right coronary artery (rca) as the patient continued to experience sob. while being placed on bilevel positive airway pressure (bipap), the impella was pulled across the atrioventricular valve (av). the medical professional attempted to cross the av a second time but was unable as the impella prolapsed. the decision was made to remove the initial device and implant a second impella which was placed without issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26613 | 1220648-2024-11187 | 19319434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319434&pc=OZD | 05/07/2024 | 07/10/2024 |  | 05/15/2024 |  |  | 05/10/2024 | 05/15/2024 |  |  | 10/31/2025 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024412010 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility reported a 72-year-old female patient was brought to the cardiac catheterization lab where the left anterior descending (lad) artery was found to be occluded. an impella cp pump was placed, a shockwave procedure was completed to the lad, and a non-compliant nc balloon was added to the lad. during ballooning, the lad was perforated and required multiple covered stents. following the procedure the patient's pulse narrowed and an echo showed an effusion was present. the patient was tapped to treat the effusion, but the patient eventually passed away. the physician stated that they believed the pump caused a free wall rupture due to increased flows. a definitive cause was not established. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings & cautions: dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the reported ventricle perforation has been completed. according to the clinical, before beginning impella cp support, the patient was brought to the cath lab where the lad artery was found to be occluded. the impella cp pump was then placed, a shockwave procedure was completed, and a nc balloon was added. during ballooning, the lad was perforated and required multiple covered stents. after the physician was satisfied with the results, the patient's pulse pressure narrowed and an echo confirmed an effusion. the patient was tapped with massive amounts of volume to treat the effusion, but the patient eventually passed away. only a data stick was returned for a product evaluation. data log analysis confirmed a downward trend in the placement signal near the end of the case. the cause of the ventricle perforation was not determined because no product was returned and not enough clinical information was given. added- h6 component code 925 f6/f8-should have been left blank in the initial report.
26614 | 1220648-2024-11188 | 19319453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319453&pc=OZD | 05/10/2024 | 07/12/2024 |  | 05/15/2024 |  |  | 05/10/2024 | 05/15/2024 |  |  | 09/13/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386965 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility reported a 56-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella cp support started bleeding at impella access site. there was the cp and reperfusion cannula in the right femoral artery, and a central line in the right femoral vein. md decided to remove the central line. hemostasis was difficult to achieve where central line was removed. heparin was stopped, one purse string was added, and several blood products were given. the impella cp was exchanged for 5.5. the patient is stable. | It was reported the care was withdrawn from the patient and the patient subsequently expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) not enough information to associate use of the device with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the following sections: section: general patient care considerations section: entering cardiac output section: potential adverse events (united states) . added- b5 mso review, h6 component code d4-corrected model number f6/f8 should have been left blank in the initial report.
26615 | 1220648-2024-11189 | 19319561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319561&pc=PYX | 05/10/2024 | 07/10/2024 |  | 05/15/2024 |  |  | 05/10/2024 | 05/15/2024 |  |  | 11/30/2025 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024435267 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Difficult to Insert | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Great Vessel Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility reported a 70-year-old female patient in st elevated myocardial infarction was implanted with an impella rp device for mechanical circulatory support. it was reported there was a perforation noted with the guidewire. the 0.027" wire was inserted straight into the pulmonary artery without a curve and appeared to have potentially perforated the pulmonary artery after assessing the computed tomography angiography of the lungs. 70ml of bloody sputum was noted in the collection chamber. protamine was given and bleed ceased after 10min. the patient remained off heparin due to bleeding. the patient was reported as stable. | [5 empty] | The investigation for the reported vascular damage has been completed. according to the clinical, before beginning impella rp support, the 0.027 guidewire was inserted straight into the pa without a curve and perforated the pa. the pump and purge cassette were returned, and no abnormalities were found. the guidewire was not returned for a visual inspection. data log analysis was not applicable for this complaint. the most likely cause of the great vessel vascular damage was use related as the guidewire was inserted straight into the pa resulting in a perforation. added- d9 device return date f6/f8 should have been left blank in the initial report. g1-corrected manufacturing site address, city, zip code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation).
26616 | 1220648-2024-11190 | 19319575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319575&pc=OZD | 01/04/2022 | 05/15/2024 |  | 05/15/2024 |  |  | 01/04/2022 | 05/15/2024 |  |  | 09/30/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2022146398 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Failure | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/04/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during implantation, the patient when asystolic when the impella was advanced across the aortic valve, requiring two rounds of cardio-pulmonary resuscitation and a temporary pacer. reportedly return of spontaneous circulation was achieved after the transvenous pacemaker was placed. also, the patient was bleeding from all the impella access sites, heparin was withheld. the patient continued oozing, a total of 5 units of packed red blood cells. 2 platelets and one plasma. when the procedure was completed the impella repositioning sheath dislodged from the arteriotomy and pulsatile bleeding was noted, the physician re-advanced the sheath, and re-sutured which reportedly resolved the bleeding. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26617 | 1220648-2024-11191 | 19319602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319602&pc=PYX | 05/10/2024 | 07/25/2024 |  | 05/15/2024 |  |  | 05/10/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2025480447 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2024 |  |  |  |  | The user facility reported a 62-year-old male was implanted with an impella rp device for mechanical circulatory support. it was reported that during placement of the central catheter, the utilized guidewire interacted with the impella rp pump, and the placement signal disappeared. flows remained stable following the loss in signal and the patient was later weaned from support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for optical signal issue has been completed. the impella device was not returned for evaluation. review of the data logs show the root cause of the optical signal issue was most likely a damaged sensor caused by an interaction with the guidewire during central line placement. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-11191: f6/f8-should have been left blank . g1 manufacturing site address.
26618 | 1220648-2024-11193 | 19319650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319650&pc=OZD | 05/11/2024 | 07/05/2024 |  | 05/15/2024 |  |  | 05/11/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025457954 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unsuccessful delivery attempt of the impella cp device to a patient in acute myocardial infarction/cardiogenic shock for mechanical circulatory support. the impella cp pump was unable to cross at the level of the iliac bifurcation, due to torturous and calcified anatomy. despite multiple attempts, the impella pump was removed, and the access site was closed without an issue. there was no report of harm to the patient. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely due to patient condition since impella pump was unable to cross at the level of the iliac bifurcation, due to torturous and calcified anatomy. d4-corrected model number.
26619 | 1220648-2024-11192 | 19319652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319652&pc=OZD | 01/03/2022 | 05/14/2025 |  | 05/15/2024 |  |  | 01/03/2022 | 05/15/2024 |  |  | 10/31/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152538 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced cardiac arrhythmia while attempting to place the swan, which was resolved by removing the swan. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26620 | 1220648-2024-11194 | 19319765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319765&pc=OZD | 01/04/2022 | 05/14/2025 |  | 05/15/2024 |  |  | 01/04/2022 | 05/15/2024 |  |  | 08/31/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022124671 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the pump exhibited low purge flow. also, the pump would not restart upon automated impella controller transfer. the team changed the purge fluid and restarted the pump with high motor current and flow saturation. after a few minutes the monitor current mean returned to the 300s and increased to p4. the flows were no longer saturated. the impella cp was replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26621 | 1220648-2024-11195 | 19319817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319817&pc=OZD | 02/20/2022 | 05/14/2025 |  | 05/15/2024 |  |  | 02/20/2022 | 05/15/2024 |  |  | 11/30/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169134 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2022 |  |  |  |  | The complainant reported a patient with aortic valvuloplasty was implanted with an impella cp device for mechanical circulatory support. while on support, site oozing was noted. the site was redressed and heparin was withheld. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26622 | 1220648-2024-11196 | 19319970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19319970&pc=OZD | 01/05/2022 | 05/15/2024 |  | 05/15/2024 |  |  | 01/05/2022 | 05/15/2024 |  |  | 09/30/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022137572 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/05/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support the sheath was left in and the patient experienced leg ischemia. the physician removed the impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation. this report is being filed as part of a retrospective review of historical records.
26623 | 1220648-2024-11197 | 19320293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19320293&pc=OZD | 01/06/2022 | 05/15/2024 |  | 05/15/2024 |  |  | 01/06/2022 | 05/15/2024 |  |  | 08/31/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022132514 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/06/2022 |  |  |  |  | The complainant reported a patient was placed onto impella cp support for acute myocardial infarction / cardiogenic shock. while on support, the patient experienced tachycardia that was resolved with drugs and defibrillation. additionally, the patient experienced sepsis that was resolved with antibiotics. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
26624 | 1220648-2024-11198 | 19320385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19320385&pc=OZD | 05/02/2024 | 07/08/2024 |  | 05/15/2024 |  |  | 05/03/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492079 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2024 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported the patient was on impella cp support and there was low pump flow after implanting. the low pump flow was due to the patient being on extracorporeal membrane oxygenation (ecmo) along with low preload after the time in the operating room for ecmo central cannulation. volume was given. the patient also had access site bleeding. it is unknown if blood products or heparin was withheld. the pump was removed due to the access site bleeding. it is unknown if the low pump flow was resolved. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed and the low pump flow has been completed. the impella was not returned for evaluation. data logs were reviewed and confirmed the failure mode and clinical narrative. logs showed after pump started, pump position was in ventricle followed low flow that triggered impella position in ventricle, auto suction response and suction alarms. the positioning issue occurred for 12 minutes then resolved. the low flow event remained for 8 hours and 15 minutes and was resolved when pump was able to run on p-2 without flow issue. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the cause of the low pump flow was patient condition related because lv was reported to have low volume and data logs indicated pump in ventricle alarms. added- e4-no f6/f8 should have been left blank. g1 reporting contact fax number was missing.
26625 | 1220648-2024-11199 | 19320493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19320493&pc=OZD | 01/03/2022 | 05/14/2025 |  | 05/15/2024 |  |  | 01/03/2022 | 05/15/2024 |  |  | 10/31/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022148987 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had an external perfusion that was placed prophylactically as the patient¿s leg appeared to look mottled. even with the external perfusion, the patient¿s right leg below the knee was pale. the physician came in and suggested removing the intraosseous (io) vascular access in the right leg, which helped. patient was successfully weaned off support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
26626 | 1220648-2024-11200 | 19320496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19320496&pc=OZD | 05/03/2024 | 05/15/2024 |  | 05/15/2024 |  |  | 05/03/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1164961 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient experiencing postcardiotomy cardiogenic shock was implanted an impella 5.5 device for mechanical circulatory support. while on support, the automated impella controller (aic) had a controller error noted as "connect to back up aic". an alternate aic was located and switched over with no loss of support or reported harm to the patient. | [5 empty] | The automated impella controller (aic) was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26627 | 1220648-2024-11201 | 19320531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19320531&pc=OZD | 02/23/2022 | 05/14/2025 |  | 05/15/2024 |  |  | 02/23/2022 | 05/15/2024 |  |  | 11/30/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168936 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, hematoma and oozing were noted. manual pressure was held and two stitches were placed in the artery by cardiovascular surgeon. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined
26628 | 1220648-2024-11202 | 19320553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19320553&pc=OZD | 05/03/2024 | 07/18/2024 |  | 05/15/2024 |  |  | 05/03/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025470347 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2024 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that following a difficult impella cp implant, the patient experienced some bleeding from their access site. it was noted that the patient's blood vessel was deep and had been difficult to puncture. the decrease in hemoglobin levels was treated with a blood transfusion and the condition stabilized. support with the implanted pump was maintained following the intervention. the patient was reported stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was determined to be patient condition because of patient anatomy and the vessels were deep and difficult to puncture.
26629 | 1220648-2024-11203 | 19320582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19320582&pc=OZD | 01/06/2022 | 05/14/2025 |  | 05/15/2024 |  |  | 01/06/2022 | 05/15/2024 |  |  | 09/30/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022146577 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump encountered a clot. due to the patient being on no pressor support, and despite not being able to stent attempted left anterior descending intervention, the physician elected to remove the pump and not replace it. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | The failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
26630 | 1220648-2024-11204 | 19320711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19320711&pc=OZD | 11/01/2022 | 05/16/2025 |  | 05/15/2024 |  |  | 11/01/2022 | 05/15/2024 |  |  | 05/31/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022085415 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella cp device for mechanical circulatory support. during implantation, the impella cp would not pass due to calcium in the femoral artery. a longer sheath was used to pass calcium in the femoral artery. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26631 | 1220648-2024-11205 | 19320824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19320824&pc=OZD | 05/03/2024 | 06/13/2024 |  | 05/15/2024 |  |  | 05/03/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474032 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2024 |  |  |  |  | The user facility reported a 64-year-old male of unknown race and ethnicity post cardiotomy cardiogenic shock/low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that during implantation of an impella 5.5, the pump could not cross a stenosis in the axillary artery. implantation was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the pump could not be crossed a stenosis in the axillary artery despite best efforts.
26632 | 1220648-2024-11206 | 19320983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19320983&pc=OZD | 01/03/2022 | 05/14/2025 |  | 05/15/2024 |  |  | 01/03/2022 | 05/15/2024 |  |  | 08/31/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121528 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/03/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, it was found that the patient's blood was hemolyzed. it is unknown how this was resolved, but the patient remained on impella support for several days after this event. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: hemolysis: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26633 | 1220648-2024-11207 | 19321134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19321134&pc=OZD | 05/04/2024 | 06/11/2024 |  | 05/15/2024 |  |  | 05/04/2024 | 05/15/2024 |  |  | 08/31/2025 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443778 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2024 |  |  |  |  | The user facility reported a 59-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that arterial blood was pooling at the insertion site. the peel away sheath was removed, angle matching was performed, and dressing was placed. forward tension sutures were applied. however, the bleeding could not be controlled and the impella was removed. the patient was reported as stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
26634 | 1220648-2024-11208 | 19321218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19321218&pc=OZD | 01/14/2022 | 05/14/2025 |  | 05/15/2024 |  |  | 01/14/2022 | 05/15/2024 |  |  | 09/30/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022141336 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/14/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the initial purge solution was identified as being sodium chloride (nacl), which was immediately replaced with dextrose 5 percent in water. the impella was monitored for twenty-four (24) hours to ensure no damage to the device due to the addition of normal saline as a purge solution. it was reported that the patient experienced an episode of ventricular tachycardia and was coded. during chest compressions, it was observed that the impella migrated into the mitral/papillary structures. repositioning allowed the device to return to a good position. it was reported that device weaning was successful and the patient survived the procedure | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
26635 | 1220648-2024-11209 | 19321610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19321610&pc=OZD | 05/05/2024 | 09/11/2024 |  | 05/15/2024 |  |  | 05/05/2024 | 05/15/2024 |  |  | 06/30/2025 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359750 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Great Vessel Perforation | [11 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported the patient undergoing coronary artery bypass graft x 1/aortic valve replacement had an impella 5.5 device implanted prior to coming off bypass for mechanical circulatory support and unexpectedly expired. the patient was on impella 5.5 support via direct approach when there was a sudden, significant output in the chest drain and rapid drop in blood pressure. at bedside, the decision to open the patient's chest was made and revealed the chest cavity full of blood. the bleeding source was noted as the anastomosis of a graft on the aorta. cardiac massage was started, and the patient urgently transferred to the cardiovascular operating room where the graft was stitched. nine units of unpacked red blood cells were given along with internal defibrillation attempted. however, ultimately time of death was called, and the impella pump was turned off as the physician made the decision to withdraw attempts of resuscitating the patient due to the heart muscle not moving and long duration of non-perfusion to rest of the body. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the vascular damage - great vessel issue has been completed since the original report was submitted. the product was not returned for investigation. the patient had an anastomosis of graft between the aorta and the graft which was found to be not stitched correctly leading to likely bleeding. surgical intervention was done. no vascular damage event was confirmed. the root cause of the access site bleeding-major issue was most likely use related due to improper technique. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-11209 was provided in sections b2, b5, h1, and h6. | The instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. added- h6 clinical code 2152, h6 component code 925 h1-type of reportable event should have been death in report #1220648-2024-11209 follow-up #1.
26636 | 1220648-2024-11210 | 19322110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322110&pc=OZD | 05/06/2024 | 06/20/2024 |  | 05/15/2024 |  |  | 05/06/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493731 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported while implanting an impella cp device, the orange ring on the peel away sheath did not break and peel away. it had to be cut with a scalpel to be removed from the catheter. the impella cp device was implanted with no report of patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported difficult removal has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the device removal issue/introducer damage was not determined since product was not returned.
26637 | 1220648-2024-11211 | 19322283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322283&pc=OZD | 01/15/2022 | 05/12/2025 |  | 05/15/2024 |  |  | 01/15/2022 | 05/15/2024 |  |  | 11/30/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-EU |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2022150081 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/15/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced bleeding at the femoral access site after implantation. it was reported to be dry later. it is unknown if there was intervention performed. it was also reported that the patient had electric shocks on refractory atrial fibrillation. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26638 | 1220648-2024-11212 | 19322387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322387&pc=OZD | 05/06/2024 | 06/13/2024 |  | 05/15/2024 |  |  | 05/06/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025441190 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support and experienced limb ischemia. the pump accessed the limb had ischemia diagnosed after implant, and the decision was made to place a perfusion catheter. the team perfused the limb. however, after 24 hours, the team escalated care to an impella 5.5 pump, which was successfully placed. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
26639 | 1220648-2024-11213 | 19322423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322423&pc=OZD | 01/15/2022 | 05/22/2025 |  | 05/15/2024 |  |  | 01/15/2022 | 05/15/2024 |  |  | 10/31/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151879 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/15/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient¿s hemoglobin had dropped due to bleeding from the nose and mouth. the source of the bleeding is unknown. sodium bicarb was added to the purge solution. systemic and purge heparin were discontinued due to bleeding concerns. the bleeding resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26640 | 1220648-2024-11214 | 19322589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322589&pc=OZD | 05/03/2024 | 07/11/2024 |  | 05/15/2024 |  |  | 05/03/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493239 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2024 |  |  |  |  | The user facility reported a 64-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported oozing around the repositioning sheath. preclose suture technique completed, and oozing was resolved. this event is not reportable as there was no lasting harm or injury that necessitated intervention. | The complainant reported that after right femoral access was obtained, contrast injection through sheath to visualize femoral artery revealed retrograde dissection. appeared to be non-flow limiting. a decision was made to obtain contralateral left common femoral arterial access. an aortogram was performed through the left common femoral and showed patent bilateral proximal lower extremity flow. this confirmed the right lower extremity dissection was nonflow limiting. no further intervention on right common femoral dissection. post pci, the patient was transferred to the inpatient unit in stable condition and later was discharged from the hospital. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation for the reported vascular damage has been completed. pump was not returned for investigation. as per the clinical report, the right femoral artery was diagnosed with a dissection before pump insertion. no surgical intervention was needed for this vascular complication. later, pump was placed though left femoral artery. the cause of the vascular damage was not determined and was determined to be unrelated to impella because the dissection occurred to the rfa while the impella was inserted in the lfa. b5 narrative reported in initial report is updated to : the complainant reported that after right femoral access was obtained, contrast injection through sheath to visualize femoral artery revealed retrograde dissection. appeared to be non-flow limiting. a decision was made to obtain contralateral left common femoral arterial access. an aortogram was performed through the left common femoral and showed patent bilateral proximal lower extremity flow. this confirmed the right lower extremity dissection was nonflow limiting. no further intervention on right common femoral dissection. post pci, the patient was transferred to the inpatient unit in stable condition and later was discharged from the hospital. d4-corrected model number f6/f8 should have been blank in the initial report.
26641 | 1220648-2024-11215 | 19322625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322625&pc=OZD | 01/06/2022 | 05/15/2024 |  | 05/15/2024 |  |  | 01/06/2022 | 05/15/2024 |  |  | 10/31/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154860 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/06/2022 |  |  |  |  | The user facility reported an 88-year-old patient presented with severe aortic stenosis (as), aortic valve area (ava) of 0.8 centimeters², was found to have a severely calcified microvascular disease during diagnostic catheterization. it was reported that while on impella cp after looking under fluoro, the long impella sheath appeared to be kinked at the level of entry into the vessel. weaning was successful, the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
26642 | 1220648-2024-11216 | 19322667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322667&pc=PYX | 05/07/2024 | 06/24/2024 |  | 05/15/2024 |  |  | 05/07/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025478605 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/07/2024 |  |  |  |  | The user facility reported a 60-year old male with post-cardiotomy cardiogenic shock was implanted with an impella rp flex for mechanical circulatory support. it was reported that the impella rp flex was being implanted for high risk of surgery, with left atrial appendage closure and aortic valve regurgitation. the 23 french sheath malfunctioned when the medical staff was removing the dilator. the sheath was replaced successfully, and no patient harm has been reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the introducer damage ¿ mechanical issue has been completed since the original report was submitted. the device was not returned and not available for product investigation. the cause of the introducer damage-mechanical issue was not determined since product/images were not returned.
26643 | 1220648-2024-11217 | 19322681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322681&pc=OZD | 01/21/2022 | 05/14/2025 |  | 05/15/2024 |  |  | 01/21/2022 | 05/15/2024 |  |  | 10/31/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158416 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient went into atrial fibrillation and had to be shocked three times. the patient was treated with balloon angioplasty and successfully weaned off support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
26644 | 1220648-2024-11218 | 19322682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322682&pc=OZD | 05/07/2024 | 11/20/2024 |  | 05/15/2024 |  |  | 05/07/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1703855 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/07/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella 5.5 pump experienced flow and suction issues during patient support. repositioning was performed, but the issue remained. the aic (automated impella controller) was subsequently stopped, and the issue resolved. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) pump issue has been completed. data logs from the console confirm that the numerous suction alarms issued by the console. there no observable correlation between the optical signal parameters and the spike of suction level alarms. the console's optical system was inspected. there was no observable defect of the front optical connector. there was no observable issue with the console. the root cause for the suction alarms could not be determined due to insufficient information. added- d9 device return date corrections to the initial mfr report: d4 model number f6/f8 should have been left blank h8-changed to reuse.
26645 | 1220648-2024-11220 | 19322753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322753&pc=OZD | 01/13/2022 | 05/14/2025 |  | 05/15/2024 |  |  | 01/13/2022 | 05/15/2024 |  |  | 09/30/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022147082 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/13/2022 |  |  |  |  | The user facility reported a 54-year-old male on impella cp for mechanical circulatory support. it was reported the patient lost pulses in the left leg. the impella cp device was weaned and explanted, and replaced in the opposite groin. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
26646 | 1220648-2024-11221 | 19322775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322775&pc=OZD | 12/15/2022 | 07/15/2024 |  | 05/15/2024 |  |  | 05/07/2024 | 05/15/2024 |  |  | 04/30/2024 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217143 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 05/07/2024 |  |  |  |  | The research team with chvp- dtu clinical study noted that while the patient was on impella cp support, the right groin had a hematoma. the hematoma did not grow in size and by the time of discharge it was resolved. additionally, the echo (echocardiogram showed the impella shallow. the doctor was notified and the pump was later explanted. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported positioning issue has been completed. pump was not returned for investigation. the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information provided on how the pump was seated shallow. the instructions for use referenced in the initial report is for the impella cp with smartassist system in section: use of echocardiography for positioning of the impella catheter and section: warnings and cautions: cautions d4-corrected model number. f6/f8-should have been left blank in the initial report. g1 reporting contact fax number. h5-changed to yes.
26647 | 1220648-2024-11222 | 19322829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322829&pc=OZD | 05/09/2024 | 07/09/2024 |  | 05/15/2024 |  |  | 05/09/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485183 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient planned for high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support, and difficulty in preparation had been encountered. during preparations for implant, the sidearm on the short sheath would not flush. after multiple failed attempts, the decision was made to utilize the long sheath for implant. there was no report of harm to the patient. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported introducer damage has been completed. the 14x13 introducer was returned for investigation. no physical damage was noted on the introducer. the introducer was not able to flush from the 3-way valve during valve check testing. the catheter was cut near the valve, but the introducer still would not flush. upon further teardown of the valve, glue was blocking the way from the catheter side of the 3-way valve. the cause of the introducer damage issue was a damaged valve, likely related to a manufacturing issue. added- d9 device return date, h6 impact code 4629 d4-corrected model number.
26648 | 1220648-2024-11223 | 19322840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322840&pc=OZD | 05/07/2024 | 07/11/2024 |  | 05/15/2024 |  |  | 05/07/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025471275 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/07/2024 |  |  |  |  | The user facility reported a 56-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient developed a hematoma at the access site. systematic heparin was withheld due to the hematoma and the hematoma reduced. additionally, upon explanting the device, there was still bleeding at the access site. two additional perclose sutures were used and a 7 french mynx was put in with good hemostatic result. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added- h6 impact code 4627, h6 component code 925 f6/f8-should have been left blank in the initial report.
26649 | 1220648-2024-11224 | 19322845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322845&pc=OXD | 01/14/2022 | 05/15/2024 |  | 05/15/2024 |  |  | 01/14/2022 | 05/15/2024 |  |  | 09/30/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OXD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142141 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/14/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient started to have bursts of ventricular tachycardia. it was reported the patient was then experiencing pulseless electrical activity and cardio-pulmonary resuscitation was started and the patient¿s pulse returned. the nurse reported the patient was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26650 | 1220648-2024-11225 | 19322919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322919&pc=OZD | 01/14/2022 | 05/14/2025 |  | 05/15/2024 |  |  | 01/14/2022 | 05/15/2024 |  |  | 09/30/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022135351 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Deliver | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/14/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced delivery issues that were resolved by removal, reinsertion, and ultimately had to be removed again. additionally, the patient experienced peripheral vessel vascular damage that was resolved with multiple balloon inflations and stent. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26651 | 1220648-2024-11226 | 19322971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322971&pc=OZD | 05/11/2024 | 07/15/2024 |  | 05/15/2024 |  |  | 05/11/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453222 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support, and the pump stopped. after several attempts, the pump restarted. there was kinking of the 9f catheter at insertion site into red impella plug. pump stop when that connection manipulated. the impella 5.5 was exchanged without any issues. the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported pump stop and product damage has been completed. no product was returned for evaluation. no data logs were returned for evaluation. clinical details noted that patient was up and mobile with impella implanted via axillary insertion. a pump stop occurred when the patient was moving and a kink in the catheter right at the red handle was noted. when the kink was manipulated, the pump stop was reproduced. the root cause of the pump stop cannot be determined as no product or data logs were returned. added- h6 impact code 4629, h6 med dev prob code 2889 d4-corrected model number.
26652 | 1220648-2024-11219 | 19322986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19322986&pc=OZD | 01/06/2022 | 05/14/2025 |  | 05/15/2024 |  |  | 01/06/2022 | 05/15/2024 |  |  | 10/31/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022145244 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/06/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was found that there was limb ischemia at the impella site. the physician began rapid weaning and explantation of the impella, which resolved the ischemia. the patient remained in stable condition. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patient's pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
26653 | 1220648-2024-11227 | 19323061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19323061&pc=OZD | 05/08/2024 | 06/13/2024 |  | 05/15/2024 |  |  | 05/08/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025488073 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The user facility reported a 78-year old male with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported during the explant of the impella cp to implant an impella 5.5, the patient developed a hematoma on the right groin. the medical staff achieved hemostasis using two perclose sutures, a femoral stop, a jackson-pratt drain, pressure dressings, and a sandbag. the patient survived the event. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26654 | 1220648-2024-11228 | 19323114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19323114&pc=OZD | 02/24/2022 | 05/14/2025 |  | 05/15/2024 |  |  | 02/24/2022 | 05/15/2024 |  |  | 12/31/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171683 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the pigtail was crossed into the left ventricle during implantation, the patient experienced long runs of ventricular arrhythmias. it is unknown how this was resolved, but the device implantation was completed and the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26655 | 1220648-2024-11229 | 19323314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19323314&pc=OZD | 02/12/2022 | 05/27/2025 |  | 05/15/2024 |  |  | 02/12/2022 | 05/15/2024 |  |  | 08/31/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022126643 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  |  |  | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced hemolysis. additionally, the impella dislodged into the aorta and was repositioned into the left ventricle. the hemolysis was determined to be under control after repositioning the device. the patient also experienced thrombus at the inlet/pigtail and the decision was made to explant the device. it was reported that there were no neurological issues after the explant, the renal output was clear/yellow, and the patient remained stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: warnings, cautions, and precautions ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ section: warnings & cautions: warnings ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26656 | 1220648-2024-11230 | 19323984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19323984&pc=OZD | 01/04/2022 | 05/14/2025 |  | 05/15/2024 |  |  | 01/04/2022 | 05/15/2024 |  |  | 10/31/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022150398 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/04/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported the insertion site was oozing. heparin and integrilin were stopped until ptt (partial thromboplastin time) was in range. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer .¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26657 | 1220648-2024-11231 | 19323995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19323995&pc=OZD | 05/09/2024 | 07/09/2024 |  | 05/15/2024 |  |  | 05/09/2024 | 05/15/2024 |  |  |  |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453871 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The user facility reported that a 65 year old, male patient on impella 5.5 support experienced hemolysis. the patient was transfused with five units of packed red blood cells, four units of fresh frozen plasma, two units of platelets, one liter of albumin and one liter of fluid, although the blood products were not for hemolysis specifically. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis has been completed. product was returned for investigation. there was heavy biomaterial wrapped around the impeller blades and a huge chunk of biomaterial was found in the cannula clogging the inlet area of the pump. data logs were returned and there was a motor current spike seen followed by saturated flows. per product evaluation, the root cause of the hemolysis issue was biomaterial found in the cannula and on the impeller bladed. added- d6b, d9 device return date, h6 component code 925 d4-corrected model number d10 concomitant medical products updated . f6/f8 should have been left blank in the initial report. g1-corrected reporting first and last name, and reporting contact email. g2 added health professional.
26658 | 1220648-2024-11232 | 19324325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19324325&pc=OZD | 02/18/2022 | 05/14/2025 |  | 05/15/2024 |  |  | 02/18/2022 | 05/15/2024 |  |  | 11/30/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167232 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/18/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, patient had blue suture hub sutured to skin but oozing was noted at sheath site. tegaderms was placed over the access site. additionally, patient had an echo done which showed device at 5.0 cm so device was repositioned and pulled back to 3.5 cm. urine remained dark red. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26659 | 1220648-2024-11233 | 19324705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19324705&pc=OZD | 01/04/2022 | 05/15/2024 |  | 05/15/2024 |  |  | 01/04/2022 | 05/15/2024 |  |  | 09/30/2023 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144971 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported that when the physician attempted to insert the impella the patient had continuous electrocardiogram changes. the physician could not find a quiet spot for the impella, and the patients pressures dropped while attempting to place the impella. the physician decided to abort implanting the impella and placed a balloon pump instead. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26660 | 1220648-2024-11234 | 19326979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19326979&pc=OZD | 04/15/2024 | 05/29/2024 |  | 05/15/2024 |  |  | 04/15/2024 | 05/15/2024 |  |  | 12/31/2024 |  | 05/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024308935 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Appropriate Term / Code Not Available | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella cp displayed a high purge pressure/system blocked alarm. reducing the p-level did not resolve the issue. tissue plasminogen activator was added to the purge and the motor current was being monitored. the patient expired while on support. there is not enough information to say not our device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported high purge pressure has been completed. no product was returned for investigation. data logs show that just under 2 hours into the case there is a rise in purge pressure and drop in purge flow over 12 minutes. it remains elevated with purge system blocked and hpp alarms for 2 hours before dropping over a span of 20 minutes. 4 and a half hours later there is another rise in purge pressure over 13 minutes, remains high with purge alarms for the remainder of the case. there are no motor current spikes prior to either rise of purge pressure. purge flow ticks stall. the root cause of the high purge pressure issue was most likely a kinked purge line but the exact location could not be determined because product was not returned. the instructions for use states for the impella cp with smartassist for use during cardiogenic shock and high-risk cpi states in section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
26661 | 1220648-2024-11108 | 19308817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19308817&pc=OZD | 01/21/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 01/21/2022 | 05/14/2024 |  |  | 10/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157422 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/21/2022 |  |  |  |  | The complainant reported a patient in st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, a kink was observed in catheter part of the pump under fluoroscopy, and the medical team reported the pump flow was low, and the patient experienced multiple episodes of ventricular tachycardia that required defibrillation and cardiopulmonary resuscitation to resolve. the physician made the decision to explant the impella cp and replace it with a new pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26662 | 1220648-2024-11110 | 19309077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19309077&pc=OZD | 01/21/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 01/21/2022 | 05/14/2024 |  |  | 10/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022153071 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/21/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella 14 french sheath was initially unable to be deployed due to the patient¿s tortuous and calcified vasculature. it was reported the 14 french sheath was replaced with a tapered reposition sheath and the impella insertion was successful. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26663 | 1220648-2024-11111 | 19309194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19309194&pc=OZD | 04/28/2024 | 07/11/2024 |  | 05/14/2024 |  |  | 04/28/2024 | 05/14/2024 |  |  | 06/30/2025 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2024369179 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/28/2024 |  |  |  |  | The complainant reported that during impella cp support for 54-year-old male patient, a leaking purge filter was reported. decision was made to explant the pump. there was no reported patient harm. | [5 empty] | The device was returned for evaluation. upon completion of the investigation, a supplemental mdr will be filed. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the purge leak has been completed. no data logs were returned. the impella was returned for evaluation. the product was inspected and the leak was reproduced in the lab. a crack was noticed at the yellow luer connection that connects tubing to purge reservoir. clinical does not mention any excessive force or improper technique. the root cause of the purge leak - pump was a damaged sidearm yellow luer but the cause of the damage is undetermined. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high risk cpi in section: cleaning. h6 impact code 4627 added d4-corrected model number. f6/f8-should have been left blank. g1-corrected mfr site name. g2-health professional should have been checked.
26664 | 1220648-2024-11112 | 19309250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19309250&pc=OZD | 04/29/2024 | 07/08/2024 |  | 05/14/2024 |  |  | 04/29/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025480765 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Failure to Pump | Pumping Stopped |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support for 78-year-old male patient, the impella cp had high purge pressure for a few minutes and then the pump stopped. the pump was able to be restarted, however, the device was ultimately explanted and a new pump was placed. there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the temporary pump stop has been completed. no data logs were returned. the pump was returned and was inspected and biomaterial was found blocking the purge gap. there were high motor current pump stops when the device was run in the lab. the pump was torn down and biomaterial was confirmed in the motor. the root cause of the temporary pump stop was determined to be biomaterial as confirmed through pump inspection. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high risk cpi in the following sections: section: entering cardiac output. section: troubleshooting the purge system. section: impella stopped. added-d9 device return date, h6 med dev pro code 1491 g1-corrected mfr site name and address h6 med dev pro code changed to 1503
26665 | 1220648-2024-11113 | 19309317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19309317&pc=OZD | 01/21/2022 | 05/14/2025 |  | 05/14/2024 |  |  | 01/21/2022 | 05/14/2024 |  |  | 09/30/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144020 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/21/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the surgeon had some difficulty positioning the impella. the physician stopped, removed the impella, and re-inserted the device to optimize its position. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26666 | 1220648-2024-11114 | 19309363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19309363&pc=OZD | 01/21/2022 | 05/14/2024 |  | 05/14/2024 |  |  | 01/21/2022 | 05/14/2024 |  |  | 09/30/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1477737 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/21/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced limb ischemia with no specified resolution; however, post procedure it was noted by the physician that the patient was doing well. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26667 | 1220648-2024-11115 | 19309417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19309417&pc=OZD | 01/21/2022 | 05/14/2024 |  | 05/14/2024 |  |  | 01/21/2022 | 05/14/2024 |  |  | 09/30/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144039 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/21/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient had consistent oozing at the groin access site. a figure 8 suture was placed with minimal improvement. pressure dressing was applied. ultimately, the decision was made to explant the impella at bedside. perclose x 2 was successfully completed, and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26668 | 1220648-2024-11109 | 19309433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19309433&pc=OZD | 01/21/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 01/21/2022 | 05/14/2024 |  |  | 10/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151897 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/21/2022 |  |  |  |  | The complainant reported a patient with heart failure in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp was accidentally pulled into the aorta. the pump was replaced with a new impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26669 | 1220648-2024-11117 | 19309477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19309477&pc=OZD | 01/21/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 01/21/2022 | 05/14/2024 |  |  | 10/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156375 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported the impella cp had alarms for purge pressure low and purge flow high. the medical team was advised to check all the connections and de-air the system, but this did not resolve the alarms. the medical team was instructed to change the cassette and purge bag. the purge solution was changed from d5w + 25u/ml heparin to d10 + 25u/ml heparin. no further issues were reported, and the medical team was instructed to monitor the purge system for any further alarms. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26670 | 1220648-2024-11118 | 19309526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19309526&pc=OZD | 01/21/2022 | 05/20/2025 |  | 05/14/2024 |  |  | 01/21/2022 | 05/14/2024 |  |  | 10/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023159236 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and while on support the labs were hemolyzed. hemolyzed labs were persistent despite drawing blood from multiple sources. urine output was brown and platelets down significantly. therefore, the impella was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26671 | 1220648-2024-11119 | 19309591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19309591&pc=OZD | 08/26/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 08/26/2022 | 05/14/2024 |  |  | 04/30/2024 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218255 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced oozing at the groin site, and they were unable to go over the p4 setting despite adding blood volume. in addition, there was a retroperitoneal bleed present, for which they administered 4 units of packed red blood cells along with fresh frozen plasma, which achieved the desired effects. the patient was taken off heparin due to the retroperitoneal bleed, and the patient remained on impella support until they were successfully weaned. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26672 | 1220648-2024-11116 | 19309668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19309668&pc=OZD | 01/21/2022 | 05/23/2025 |  | 05/14/2024 |  |  | 01/21/2022 | 05/14/2024 |  |  | 09/30/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022137776 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/21/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support the patient experienced an episode of atrial fibrillation, however pressures remained adequate. it was reported that the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26673 | 1220648-2024-11120 | 19309791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19309791&pc=OZD | 01/21/2022 | 05/14/2024 |  | 05/14/2024 |  |  | 01/21/2022 | 05/14/2024 |  |  | 09/30/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022139869 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the right lower extremities started to become compromised, an external bypass procedure was performed to provide more blood flow to the lower leg. there was improvement in the extremities after the bypass. the patient was sent to unit on impella support. it was noted that the physician chose to explant the impella due to distal limb ischemia. the device was removed, balloon tamponade with manual compression performed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26674 | 1220648-2024-11121 | 19309858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19309858&pc=OZD | 08/28/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 08/28/2022 | 05/14/2024 |  |  | 03/31/2024 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207308 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, there were intermittent alarms for suction. the medical team temporarily decreased p-level to p4 and administered normal saline before going back up to p6. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26675 | 1220648-2024-11122 | 19309896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19309896&pc=OZD | 01/22/2022 | 05/14/2024 |  | 05/14/2024 |  |  | 01/22/2022 | 05/14/2024 |  |  | 07/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022116641 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/22/2022 |  |  |  |  | The user facility reported a 58-year-old patient needing mechanical circulatory support. it was reported that the patient was placed onto impella cp support and while on support, the patient experienced intermittent ventricular fibrillation (vf) with no specified resolution. the patient also experienced slight access site bleeding with no reference for any intervention needed. additionally, the pump experienced intermittent, ongoing suction with no specified resolution. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
26676 | 1220648-2024-11123 | 19310015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310015&pc=OZD | 01/21/2022 | 05/14/2025 |  | 05/14/2024 |  |  | 01/21/2022 | 05/14/2024 |  |  | 09/30/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140459 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, impella alarms for pump positioning were noted. it was stated that the pump was in good position under transthoracic echocardiogram and fluoroscopy. the decision was made to remove the initial impella and implant a new device to avoid the potential for clot ingestion. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26677 | 1220648-2024-11124 | 19310181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310181&pc=OZD | 01/22/2022 | 05/14/2024 |  | 05/14/2024 |  |  | 01/22/2022 | 05/14/2024 |  |  | 10/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023159242 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/22/2022 |  |  |  |  | The user facility reported a 78-year-old patient needing mechanical circulatory support. it was reported that the patient was on impella cp support due to stemi (st-segment elevation myocardial infarction). while on support no palpable pedal pulse was noted; however, color doppler did show flow. the following day, while on support, still no palpable pedal pulse, only doppler pulses. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26678 | 1220648-2024-11125 | 19310230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310230&pc=OZD | 01/22/2022 | 05/14/2024 |  | 05/14/2024 |  |  | 01/22/2022 | 05/14/2024 |  |  | 11/30/2022 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021016976 |  | 0048-0032 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/22/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced an infection in the process. the clinical staff conversed to the point of thinking the infection was from an ischemic bowel; however, it was never confirmed nor denied as the pending computed tomography scan results were not disclosed in the patient notes. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
26679 | 1220648-2024-11126 | 19310278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310278&pc=OZD | 01/27/2022 | 05/14/2025 |  | 05/14/2024 |  |  | 01/27/2022 | 05/14/2024 |  |  | 10/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156247 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/27/2022 |  |  |  |  | The complainant reported the patient was placed onto impella cp support for pre-operational placement for high-risk surgery. while on support, the patient experienced a thrombus that was resolved by aspiration and stenting. additionally, the patient experienced unspecified arrhythmias with no resolution specified. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | Thrombus. analysis summary the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. cardiac arrythmia. analysis summary: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical. information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
26680 | 1220648-2024-11127 | 19310310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310310&pc=OZD | 04/26/2024 | 07/16/2024 |  | 05/14/2024 |  |  | 04/29/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025488752 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Unspecified Infection | [11 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the 73-year-old male patient was in septic shock and experienced hemolysis since impella implantation. the lactate dehydrogenase (ldh) was noted to be very high and four blood bags were administered. care was withdrawn after five days of support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the hemolysis and the infection has been completed. product was not returned for investigation. data logs were returned and investigated. there were no motor current spikes, suction events or positioning issues seen in the data logs. the clinical information mentions that the patient was in a very poor general condition and logs show placement signal trending downwards. therefore, the root cause of the hemolysis issue was most likely patient condition related. the root cause of the infection could not be determined without additional clinical details. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b2-selected death and added date of death, h6 impact code 1802 h1 changed to death g1-corrected mfr site name and address.
26681 | 1220648-2024-11128 | 19310382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310382&pc=OZD | 01/30/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 01/30/2022 | 05/14/2024 |  |  | 11/30/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161826 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Insufficient Flow or Under Infusion | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, ventricular tachycardia was noted. the patient was then cardioverted, resulting in suction/low flow events when the patient's blood pressure dipped. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26682 | 1220648-2024-11129 | 19310392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310392&pc=OZD | 01/23/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 01/23/2022 | 05/14/2024 |  |  | 10/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149195 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2022 |  |  |  |  | The complainant reported a patient noted as high-risk surgery was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed large hematoma. manual pressure was applied for five (5) minutes, followed by femstop application at belt pressure to resolve hematoma. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26683 | 1220648-2024-11130 | 19310400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310400&pc=OZD | 04/26/2024 | 07/16/2024 |  | 05/14/2024 |  |  | 04/29/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025493688 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for 64-year-old male patient, the 14f cannula was bent close to the motor housing. there was no delivery issue prior to damage. the pump was replaced and no patient harm was reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the reported product damage has been completed. product was returned for investigation. there was a cannula kink observed near the motor housing end of the cannula. clinical information about patient dilation prior to implant, vessel condition, excessive force applied has not been provided. therefore based on the clinical information provided and product evaluation, the root cause of the product damage issue was a kinked cannula due to improper technique. the instructions for use referenced in the initial report is for the impella cp with smartassist system in section: warnings & cautions: cautions. g1-corrected mfr site name and address
26684 | 1220648-2024-11131 | 19310446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310446&pc=OZD | 02/01/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 02/01/2022 | 05/14/2024 |  |  | 10/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023159617 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient began experiencing bleeding at the impella access site and their activated clotting time was elevated. heparin was stopped, a femstop was placed, and the patient received blood products. additionally, a clot was discovered in the left ventricle with an echocardiogram. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26685 | 1220648-2024-11132 | 19310463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310463&pc=OZD | 01/28/2022 | 03/07/2025 |  | 05/14/2024 |  |  | 01/28/2022 | 05/14/2024 |  |  | 11/30/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162588 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced bleeding at the anastomoses of the graft insertion site. hemostasis was obtained with additional sutures and blood products. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
26686 | 1220648-2024-11133 | 19310471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310471&pc=OZD | 04/30/2024 | 06/11/2024 |  | 05/14/2024 |  |  | 04/30/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025451656 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support as escalation of care from an impella cp device. the impella 5.5 could not completely be advanced through the subclavian artery. implant was aborted due to tortuous patient anatomy. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist for use during cardiogenic shock in section: warnings & cautions: warnings. in accordance with updated procedures, sections d6 have been revised from the initial submission.
26687 | 1220648-2024-11134 | 19310564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310564&pc=OZD | 02/04/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 02/04/2022 | 05/14/2024 |  |  | 08/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022132672 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/04/2022 |  |  |  |  | The user facility reported a 61-year-old male on impella 5.5 for mechanical circulatory support. while under impella 5.5 support, heparin was taken out of purge and systemically due to the patient experiencing heparin induced thrombocytopenia. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
26688 | 1220648-2024-11135 | 19310572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310572&pc=OZD | 04/30/2024 | 07/11/2024 |  | 05/14/2024 |  |  | 04/30/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103168 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | Activation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella aic (automated impella controller) prior to the case starting received the following message "console failure on start-up". aic was turned off and a new one was used. there was no harm to the patient. | [5 empty] | The impella device was received from the customer on (b)(6) 2024 and therefore, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console controller failure alarm has been completed. data logs were reviewed which confirm that the yellow controller alarm was issued on the day of the complaint. in addition, there were loss of communication with keypad or between the kbd and i2c for kbd hw revision 1" alarms. the console was returned for evaluation. the console was internally inspected and ribbon cables were found to be intact, suggesting no physical damage. power supply was inspected and found to be within specification. the failure mode could not be replicated. the root cause yellow alarm could not be determined due to not being able to reproduce. added- d9 device return date d4-corrected model number. f6/f8-should have been left blank in the initial report.
26689 | 1220648-2024-11136 | 19310619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310619&pc=OZD | 01/23/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 01/23/2022 | 05/14/2024 |  |  | 11/30/2022 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021020239 |  | 0048-0032 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a loss of pedal pulses in both legs. additionally, the patient experienced a gastrointestinal bleed requiring blood products to be administered. patient's condition was stable post procedure | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event (limb ischemia) are as follows: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ impella cp with smart assist system instructions for use for the related event (vascular dissection) are as follows: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26690 | 1220648-2024-11137 | 19310624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310624&pc=OZD | 05/03/2024 | 06/28/2024 |  | 05/14/2024 |  |  | 05/03/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025449237 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2024 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that following impella cp implant, a hematoma developed proximal to the sheath. manual pressure was initially applied to manage the condition, but the decision was ultimately made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
26691 | 1220648-2024-11138 | 19310679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310679&pc=OZD | 05/05/2024 | 07/03/2024 |  | 05/14/2024 |  |  | 05/05/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494616 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support, and during support, the patient's plasma free hemoglobin levels became elevated. the p- level was initially reduced in response to the suspected hemolysis, in which the implanting physician suspected to be from delayed start of anti-coagulation. the decision was ultimately made to explant the impella pump, and post removal of the impella pump, an intra-aortic balloon pump was placed. the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis has been completed. data logs were reviewed which showed the pump ran without issue during support. there were some intermittent suction alarms during the case but resolved quickly. the product was not returned for evaluation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
26692 | 1220648-2024-11139 | 19310741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310741&pc=OZD | 05/06/2024 | 06/26/2024 |  | 05/14/2024 |  |  | 05/06/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025494694 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2024 |  |  |  |  | The user facility reported a 84-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an attempted impella cp implant and it was unsuccessful. the physician made the decision to abort the implant due to vascular anatomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues been completed since the original report was submitted. the device was not returned for investigation. the root cause of delivery issue is determined to be patient condition because of the vascular anatomy of the patient. g3.
26693 | 1220648-2024-11140 | 19310804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310804&pc=OZD | 01/28/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 01/28/2022 | 05/14/2024 |  |  | 10/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022148224 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was experiencing oozing from all sites and elevated activated clotting time. heparin was removed from the purge, and the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26694 | 1220648-2024-11141 | 19310831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310831&pc=OZD | 01/22/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 01/22/2022 | 05/14/2024 |  |  | 10/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156924 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/22/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, a slight hematoma was observed. manual pressure was initially held and the hematoma appeared to be soft and contained. the patient¿s hemoglobin was noted to be low, and two units of blood was transfused. the impella cp was successfully weaned and explanted. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26695 | 1220648-2024-11142 | 19310877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310877&pc=OZD | 05/07/2024 | 07/11/2024 |  | 05/14/2024 |  |  | 05/07/2024 | 05/14/2024 |  |  | 09/30/2025 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403895 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/07/2024 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported that a patient implanted with an impella cp had a hematoma. an angiogram was performed and retroperitoneal bleed on the impella side was revealed. the impella was removed, the bleed was tamponade, and mass transfusion protocol was started. pressers were weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torqueing the impella, adjustments should be performed under imaging guidance. ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). | The investigation for the reported access site bleed and vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and vascular damage could not be determined. the instructions for use referenced in the initial report is for the impella cp with smart assist system in sections: section: general patient care considerations. section: entering cardiac output. section: potential adverse events (united states) d4-corrected model number.
26696 | 1220648-2024-11143 | 19310882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310882&pc=OZD | 04/29/2024 | 06/25/2024 |  | 05/14/2024 |  |  | 04/29/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025478938 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the 78-year-old male patient experienced massive bleeding problems at the puncture site and others. the angle of the impella was adjusted and the bleeding subsided. protamine was also administered. the device was not removed as a result of the failure mode. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved by changing the impella angle.
26697 | 1220648-2024-11144 | 19310903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310903&pc=OZD | 05/08/2024 | 07/09/2024 |  | 05/14/2024 |  |  | 05/08/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493129 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while the impella cp was being explanted, a blood clot was seen through the sheath despite aspirating. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the thrombus has been completed. the impella product was returned and was visually inspected and no abnormalities were found. without return of data logs nor additional clinical details, the root cause of the thrombus could not be determined. added- d9 device return date f6/f8 should have been left blank in the initial report.
26698 | 1220648-2024-11145 | 19310916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310916&pc=OZD | 01/04/2022 | 07/16/2024 |  | 05/14/2024 |  |  | 04/29/2024 | 05/14/2024 |  |  | 09/30/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2022131525 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility |  |  | 76 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant from the study reported that during impella cp support, the patient experienced post percutaneous coronary intervention (pci) hemolysis, that was moderate in severity. the plasma free hemoglobin was increased up to 266 mg/l. the patient is noted to have recovered. the event was determined to be possibly related to the device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported hemolysis has been completed. the product and data were not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ b.3 revised as previously entered incorrectly on manufacturer device report 1220648-2024-11145. e.1 address line 1 was revised as previously entered incorrectly on manufacturer device report 1220648-2024-11145. g.1 revised manufacturing site facility name, address line 1 and city as previously entered incorrectly on manufacturer device report 1220648-2024-11145. g.2 revised as selection was inadvertently omitted from manufacturer device report 1220648-2024-11145. h.6 code 4643 was reported incorrectly on manufacturer device report 1220648-2024-11145. h.10 additional manufacturer narrative instructions for use were revised from manufacturer device report 1220648-2024-11145 in accordance with updated procedures.
26699 | 1220648-2024-11146 | 19310933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310933&pc=OZD | 01/30/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 01/30/2022 | 05/14/2024 |  |  | 10/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157610 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/30/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hematoma and oozing at the access site. the hematoma was unsuccessfully managed with manual compression and femstop, so the patient was given packed red blood cells. an additional impella cp was placed in the left groin, and the impella cp in the right groin was removed. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26700 | 1220648-2024-11147 | 19310956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310956&pc=OZD | 04/30/2024 | 07/15/2024 |  | 05/14/2024 |  |  | 04/30/2024 | 05/14/2024 |  |  | 08/31/2025 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024402918 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | VA ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the purge side arm was defective and a leak was observed. additionally, the pump stopped due to high motor current. there was no reported patient harm. | Additional information was provided which states the physician decided to remove the impella cp and keep the patient on va ecmo support only. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the pump stop and purge leak has been completed. during support, there was a purge leak and a high motor current pump stop. no product or data logs were returned. the root cause of the purge leak - pump was most likely due to a damaged sidearm but the cause of the damage was not determined. the root cause of the pump stop was not determined because no product or data logs were returned. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high risk cpi in the following sections: section: using the automated impella controller with the impella catheter. section: troubleshooting the purge system. section: impella stopped. it also states in impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ added-b5 addl narrative, d10-concomitant device, h6 impact code 4627, 4642 g1-corrected mfr site name and address.
26701 | 1220648-2024-11148 | 19310958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310958&pc=OZD | 05/08/2024 | 07/12/2024 |  | 05/14/2024 |  |  | 05/09/2024 | 05/14/2024 |  |  | 11/30/2025 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025437930 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility |  |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The user facility reported a male patient of an unknown age was implanted with an impella cp device for mechanical circulatory support. it was reported that bleeding occurred in patient while on impella cp support. the bleeding was not only at the impella access site, but all over the body. there was bleeding in the thoracic cavity from the time of admission due to trauma. the bleeding in the thoracic cavity was increasing and blood transfusions were required. heparin was reduced but bleeding did not get better. the impella was removed, and bleeding improved. the patient was reported to be in serious condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for access site bleeding and vascular damage has been completed. the pump was not returned for investigation. bleeding was reported from the impella access site and other body areas. bleeding did not improve even after stopping heparin. additionally, bleeding was reported in the thoracic cavity when the patient came for treatment, and thoracic bleeding increased. the cause of the access site bleeding issue was most likely patient condition related since bleeding was reported from multiple sites. the cause of the vascular damage was determined to be unrelated to the impella because the patient already had thoracic bleeding before the pump was placed. added g2 report source- foreign d4 model number corrected f6/f8-should have been left blank in the initial report.
26702 | 1220648-2024-11149 | 19310988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310988&pc=OZD | 05/03/2024 | 06/19/2024 |  | 05/14/2024 |  |  | 05/03/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025485251 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2024 |  |  |  |  | The user facility reported a 74-year-old female patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that an impella cp pump was unable to cross the patient's valve during attempted delivery due to severe aortic stenosis. the implant was subsequently aborted. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings & cautions: to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition since the pump could not be placed due to severe aortic stenosis.
26703 | 1220648-2024-11150 | 19310993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19310993&pc=OZD | 05/09/2024 | 06/26/2024 |  | 05/14/2024 |  |  | 05/09/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025481405 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The user facility reported an 84-year-old female patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that during support, the patient access site bleeding. the patient received 2 units of packed red blood cells as treatment and additional sutures and femstop were used to achieve hemostasis. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved by adding additional suture.
26704 | 1220648-2024-11151 | 19311037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19311037&pc=OZD | 05/04/2024 | 06/26/2024 |  | 05/14/2024 |  |  | 05/04/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025458665 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp/5.5 device for mechanical circulatory support. the patient developed ischemia during impella cp support. when contrast was injected through the re-access port, flow was noted to be sluggish down the leg. as a result of the ischemia, the decision was made to explant the pump. pulses were detectable following pump explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿ | The investigation into the access site bleeding ¿ minor and the limb ischemia ¿ reversible issues been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
26705 | 1220648-2024-11152 | 19311051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19311051&pc=OZD | 03/14/2024 | 07/18/2024 |  | 05/14/2024 |  |  | 05/05/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025446003 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/05/2024 |  |  |  |  | Complainant in japan was on impella cp smart assist support for non-st-elevation myocardial infarction and left coronary artery aneurysm. during support, the patient developed hemolysis. the impella cp was removed due to the hemolysis and the need for early intervention for cardiac rehab and the patient was escalated to the impella 5.5. postoperatively, the patient developed acute kidney injury and started continuous renal replacement therapy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events include: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ | The investigation for the hemolysis and renal failure has been completed. no product was returned for evaluation. data logs were reviewed and no positioning issues or ingestion of biomaterial was noted. the root cause of the renal failure cannot be determined due to limited clinical information. d4-corrected model number f6/f8 should have been left blank in the initial report.
26706 | 1220648-2024-11153 | 19311064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19311064&pc=OZD | 05/07/2024 | 06/21/2024 |  | 05/14/2024 |  |  | 05/07/2024 | 05/14/2024 |  |  | 05/31/2025 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024346464 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/07/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as escalation of care from an impella cp. the patient was given blood products due to down trending hemoglobin and platelet. there was concern over impella 5.5 interaction with blood. the patient is stable and remains on impella support on the date of this report. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported thrombocytopenia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombocytopenia could not be determined. d4-corrected udi number.
26707 | 1220648-2024-11154 | 19311067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19311067&pc=OZD | 05/09/2024 | 06/24/2024 |  | 05/14/2024 |  |  | 05/09/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025478381 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The complainant reported a patient of unknown age and gender had an attempted impella 5.5 implant for mechanical circulatory support. medical staff had issues having the pump connect ¿ an optical sensor system error was showing. the staff disconnected/reconnected the sensor, which resolved the issue. approximately 50 minutes later and still prior to the implant, the team received a catheter disconnected error. the staff then opted to use a different pump for the implant. no patient harm has been reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into the pump not detected issue has been completed since the original report was filed. the pump was returned for investigation. no evidence of damage was noted on the returned product. the pump passed the laser continuity test, and all 5s optical parameters were within specification. data logs show the placemat signal was not reliable, with zero snr values during the start of the case. the cause of the optical signal failure was an air gap from between the aic and the patient cable plug or within the middle of the red handle ferrule.
26708 | 1220648-2024-11155 | 19311082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19311082&pc=OZD | 01/23/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 01/23/2022 | 05/14/2024 |  |  | 10/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156971 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2022 |  |  |  |  | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced heparin-induced thrombocytopenia (hit). in addition, there were high purge pressure alarms noted on impella. as a result, patient was put on argatroban. the impella was successfully weaned and explanted. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26709 | 1220648-2024-11156 | 19311162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19311162&pc=OZD | 05/08/2024 | 06/13/2024 |  | 05/14/2024 |  |  | 05/08/2024 | 05/14/2024 |  |  | 10/31/2025 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024418220 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and upon inserting the pump, the patient had ventricular fibrillation (vf). the 5.5 was attempted to be repositioned in the left ventricle and the patient had multiple runs of vf. they were shocked during each episode. the pump was repositioned successfully and no further issues were noted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation into the vf/vt/af/at (v-fib) issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
26710 | 1220648-2024-11157 | 19311196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19311196&pc=OZD | 05/08/2024 | 06/24/2024 |  | 05/14/2024 |  |  | 05/08/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450325 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Defective Device | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was inserted per protocol with no difficulty. support had been going for approximately 45 minutes with stable hemodynamics when patient then acutely lost pulsatility. a hematoma in the right groin developed. the groin swelled very quickly despite pressure and patient became acutely unstable. volume was administered and four units of packed cells. the physician gained contralateral access and was able to use an omni catheter to get up and over the bifurcation and place a wire. the vascular surgeon joined the case at this point and was able to use an 8 mm balloon to occlude the arteriotomy. patient still hemodynamically unstable until rapid infusion of four units of packed cells completed. impella had to be removed to manage the vascular complication. x-ray imaging showed a kink in the 14 french sheath body. the kink caused the 14 french sheath body to crack and allow for a large hematoma to develop. the cp was removed and per close deployed successfully with balloon tamponade. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation into the vascular damage - peripheral vessel and the introducer damage ¿ mechanical issues has been completed since the original report was submitted. the impella device and the introducer were not returned for investigation. it was reported that during removal of the sheath active bleeding was observed out of the crack in the body of sheath prior to peeling it away. the cause of the vascular damage issue was most likely an introducer sheath split along the score lines due to excessive manipulation of sheath during support per clinical.
26711 | 1220648-2024-11158 | 19311224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19311224&pc=OZD | 05/09/2024 | 07/10/2024 |  | 05/14/2024 |  |  | 05/09/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025453690 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support, which lasted just under 18 hours. during the 18 hours, the patient was transferred to the tertiary center for the care escalation to an impella 5.5 pump. the impella cp pump was placed via the femoral artery through a stent that was delivered to native anatomy that was heavily calcified. the team explanted the impella cp and was placing the perclose when an observation was made that the femoral stent had migrated into the distal aorta. vascular surgery team involvement remedied the issue. the stent was snared from aorta, new stent placed, and groin closed via a cutdown. of note, the physician stated the fault of stent migration was not due to impella cp device, but rather the stent being under deployed. | [5 empty] | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system. section: warnings & cautions: warnings: avoid manual compression of the inlet and outlet areas of the cannula assembly. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the reported insertion issues has been completed. per the clinical information, the stent was implanted due to occluded iliac, followed by successful pump implant. after the medical procedure was completed, the pump was explanted and while using preclose to close the access site, the stent was dislodged to aorta while re-wiring. the failure is best captured under access site bleeding. the root cause of the access site bleeding issue was most likely use related. d3 manufacturer fax number missing. d4-corrected model number.
26712 | 1220648-2024-11159 | 19311257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19311257&pc=OZD | 05/09/2024 | 06/24/2024 |  | 05/14/2024 |  |  | 05/09/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025466192 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced access site bleeding. the bleed at the access site prompted a hematoma, which was treated by the infusion of one unit of blood. the support remains on and has not been explanted to date of this report. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation into the access site bleeding major issue has been completed since the original report was submitted. the product was not returned for investigation. the root cause of access site bleeding is determined to be patient movement because the doctor mentioned that the vasculature was most likely damaged when the patient kept moving his leg.
26713 | 1220648-2024-11160 | 19311937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19311937&pc=OZD | 08/16/2022 | 03/10/2025 |  | 05/14/2024 |  |  | 08/16/2022 | 05/14/2024 |  |  | 10/31/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022148677 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/16/2022 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, a hematoma was identified at the access site, requiring a blood transfusion to be administered to the patient in order to resolve this issue. the physician then reported being unable to unload as expected with the impella due to significant aortic regurgitation (ar). it was stated that after repositioning, there was no improvement in the ar. weaning was successful and the medical staff replaced the impella with an intra-aortic balloon pump (iabp). it was reported that the patient survived the procedures. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding/hematoma: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | Corrected information provided in b3, g3. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26714 | 1220648-2024-11161 | 19311972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19311972&pc=OZD | 05/09/2024 | 06/27/2024 |  | 05/14/2024 |  |  | 05/09/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025489406 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The complainant reported an 83-year-old male patient with a high risk of percutaneous coronary intervention acute was implanted with a impella cp for mechanical circulatory support. the complainant reported an access site bleed. medical staff remedied the bleed by sutures and a femoral stop. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation into the access site bleeding ¿ major issue been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved by adding additional 8 figure suture. g3.
26715 | 1220648-2024-11162 | 19312002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19312002&pc=OZD | 05/09/2024 | 07/09/2024 |  | 05/14/2024 |  |  | 05/09/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025473999 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The complainant reported a 65-year-old male patient with post-cardiotomy cardiogenic shock was implanted with a impella 5.5 for mechanical circulatory support. the complainant reported a high purge pressure alarm. the purge cassette was changed with no resolution. medical staff applied the tissue plasminogen activator protocol to resolve the issue. the patient remains on support. | Pump explanted and the patient survived. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the reported high purge pressure has been completed. product was not returned for investigation. data logs show a gradual rise in purge pressure over two hours with purge pressure high alarms. this trend continued for 14 hours before returning to normal specification limits. the doctor initiated tissue plasminogen activator after observing the high purge pressure. the cause of the high purge pressure issue was most likely biomaterial. added- b5 and d6b f6/f8 should have been left blank in the initial report.
26716 | 1220648-2024-11163 | 19312158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19312158&pc=OZD | 04/14/2024 | 06/21/2024 |  | 05/14/2024 |  |  | 04/14/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2025459067 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Malfunction | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 support for 66-year-old male, the pump stopped after hours of high purge pressure and low pump flow. the patient had undergone lysis therapy which had been unsuccessful. therefore, the pump was explanted. care was withdrawn and the patient expired. there is not enough evidence to say not our device that contributed to the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into the pump stop issue has been completed since the original report was submitted. the device was not returned for investigation. the data logs confirmed high purge pressure and low purge flow that did not resolve and there were also purge system blocked alarms before the pump stop. the root cause of pump stop was most likely high purge pressure, but the cause of the high purge pressure was unknown because no product was returned.
26717 | 1220648-2024-11164 | 19312472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19312472&pc=OZD | 02/16/2024 | 05/14/2024 |  | 05/14/2024 |  |  | 04/15/2024 | 05/14/2024 |  |  | 11/30/2025 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025436562 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp support, there was a sudden increase in purge flow, and normalization of purge flow but shortly after an increase in motor current and left ventricle (lv) curve completely over the aortic wave form. investigation showed saturated flow. the user was advised of increased risk of pump stop. care was withdrawn and the patient expired. there is not enough information to associate the use of the device with the patient's outcome. | [5 empty] | The investigation into the reported 'saturated flow" has been completed. the product was returned for investigation. the pump was visually inspected and found that the pump had significant biomaterial stuck outside/below the outflow cage and motor housing was observed. also, large purge gap was observed with the bearing wear occurring at the end on the 12th day of support. the biomaterial likely expelled out from the impeller due to large purge gap since the pump was used for greater than design life of eight days (12 days in support). there was no optical signal issue observed with the return pump. there were no other abnormalities found. the data logs were reviewed and motor current spike followed by saturated flows on the 10th day, 40 hours before explant was observed. saturated flows were seen at p-8 and other p-levels after motor current spike. no optical signal issues observed in data logs. the cause of the saturated flow issue was due to biomaterial ingestion per log and return product review.
26718 | 1220648-2024-11165 | 19312544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19312544&pc=OZD | 08/21/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 08/21/2022 | 05/14/2024 |  |  | 05/31/2024 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229038 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, the patient experienced a hematoma at the access site. blood products were given, and the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26719 | 1220648-2024-11166 | 19312582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19312582&pc=OZD | 08/23/2022 | 05/14/2024 |  | 05/14/2024 |  |  | 08/23/2022 | 05/14/2024 |  |  | 04/30/2024 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221594 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a small amount of oozing from the groin site. the reporter did not indicate how the oozing was resolved. however, the complainant reported the patient was stable and they were preparing the patient for weaning. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26720 | 1220648-2024-11168 | 19313034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19313034&pc=OZD | 04/19/2024 | 06/21/2024 |  | 05/14/2024 |  |  | 04/19/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025447332 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/19/2024 |  |  |  |  | The user facility reported an 83-year-old female noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that excessive bleeding was noted around the repositioning sheath during impella cp support. in response to the noted condition, the decision was made to explant the pump. the patient was reported stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. h.6 code 4114 was reported incorrectly on the previously submitted report as the device was returned. h.10 additional manufacturer narrative on the previously submitted report stated that the device was not returned but it was returned. | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the repositioning sheath was oozing blood at the posterior of the sheath where the tuohy would be locked. the product was inspected and a leak test was performed which failed. it was discovered that the o-ring valve was damaged with an abnormal groove along the side, which compromised the hemostatic seal by creating a leak path. the root cause of the access site bleeding - major was determined to be a damaged valve. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures.
26721 | 1220648-2024-11169 | 19313159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19313159&pc=OZD | 05/09/2024 | 06/26/2024 |  | 05/14/2024 |  |  | 05/09/2024 | 05/14/2024 |  |  |  |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494618 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2024 |  |  |  |  | The complainant reported a 70-year-old female patient with acute myocardial infarction / cardiogenic shock was implanted with a impella cp for mechanical circulatory support. the complainant reported that the patient was taken to the operating room for escalation to an impella 5.5. vascular surgery was consulted to assist with the removal of the impella cp ¿ a clot in the descending aorta near the catheter was noted by anesthesia during the pre-procedure transesophageal echocardiogram. a clot in an appendage was also noted. medical staff decided to perform a bilateral femoral cutdown with a manual removal of the catheter. the device was removed successfully along with all material and subsequent vascular procedures. the patient survived the event. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation into the thrombus issue been completed since the original report was submitted. the device was not returned for investigation. as the necessary information was not provided, the root cause of the thrombus was not determined.
26722 | 1220648-2024-11170 | 19313203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19313203&pc=PYX | 08/14/2021 | 06/03/2024 |  | 05/14/2024 |  |  | 05/02/2024 | 05/14/2024 |  |  | 04/30/2021 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1435321 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility |  |  |  | 40 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/02/2024 |  |  |  |  | The complaints team was forwarded a publication entitled, "use of right ventricular assisted device for right heart failure in a patient with acute respiratory distress syndrome" in which there was a neurological event with possible relationship to impella. the patient had a computed tomography scan performed and the brain was consistent with cerebral edema and multifocal hemorrhages in bilateral cerebral hemispheres, with the largest lesion measured approximately one centimeter. the cause of the cerebral edema and multifocal hemorrhages is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | G.2 revised as selection was omitted from the previously submitted report. | Citation: jahngir, m. u., & nabizadeh-eraghi, p. (2021). use of right ventricular assisted device for right heart failure in a patient with acute respiratory distress syndrome. cureus. https://doi.org/10.7759/cureus.17671 h.6 added code 925 as it was inadvertently left off the previously submitted manufacture device report. code 4114 was reported incorrectly on the previously submitted report. new code has been added to type of investigation codes h.10 additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. literature citation was omitted from the previously submitted reports. attachments - literature attachment was omitted from the previously submitted reports.
26723 | 1220648-2024-11171 | 19313862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19313862&pc=OZD | 08/07/2015 | 06/20/2024 |  | 05/14/2024 |  |  | 05/02/2024 | 05/14/2024 |  |  | 03/31/2017 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1190507 |  | 0048-0003 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | IMPELLA RP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | A publication entitled: ¿severe mitral regurgitation and biventricular heart failure successfully treated with biventricular percutaneous axial flow pumps as a bridge to mitral valve surgery" reported a 61-year-old male of unknown ethnicity in which a bipella rp-cp support was used. the impella cp was explanted for hemolysis, and replaced by intra-aortic balloon pump (iabp). support was proceeded with the impella rp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. therefore, the root cause of the hemolysis issue was not determined since product was not returned, data logs were not returned, and insufficient clinical information was provided.
26724 | 1220648-2024-11172 | 19314090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19314090&pc=OZD | 08/24/2022 | 06/02/2025 |  | 05/14/2024 |  |  | 08/24/2022 | 05/14/2024 |  |  | 01/31/2024 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023186668 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was an ongoing placement signal low alarm that transitioned to the correct impella position in the ventricle. the patient remained on impella support as a bridge to transplant. no further information was available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26725 | 1220648-2024-11173 | 19314167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19314167&pc=PYX | 11/11/2023 | 06/13/2024 |  | 05/14/2024 |  |  | 04/23/2024 | 05/14/2024 |  |  | 03/31/2025 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024388958 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating an allegation of pump flow: ¿in or, rvad stopped flowing. surgeon opened chest and started manual compressions. it was identified that rp impella clotted and unable to exchange the circuit secondary to hemodynamic instability." | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the low pump flow has been completed. according to the clinical, on the second day of impella rp support the pump stopped flowing. after opening the patient's chest and beginning manual compressions it was discovered that the pump was clotted. the decision was then made to explant the pump and replace it with another device. no product was returned for a visual inspection. data log analysis confirmed a sudden motor current spike followed by reduced flows with no change in p-level. the most likely cause of the low pump flow was biomaterial. the instructions for use referenced in the initial report is for the impella rp with smartassist system in section: using the automated impella controller with the impella catheter and section: using the automated impella controller with the impella rp with smartassist system catheter. d4-updated model number g1-corrected mfr site name and address in accordance with updated procedures, section h10 has been revised from the initial submission.
26726 | 1220648-2024-11174 | 19314900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19314900&pc=OZD | 04/15/2022 | 02/24/2025 |  | 05/14/2024 |  |  | 04/15/2022 | 05/14/2024 |  |  | 01/31/2024 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182849 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2022 |  |  |  |  | The user facility reported a 53-year-old male patient needing an impella cp for mechanical circulatory support. it was reported that the impella cp pulled across the valve during repositioning. the pump was explanted. no additional information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined. corrections to the initial mfr report: f6/f8 should have been left blank. g1 mdr reporting contact email.
26727 | 1220648-2024-11175 | 19316061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19316061&pc=OZD | 02/24/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 02/24/2022 | 05/14/2024 |  |  | 09/30/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144859 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2022 |  |  |  |  | The user facility reported a 59-year-old male on impella cp for pre-percutaneous coronary intervention (pci). the patient was stable and recommendations were reviewed for correcting hemolysis. the patient had pink/red urine. the physician was concerned about hemolysis and had debated about explanting the impella cp. the patient was successfully escalated to an impella 5.0 via axillary insertion. femoral artery was successfully closed by surgical cutdown after removal of impella cp. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26728 | 1220648-2024-11176 | 19316072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19316072&pc=OZD | 03/15/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 03/15/2022 | 05/14/2024 |  |  | 11/30/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161993 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2022 |  |  |  |  | The user facility reported a 71-year-old female on impella cp for mechanical circulatory support. it was reported the patient was on impella cp support and a dissection was noted in the patient¿s aorta. the device was weaned and explanted. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26729 | 1220648-2024-11177 | 19316318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19316318&pc=OZD | 08/31/2022 | 05/14/2024 |  | 05/14/2024 |  |  | 08/31/2022 | 05/14/2024 |  |  | 11/30/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023164375 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Suction Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/31/2022 |  |  |  |  | The user facility reported a 65-year-old female was implanted with an impella cp for acute myocardial infarction /cardiogenic shock. it was reported that while on support, the pump experienced suction alarms with no specified resolution. no other information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26730 | 1220648-2024-11178 | 19316372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19316372&pc=OZD | 09/01/2022 | 05/15/2025 |  | 05/14/2024 |  |  | 09/01/2022 | 05/14/2024 |  |  | 09/30/2023 |  | 05/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022145562 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced bleeding at the access site. the physician successfully performed an impella wean and explanted the device. hemostasis was achieved with manual pressure and femstop. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26731 | 1220648-2024-11038 | 19296466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19296466&pc=OZD | 04/14/2024 | 06/11/2024 |  | 05/13/2024 |  |  | 04/14/2024 | 05/13/2024 |  |  | 11/30/2025 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424917 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2024 |  |  |  |  | The user facility reported a 79-year-old female with acute myocardial infarction and cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the patient developed a hematoma and bleeding, and the physician manually expressed the hematoma. manual pressure was held and a femoral stop was applied. the patient was administered one unit of blood products. the patient remained stable, but eventually expired under support. there is not enough information to associate the use of the impella device with the patient's outcome. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26732 | 1220648-2024-11039 | 19296500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19296500&pc=PYX | 04/14/2024 | 06/06/2024 |  | 05/13/2024 |  |  | 04/14/2024 | 05/13/2024 |  |  | 10/31/2025 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024423939 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Malfunction | Yes | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2024 |  |  |  |  | The user facility reported a 64-year-old male with a high risk of percutaneous coronary intervention and with chest pain was unsuccessfully implanted with an impella rp for mechanical circulatory support. the medical staff tried different techniques to implant the impella device. however, the implant was ultimately aborted. the patient eventually expired. there is not enough evidence to rule out the impella device as a factor in the patient's death. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella rp smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella rp smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
26733 | 1220648-2024-11040 | 19296594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19296594&pc=OZD | 02/12/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 02/12/2022 | 05/13/2024 |  |  | 09/30/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144897 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, bleeding at the site was noted. a forward suture was used and gauze placed to match access angle. fem-stop was placed to maintain hemostasis. patient remained on support for 2 days and survived the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26734 | 1220648-2024-11041 | 19296626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19296626&pc=OZD | 02/26/2022 | 05/13/2024 |  | 05/13/2024 |  |  | 02/26/2022 | 05/13/2024 |  |  | 08/31/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022134156 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/07/2024 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced bleeding at the access site and a hematoma. the patient received 2 units of blood. the bleeding subsided and systematic heparin was initiated. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26735 | 1220648-2024-11042 | 19296655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19296655&pc=OZD | 03/02/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 03/02/2022 | 05/13/2024 |  |  | 11/30/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161639 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2022 |  |  |  |  | The user facility reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, staff noted the patient had oozing at the groin site that was resolved with adjusting the angle. also noted was a hematoma that developed at the access site. the surgeon decided to explant the impella cp. manual pressure was held by staff until hemostasis was achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26736 | 1220648-2024-11043 | 19296683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19296683&pc=OZD | 03/09/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 03/09/2022 | 05/13/2024 |  |  | 07/31/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022120011 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Unknown |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2022 |  |  |  |  | The complainant reported that plans were made for a patient with right heart dysfunction to change from the impella cp to the impella 5.5 due to the escalation of three (3) drips. it was reported that the patient experienced an episode of atrial fibrillation (afib) that required cardioversion. weaning successful, the device was explanted, and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26737 | 1220648-2024-11044 | 19296709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19296709&pc=OZD | 04/23/2024 | 07/08/2024 |  | 05/13/2024 |  |  | 04/23/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1545370 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during training, the automated impella controller (aic) had the same alarm occur with two different pumps. the alarm message stated that there was a problem with the console after trying to raise the p level from zero. there was no reported patient involvement. | [5 empty] | The investigation for the reported console issues has been completed. the console was returned for evaluation. the console was run at a variety of p levels but the failure modes from the field were not reproduced. data logs were reviewed finding that there were multiple motor current high (alarm #13) and pump stop alarms. the persistent pump stops on start up were most likely due to dirty test pumps since motor current high and speed deviation alarms would occur with dirty pumps. added- d9 device return date g1-corrected mfr site name and address. | The automated impella controller (aic) was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26738 | 1220648-2024-11046 | 19296770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19296770&pc=OZD | 11/01/2023 | 06/04/2024 |  | 05/13/2024 |  |  | 04/23/2024 | 05/13/2024 |  |  | 08/31/2025 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024395599 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured bleeding (left) groin impella site with a "possible" relationship to the impella device used: ¿11/1: impella placed from l groin today. slow bleeding from l groin impella site noted by primary team despite pressure on the wound. vascular surgery consulted for evaluation. hematoma to left groin with hemostasis after holding pressure for ~45 min. lidocaine/epi injected near site for superficial bleeding. baseline hemoglobin (10/31): 15.0 11/1: 12.0 11/8: 7.1.¿ the patient recovered with sequelae. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
26739 | 1220648-2024-11045 | 19296772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19296772&pc=OZD | 11/05/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 11/05/2022 | 05/13/2024 |  |  | 06/30/2024 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240896 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2022 |  |  |  |  | The complainant reported a patient, in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient developed a thrombus in the left ventricle. it was also reported while on impella cp support, there were suction alarms from the impella, and the patient's mean arterial pressure (map) dropped from 90 mmhg to 50 mmhg. the patient received 500ml of crystalloid fluids to resolve the suction alarm, and the patient's map returned to 90 mmhg. the patient remained on impella cp support, and the thrombus was monitored by the medical team. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26740 | 1220648-2024-11047 | 19296808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19296808&pc=OZD | 04/07/2024 | 06/03/2024 |  | 05/13/2024 |  |  | 04/23/2024 | 05/13/2024 |  |  | 11/30/2025 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434117 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | PHYSICIAN ASSISTANT | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured recurring non-sustained ventricular tachycardia with a "possible" relationship to the impella device used: ¿the patient experiences episodes of recurring ventricular tachycardia including accelerated idioventricular rhythm. the patient has a history of ventricular tachycardia storms and has an icd in place that has historically appropriately shocked the vt rhythm. furthermore, the pt. has a history of vt ablation ((b)(6) 2024) with modified pathway for atypical av nodal reentrant tachycardia. electrolytes were replaced for index vt occurrences.¿ the patient outcome "not recovered/not resolved". | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. the instructions for use referenced in the initial report is from ifu for the impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states) d4-updated model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿
26741 | 1220648-2024-11048 | 19296858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19296858&pc=OZD | 12/04/2023 | 06/04/2024 |  | 05/13/2024 |  |  | 04/16/2024 | 05/13/2024 |  |  | 10/31/2025 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024406222 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Ventricular Fibrillation | Cardiogenic Shock | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/16/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured worsening of acute renal failure with a "possible" relationship to the impella device used: ¿acute kidney injury: creatinine of 3.93-3.97 - 3.19 - 2.96-2.59- 2.32-2.1-1.55 - 1.40 - 1.25- 1.19 - 1.2 - 1.5 - 1.4 - 1.36 - 1.29 - 1.23-1.17-1.42 - 1.64. likely cardiorenal the setting of cardiogenic shock. discussed with and consulted nephrology. monitored urine output and kidney function closely. avoided nephrotoxic agents. uo improving. continued po torsemide. losartan held. s/p renal bx on (b)(6) 2023 which showed fsgs - no need for further specific rx. creatinine stabilized around 1.7 on (b)(6) 2024 prior to discharge.¿ additionally, we received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured recurrent non-sustained ventricular tachycardia with a possible relationship to the impella device used: ¿patient had non-sustained vtach after cough. impella repositioned twice on (b)(6) 2023. amiodarone given. multiple episodes: (b)(6) 2023.¿ the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿ potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported renal failure and ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the renal failure and vt could not be determined. the instructions for use referenced in the initial report is for the impella 5.5 with smart assist for use during cardiogenic shock. d4-updated model number.
26742 | 1220648-2024-11049 | 19296898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19296898&pc=OZD | 03/18/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 03/18/2022 | 05/13/2024 |  |  | 06/30/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022102843 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced multiple episodes of ventricular tachycardia (vt). subsequently it was reported that the patient improved and was weaned off support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | Cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
26743 | 1220648-2024-11050 | 19296920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19296920&pc=OZD | 02/26/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 03/15/2022 | 05/13/2024 |  |  | 04/30/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP - JP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022075997 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  | Asian | Injury | Yes | No | Restricted Flow rate | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient¿s condition was unstable due to ventricular tachycardia (vt) when leaving the cath room. both the impella and extracorporeal membrane oxygenation (ecmo) were in a situation where it was difficult to flow due to the lack of volume. additionally, it was reported that the patient developed bleeding and a hematoma. the patient required blood transfusions. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26744 | 1220648-2024-11051 | 19296935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19296935&pc=OZD | 08/24/2023 | 05/29/2024 |  | 05/13/2024 |  |  | 04/17/2024 | 05/13/2024 |  |  | 07/31/2025 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024374139 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured ischemic stroke with a "possible" relationship to the impella device used: ¿stroke code called for left-sided weakness. cervicocerebral angiogram did not show large vessel occlusion but, per neuro but cannot rule out stroke that had happen. cth did not show acute stroke on (b)(6). however, on (b)(6) cth shows numerous lesions which are possibly acute infarcts within the bilateral corona radiata, bilateral basal ganglia, and right thalamus which could have happened in the setting of shock. mri brain stroke protocol completed on 9/1, showing bilateral acute strokes in the watershed and supratentorial cortex and bilateral cerebellum, consistent with previously thought cardioembolic and hypoperfusion associated stroke," per neurology. (b)(6) seems to be improving with waxing and waning attention span. 9/12 dc to a rehab facility¿. the patient recovered with sequelae. | [5 empty] | The investigation into the stroke/cva issue has been completed since the original report was submitted. the pump was not returned for investigation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26745 | 1220648-2024-11052 | 19296956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19296956&pc=OZD | 04/08/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 04/08/2022 | 05/13/2024 |  |  | 01/31/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022034510 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there was heavy groin access site bleeding after peeling away the sheath. as a result, the pump was explanted and a dry seal was placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26746 | 1220648-2024-11053 | 19297015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19297015&pc=OZD | 04/23/2022 | 05/13/2024 |  | 05/13/2024 |  |  | 04/23/2022 | 05/13/2024 |  |  | 01/31/2024 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186272 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/23/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support for percutaneous coronary intervention. while on support, the patient experienced bleeding from all sites, including the nose. the physician held off the systemic heparin due to the bleeding. subsequently, it was reported that the patient became hemodynamically stable. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions. as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: use of impella with transcatheter aortic valves. as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
26747 | 1220648-2024-11054 | 19297039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19297039&pc=PYX | 04/14/2024 | 06/04/2024 |  | 05/13/2024 |  |  | 04/14/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025450298 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2024 |  |  |  |  | The user facility reported that a patient implanted with an impella rp and developed access site bleeding. the bleeding was resolved with a suture, pressure dressings, and clotting agents at the insertion site. later, the patient developed elevated white blood cells and urosepsis was suspected. therefore, the patient was treated with antibiotics. the patient expired while on support. there is not enough information to associate the use of the device with the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ impella rp flex with smart assist system with automated impella controller section: potential adverse events (united states) .¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer: ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into the access site bleeding ¿ major and the infection/sepsis issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of access site bleeding is determined to be use issue because the bleeding was significantly decreased upon second suture placed after sheath exchange. because evidence related to the issue was not provided, the root cause of the infection/sepsis could not be determined.
26748 | 1220648-2024-11055 | 19297043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19297043&pc=OZD | 04/23/2024 | 06/24/2024 |  | 05/13/2024 |  |  | 04/25/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025456865 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the male patient suffered a cerebral infarction post implantation. the stroke was embolic. the patient had hemiparesis on the left. the cerebral infarction may have been due to calcifications in the aorta. the paresis was noted to have regressed and the patient is stable. the device was not removed as a result of the issue. | [5 empty] | The investigation into the stroke/cva issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided, the root cause of the stroke/cva was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26749 | 1220648-2024-11056 | 19297097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19297097&pc=OZD | 04/24/2022 | 05/13/2024 |  | 05/13/2024 |  |  | 04/24/2022 | 05/13/2024 |  |  | 07/31/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022116616 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/24/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a coughing episode resulting in groin site oozing. the team held pressure at the arterial insertion site and rechecked the angle of insertion. subsequently, the oozing and hematoma were reportedly resolved. the impella was removed at bedside, the physician left the 14 french peel away in. it was reported that the patient was to have surgical repair. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
26750 | 1220648-2024-11057 | 19297154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19297154&pc=OZD | 04/28/2022 | 05/13/2024 |  | 05/13/2024 |  |  | 04/28/2022 | 05/13/2024 |  |  | 11/30/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165822 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a small hematoma was noted, but it reportedly resolved with no palpable impact per the physician. additionally, oozing was noted at the impella site. the patient required surgery due to complications of bleeding. heparin was also decreased and reportedly the bleeding became controlled. post-surgery, the graft was bleeding from the anastomosis. to resolve the bleeding systemic heparin was discontinued and the patient was reported to be stable again. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations . ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer . as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26751 | 1220648-2024-11058 | 19297169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19297169&pc=OZD | 04/24/2024 | 07/03/2024 |  | 05/13/2024 |  |  | 04/25/2024 | 05/13/2024 |  |  | 01/31/2025 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024313089 |  | 0048-0014 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the 68-year-old male patient had developed an ischemic leg. the leg was cold and showed livedo reticularis. a trial wean was performed and was positive. the impella was therefore explanted. the patient was noted to be stable following impella removal, and a small dose of norepinephrine was kept running. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the potential adverse events (united states) section.
26752 | 1220648-2024-11059 | 19297194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19297194&pc=OZD | 09/15/2022 | 05/13/2025 |  | 05/13/2024 |  |  | 09/15/2022 | 05/13/2024 |  |  | 04/30/2024 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221081 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/15/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during support there were intermittent suction events noted and the patient received 4l of infused volume with a combination of blood products, albumin, crystalloid fluids. the patient was on venoarterial extracorporeal membrane oxygenation (va ecmo) with the impella cp to unload the left ventricle. however, at p2 there were still suction events. it was recommended to the care team to use the echocardiogram for position verification. it was noted that the impella cp was later explanted, however, it was noted to have been explanted independently. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26753 | 1220648-2024-11060 | 19297371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19297371&pc=OZD | 02/26/2022 | 05/13/2024 |  | 05/13/2024 |  |  | 02/26/2022 | 05/13/2024 |  |  | 11/30/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166678 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Positioning Failure | Restricted Flow rate | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/26/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the device kept getting stuck in mitral apparatus and suction alarms were noted despite three repositioning attempts. additionally, the patient experienced hemolysis. the physician elected to explant the device and used manual pressure to help on the closure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26754 | 1220648-2024-11061 | 19297396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19297396&pc=OZD | 01/28/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 01/28/2022 | 05/13/2024 |  |  | 10/31/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157131 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, mixed venous saturation, and cardiac index increased. the groin access site had continuous oozing and a venous triple lumen was placed next to the impella insertion site. access site bleed continued despite preclose tightening. it is stated that perforation of artery was the potential root cause of the bleeding. a vascath central venous catheter was placed on the left venous femstop. the patient began continuous renal replacement therapy. as hemodynamics improved, the physician made the decision to remove the impella. weaning was successful, and the device was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26755 | 1220648-2024-11062 | 19297416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19297416&pc=OZD | 06/12/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 06/12/2022 | 05/13/2024 |  |  | 06/30/2022 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1488546 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/12/2022 |  |  |  |  | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. while on support, it was noted that patient experienced heparin induced thrombocytopenia (hit) and the device had high purge pressure alarm. hit was resolved by replacing heparin with sodium bicarbonate and high purse pressure alarm was resolved by adding heparin to the purge solution. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26756 | 1220648-2024-11063 | 19297615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19297615&pc=OZD | 01/28/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 01/28/2022 | 05/13/2024 |  |  | 10/31/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149930 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported that the impella cp was removed from lfa due to limb ischemia and the patient was escalated to a 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
26757 | 1220648-2024-11066 | 19297830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19297830&pc=OZD | 04/26/2024 | 06/03/2024 |  | 05/13/2024 |  |  | 04/26/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1378196 |  | 0042-0000-UK |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following optical update, the automated impella controller (aic) displayed a yellow 'controller error (dfective optical sensor lamp) alarm. a new optical bench was used and the console started normally but after 5 minutes the error appeared again. the 5 v supply from the the pbm was checked and was noted to be ok. there was no reported patient involvement as the failure occurred curing maintenance. | [5 empty] | The investigation for the reported console (aic) yellow alarm issues has been completed. the console was returned for evaluation. upon visual inspection, there was a broken optical fiber connection from the lamp assembly (defect of the optical bench). data logs were reviewed which were consistent with complaint and show controller error alarm (#1039, due to defective optical sensor lamp), preceded by ¿5s¿ diagnostic data indicating loss of light source (but no loss of communication to optical bench). the cause of the aic issue was a defective optical bench. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
26758 | 1220648-2024-11068 | 19297857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19297857&pc=OZD | 04/27/2024 | 12/19/2024 |  | 05/13/2024 |  |  | 04/27/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1326577 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | Defective Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for 62-year-old male patient, the "condition wheel" on the automated impella controller (aic) is defective. there was no reported patient harm. the patient is noted to be stable and on ecmo support. | [5 empty] | The automated impella controller (aic) was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the kbd/rotary knob issues has been completed. data logs were reviewed which showed while running pump, the user opened up the menu to choose p-level and set the pump to p-2. the p-level menu is opened up again two more times and each time 2 is entered despite the fact that was already the p-level. this was most likely a case of the user trying to change the p-level but being unable to select the number they wanted indicating a the rotary knob not operating as expected. the console was returned evaluation. upon functional testing, the failure mode was reproduced at fs with the rotary knob skipping over selections. after replacing the impellatronic board, the rotary knob performed as expected. the cause of the malfunctioning rotary knob was a defective impellatronic board.
26759 | 1220648-2024-11065 | 19297875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19297875&pc=OZD | 04/26/2024 | 07/12/2024 |  | 05/13/2024 |  |  | 04/26/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1181836 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) did not recognize the purge disc. therefore, the console was exchanged. there was no reported patient harm. | [5 empty] | The investigation for the reported purge disc/flag issues has been completed. the console was returned for evaluation. this was not reproduced as the console was able to recognize the purge disc but there was glucose residue around the purge pressure sensor that could have affected purge disc recognition. the logs were reviewed which confirmed the inability to detect the purge disc. case wizard was then started up but the purge disc would still not be recognized. purge cassette swapped for sn (b)(6) but no difference in the results. the cause of the purge disc's intermittent inability to be recognized was most likely contamination from dried glucose. added-d9 device return date g1-corrected mfr site name and address. | The automated impella controller (aic) was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26760 | 1220648-2024-11070 | 19297985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19297985&pc=OZD | 01/29/2022 | 05/21/2025 |  | 05/13/2024 |  |  | 01/29/2022 | 05/13/2024 |  |  | 08/31/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022133655 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there was an intermittent aic "impella in aorta" alarm. an echocardiogram was performed, the impella pig tail part of the catheter was still across the valve. the absence of pulses to the distal lle on doppler was noted. the decision was made to discontinue impella support and the impella cp was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26761 | 1220648-2024-11072 | 19298028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298028&pc=OZD | 05/03/2024 | 07/10/2024 |  | 05/13/2024 |  |  | 05/03/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025458192 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Difficult to Remove | Failure to Deliver | Patient-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2024 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant was attempting to deliver a cp to support a patient with savr history and low ef, as well as as (aortic stenosis). the aortic valve was stenotic of such a degree that the pump could not deliver across to the lv (left ventricle). the pump delivery was aborted and the md made choice to explant the wire also, but it would not exit the sheath. the team had to obtain contralateral leg access and snare the product out of the patient. | [5 empty] | The investigation for the reported delivery and removal issues has been completed. the pump and guidewire were not returned. the cause of the delivery issue was patient condition related due to pump unable to cross bioprosthetic valve. the cause of the product damage (guidewire) issue- peripheral vessels was not determined since product was not returned and due to insufficient clinical information. b1-should be product problem only instead of adverse event. d4-corrected model number. f6/f8 should have been left blank. g1 reporting contract e-mail corrected. h1-type of reportable event corrected to product problem. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26762 | 1220648-2024-11073 | 19298104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298104&pc=OZD | 01/30/2022 | 05/13/2024 |  | 05/13/2024 |  |  | 01/30/2022 | 05/13/2024 |  |  | 11/30/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161313 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 01/30/2022 |  |  |  |  | The user facility reported a patient needing circulatory mechanical support. it was reported that while the patient was on impella cp support, the patient had ventricular fibrillation (vf) episodes. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26763 | 1220648-2024-11074 | 19298124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298124&pc=OZD | 01/30/2022 | 05/13/2024 |  | 05/13/2024 |  |  | 01/30/2022 | 05/13/2024 |  |  | 09/30/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022139482 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/30/2022 |  |  |  |  | The user facility reported a 78-year-old patient needing mechanical circulatory support. it was reported that the patient was placed onto impella cp support for acute myocardial infarction / cardiogenic shock. while on support, the patient experienced access site bleeding that was resolved by lowering heparin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26764 | 1220648-2024-11064 | 19298147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298147&pc=OZD | 01/28/2022 | 04/16/2025 |  | 05/13/2024 |  |  | 01/28/2022 | 05/13/2024 |  |  | 05/31/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022087250 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/28/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be implanted with an impella cp device for mechanical circulatory support. impella placement was unsuccessful as the device was unable to pass through the patient¿s vasculature and the procedure was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26765 | 1220648-2024-11067 | 19298187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298187&pc=OZD | 01/28/2022 | 05/13/2024 |  | 05/13/2024 |  |  | 01/28/2022 | 05/13/2024 |  |  | 09/30/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144850 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/28/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support slight oozing at the impella access site was noted, this was managed with pressure dressing and the patient was given two (2) units of blood. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26766 | 1220648-2024-11069 | 19298199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298199&pc=OZD | 04/27/2024 | 06/11/2024 |  | 05/13/2024 |  |  | 04/27/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1154857 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the 50-year-old male patient was transferred to different hospital. during ecmo insertion, the automated impella controller (aic) displayed a 'controller failure' yellow alarm. the console was therefore exchanged and the issue resolved. there was no reported patient harm. | [5 empty] | The investigation for the reported console (aic) yellow alarm has been completed. the console was returned for evaluation. reported controller error alarm was identified as related to communication issues between kbd and 4-in-1 carrier. the issue was not reproduced during testing but was most likely related to an undetermined malfunction of kbd. it is most likely that the momentary communication glitch originated at the kbd assembly. console logs from the reported day of event are mostly consistent with complaint and show a (yellow) controller error alarm (#24, due to loss of communication with keypad), preceded by several error messages indicating kbd communication errors ¿imc-kbd error: 0xa2 0x52 0x00¿. the cause of the aic issue was a defective kbd/display. added- d6a/d6b. | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26767 | 1220648-2024-11075 | 19298272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298272&pc=OZD | 05/04/2024 | 09/26/2024 |  | 05/13/2024 |  |  | 05/04/2024 | 05/13/2024 |  |  | 07/31/2025 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440658 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Death | Yes | Yes | Positioning Failure | Migration or Expulsion of Device | Adverse Event Without Identified Device or Use Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/04/2024 |  |  |  |  | The complainant reported a 71-year old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. after the implant, there were positioning issues. the impella slid across the bio-prosthetic aortic valve multiple times, however, medical staff was able to reposition it. when the impella slid again, the pigtail was also in the aorta - there was no longer the ability for impella to be advanced. the pump was explanted. after the explant, other hospitals would no longer accept the patient due to the patient being unstable during transport. escalation to other impella devices was also not possible due to the positioning issues previously experienced. the patient was put on comfort measures only and expired after approximately four hours of support. it is unknown if the impella was correlated to the cause of death. | [5 empty] | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-11075 was provided in sections b2, g3, and h1, and h6. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the positioning issues been completed since the original report was submitted. the device was not returned for investigation. based on clinical description, the root cause of positioning issue was most likely use related.
26768 | 1220648-2024-11071 | 19298286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298286&pc=OZD | 05/03/2024 | 07/08/2024 |  | 05/13/2024 |  |  | 05/03/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025444074 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 24 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2024 |  |  |  |  | The user facility reported a 24-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan reported that impella 5.5 was not able to be placed due to the patient¿s vascular anatomy. there was no patient harm. | [5 empty] | The investigation for the delivery issue has been completed. the impella was not returned for evaluation. however, clinical details were sufficient to determine the root cause of the issue. the root cause of the delivery issue was most likely patient condition related, impella 5.5 was not able to be placed due to the patient¿s vascular anatomy. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ d4-corrected model number, corrected udi number f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26769 | 1220648-2024-11077 | 19298367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298367&pc=OZD | 05/05/2024 | 06/24/2024 |  | 05/13/2024 |  |  | 05/05/2024 | 05/13/2024 |  |  | 07/31/2025 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387769 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/05/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pigtail of an implanted impella cp pump became wrapped around the mitral valve chordae during patient support. attempts to reposition the pump were unsuccessful and the decision was ultimately made to escalate to impella 5.5 support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was submitted. the product was not returned for investigation. data logs showed no positioning alarm when pump against mitral apparatus. based on the clinical description and the data analysis, the root cause of positioning issue was most likely use related.
26770 | 1220648-2024-11078 | 19298390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298390&pc=OZD | 01/31/2022 | 05/13/2024 |  | 05/13/2024 |  |  | 01/31/2022 | 05/13/2024 |  |  | 08/31/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022124758 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/31/2022 |  |  |  |  | The user facility reported a 51-year-old patient needing mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient had impella limb ischemia. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26771 | 1220648-2024-11079 | 19298414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298414&pc=OZD | 05/02/2024 | 07/03/2024 |  | 05/13/2024 |  |  | 05/02/2024 | 05/13/2024 |  |  | 11/30/2025 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024436203 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The user facility reported a 76-year-old male noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that after the pump was removed, the physician tried to close the artery with a perclose and it did not take. the perclose tore the common femoral artery. a gore dry seal sheath was placed to maintain hemostasis until vascular repair in the operating room performed. the patient was reported as stable. | [5 empty] | The investigation for the reported vascular damage has been completed. according to the clinical, shortly after beginning impella cp support the decision was made to remove the pump. after removing it without issue the physician attempted to close the artery with perclose, but accidentally tore the common femoral artery. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the peripheral vessel vascular damage was use related. g1a: reporting contact fax number f6/f8-should have been left blank in the initial report. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
26772 | 1220648-2024-11076 | 19298453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298453&pc=OZD | 05/05/2024 | 06/13/2024 |  | 05/13/2024 |  |  | 05/05/2024 | 05/13/2024 |  |  | 09/30/2025 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445687 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | VENTILATOR | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/05/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient lost pulse in the right foot. in response to the noted condition, the decision was made to explant the impella cp pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
26773 | 1220648-2024-11080 | 19298481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298481&pc=OZD | 02/04/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 02/04/2022 | 05/13/2024 |  |  | 10/31/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158893 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced suction alarms that were resolved with p level reduction and bolus. additionally, the patient experienced a thrombus that was resolved with thrombectomy, percutaneous transluminal angioplasty, and lazer. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26774 | 1220648-2024-11081 | 19298551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298551&pc=OZD | 05/05/2024 | 06/20/2024 |  | 05/13/2024 |  |  | 05/05/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025482751 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Ventricular Fibrillation | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/05/2024 |  |  |  |  | Additional information was provided which states that impella was initially placed deep, then pulled back a total 5cm which removed a lot of slack. then pulled the pump back to the elbow on the valve. waveforms continued to be less than optimal and the doctor attempted to torque pump away from mitral valve/septum but was unable to translate across pump. the patient was transferred back to cath lab for pump reposition/replacement. later that evening, the pigtail deviated inferolaterally and was stuck in pap. attempted manipulation bedside by doctor unsuccessful and the patient was again brought back to cath lab to have another doctor reposition pump and was successful. | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and during the impella implant the patient had ventricular tachycardia (vt). the impella was placed in right femoral groin, left intentionally shallow as when it was advanced it caused vt. support continued. additionally, the patient had a hematoma at the access site. manual pressure was held and the hematoma resolved. |  |  |  |  | The investigation for the reported ventricular tachycardia and positioning issues has been completed. without additional information and the device, the root cause of the vt could not be determined. however, clinical details were sufficient to determine the root cause of the positioning issue. the root cause of positioning issue is determined to be use issue because the pump was too deep. the patient was taken to or and the pump was pulled back. the instructions for use for the impella cp with smartassist system in table 6.31 in hospital adverse events lists ventricular arrythmia has an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." added-b5 additional narrative, b1 selected product problem, h6 med dev prob code 1158 d4-corrected model number. g1-corrected mfr site name and address. h6 clinical code removed 1884. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿
26775 | 1220648-2024-11082 | 19298554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298554&pc=OZD | 09/01/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 09/01/2022 | 05/13/2024 |  |  | 06/30/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022107027 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported an 80-year-old male on impella cp for mechanical circulatory support. during a trans-thoracic echo, a possible clot formation was located near the inlet window. since the patient was being weaned off support, the decision was made to remove the impella cp. impella cp device was successfully explanted and groin site perclosed, sutured, cinched, and trimmed. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
26776 | 1220648-2024-11083 | 19298595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298595&pc=OZD | 05/03/2024 | 07/02/2024 |  | 05/13/2024 |  |  | 05/03/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025453616 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The user facility reported a 64-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. two days after implant, the patient showed left sided weakness and the neurology team expressed concerns of a sub clinical stroke. the patient's blood pressure was raised in response to the stroke and impella cp support continued for two more days. | [5 empty] | The investigation into stroke/cva issue has been completed since the original report was submitted. the device was not returned for investigation. insufficient clinical information was provided. therefore, the true root cause of the stroke/cerebrovascular accident could not be determined. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-11083 was submitted. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
26777 | 1220648-2024-11084 | 19298601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298601&pc=OZD | 09/02/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 09/02/2022 | 05/13/2024 |  |  | 05/31/2024 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231103 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/02/2022 |  |  |  |  | The use facility reported a 58-year-old male on impella cp support for a percutaneous coronary intervention (pci). the patient developed oozing at the impella cp access site. anticoagulation protocol and partial thromboplastin time (ptt) goals were discussed. care team sent for ptt and will assess the need for systemic heparin. systemic heparin was added and was later turned off. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26778 | 1220648-2024-11085 | 19298769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298769&pc=OZD | 05/02/2024 | 07/09/2024 |  | 05/13/2024 |  |  | 05/02/2024 | 05/13/2024 |  |  | 09/30/2025 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024406420 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The user facility reported a 77-year-old female noted as high risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella implant was aborted due to calcium in the aorta preventing the cannula from passing. no patient harm was reported. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for investigation. the root cause of the delivery issue was determined to be patient condition because the pump was unable to pass through the vasculature due to possible calcification of the vessels. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-11085 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since manufacturer device report 1220648-2024-11085 was submitted. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
26779 | 1220648-2024-11086 | 19298932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298932&pc=OZD | 05/06/2024 | 06/11/2024 |  | 05/13/2024 |  |  | 05/06/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1110485 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/06/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) was unable to read the purge disc. multiple troubleshooting steps were taken. the aic was exchanged. there was no patient involvement. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge flag had broken off on the console. the root cause of this issue was a broken purge flag. d4-updated model number. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
26780 | 1220648-2024-11087 | 19298975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298975&pc=OZD | 05/06/2024 | 06/25/2024 |  | 05/13/2024 |  |  | 05/06/2024 | 05/13/2024 |  |  | 11/30/2025 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024425794 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. white the patient was on support, there was an expanding hematoma at the access site. staff was unable to resolve with pressure. the impella cp was removed. the patient was stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26781 | 1220648-2024-11088 | 19298999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19298999&pc=OZD | 05/06/2024 | 07/08/2024 |  | 05/13/2024 |  |  | 05/06/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025452705 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient exhibited hemolysis. the impella cp was explanted. | [5 empty] | The investigation for the hemolysis was completed. the impella was not returned for evaluation. however, data logs were reviewed which showed pump ran without issue during support. there were some intermittent suction alarms during the case but resolved quickly. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the instructions for use referenced in the initial report is for the impella cp with smartassist system in the following sections: section: potential adverse events. (united states) section: suction. section: hemolysis. section: use of echocardiography for positioning of the impella catheter f6/f8-should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26782 | 1220648-2024-11089 | 19299010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19299010&pc=OZD | 05/07/2024 | 07/01/2024 |  | 05/13/2024 |  |  | 05/07/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485946 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient's leg became cold during the impella cp support. in response to the condition, the decision was made to remove the pump and place a new impella cp device via transcaval access. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. d.4 revised model number from the initial submission in accordance with updated procedures. h.6 code 4114 was reported incorrectly on the previously submitted report. a new code has been added to type of investigation codes. h.10 additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system-section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
26783 | 1220648-2024-11090 | 19299033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19299033&pc=OZD | 05/04/2024 | 06/24/2024 |  | 05/13/2024 |  |  | 05/04/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025463359 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy implanted with an impella cp device for mechanical circulatory support experienced positioning issue and access site bleeding. from echocardiogram, the impella pigtail was noted to be slightly in the papillary muscle. the physician attempted to reposition the catheter; however, it was unsuccessful. the decision was made to leave the position as is, no complications thus far, and support continued. subsequently noted, there was bleeding to impella access site after the patient became agitated. sedation was adjusted, and the stie was redressed with no further issues noted. | [5 empty] | The investigation into the positioning issues and the access site bleeding - major issues has been completed since the original report was submitted. the product was not returned for investigation. based on the clinical details provided, the cause of the access site bleeding issue was most likely patient movement, and the cause of the hemolysis was improper positioning since the complication was successfully managed by repositioning the catheter. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26784 | 1220648-2024-11091 | 19299095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19299095&pc=OZD | 05/04/2024 | 07/03/2024 |  | 05/13/2024 |  |  | 05/07/2024 | 05/13/2024 |  |  | 08/31/2025 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387386 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Valvular Insufficiency/ Regurgitation | Aortic Valve Insufficiency/ Regurgitation | Intraoperative Cardiac Valve Injury | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/07/2024 |  |  |  |  | The complainant reported a patient of unknown age and gender was implanted with an impella cp for mechanical circulatory support. the complainant reported that the patient suffered a prosthetic valve injury during a planned explant of the cp pump. medical staff noted aortic regurgitation following the pump removal. medical staff also implanted a transcatheter aortic valve. the patient remained stable following the intervention. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system: section: warnings & cautions: warnings. section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion. ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation for the reported heart valve damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the heart valve damage could not be determined. added-e4: no, h6 component code 925 f6/f8-should have been left blank. g1a reporting contact fax number missing.
26785 | 1220648-2024-11093 | 19299302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19299302&pc=OZD | 05/07/2024 | 07/16/2024 |  | 05/13/2024 |  |  | 05/07/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025464465 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support and experienced a positioning issue. approximately 23 minutes into support, the impella pump came across the aortic valve and was unable to re-cross the valve. the decision was made to remove and replace with another impella cp device. there was no report of patient harm to the patient. | [5 empty] | The investigation for the reported positioning issue has been completed. pump was not returned for investigation. as per the clinical, impella was pulled out of aorta and unable to re-cross the valve. the cause of the positioning issue was most likely use related. d4-corrected model number. h6 impact code changed from 4627 to 4629. | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26786 | 1220648-2024-11094 | 19299365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19299365&pc=OZD | 05/07/2024 | 03/24/2025 |  | 05/13/2024 |  |  | 05/07/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1536903 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/07/2024 |  |  |  |  | The complainant reported an 80-year old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the automated impella controller (aic) had a battery alarm occur, followed by a shutdown and a controller failure alarm. the aic was replaced and support continued. no patient harm has been reported, and the patient remains on support. | [5 empty] | The investigation of automated impella controller (aic) - shutdown or restart issue and aic - battery comm. failure (yellow alarm) has been completed. according to device and data analysis the reported battery low alarm and unexpected shutdown was determined to have most likely been caused by an intermittent malfunction of the power battery manager (pbm). d9 date device returned to manufacturer added. d2 common device name revised on manufacturer device report 1220648-2024-11094 in accordance with updated procedures. d6a should have been left blank on manufacturer device report 1220648-2024-11094 f6/f8-should have been left blank on manufacturer device report 1220648-2024-11094 g1 reporting contact email revised on manufacturer device report 1220648-2024-11094 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-11094. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
26787 | 1220648-2024-11095 | 19299570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19299570&pc=OZD | 05/05/2024 | 06/11/2024 |  | 05/13/2024 |  |  | 05/08/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1696098 |  | 0042-0040-US |  | Yes |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The complainant reported a 50-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that while the automated impella controller (aic) was unplugged, a battery temperature high alarm sounded. no patient harm has been reported. | [5 empty] | The investigation for the reported console (aic) "battery temp alarm high" has been completed. the console was returned for evaluation. upon functional testing, the console¿s batteries and power management system functioned normally without issues. console logs from the complaint date revealed the ¿battery temperature high¿ alarm (#49, due to battery temperature over 60degc) that self-resolved in 10 seconds. there was no preceding battery temperature alarm #48, which occurs when battery temperature is between 50degc and 60degc. battery and power data embedded in left data logs do not show any temperature rise or elevated battery temperatures, as left data log data is recorded every 60s. the battery level low alarm was due to a momentary communication glitch between the batteries and pbm, where a single sample from battery data (in this case value of battery temperature) was corrupt. d2-corrected common device name. d6a-should be blank. this console is not implantable. d8-device available for evaluation changed to yes | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
26788 | 1220648-2024-11096 | 19299676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19299676&pc=OZD | 05/08/2024 | 07/17/2024 |  | 05/13/2024 |  |  | 05/08/2024 | 05/13/2024 |  |  | 09/30/2025 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395292 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The complainant reported a 59-year-old male patient with a high risk of percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. the complainant reported that after approximately 502 hours of support, the patient expired from ventricular tachycardia (vt). the cause of the patient¿s vt remains unknown. | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. the instructions for use for the impella cp with smartassist system lists ventricular arrythmia has an adverse event that occurred while in use with the impella device. it states that "only a fraction of these rates were attributed to the impella device and the events resolved without residual effect in most of the cases, unless the event of death occurred. overall, the results did not show any evidence of increased morbidity associated with the impella support in cardiomyopathy, myocarditis, and peripartum cardiomyopathy (ppcm) patients." it also states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ f6/f8 should have been left blank in the initial report. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
26789 | 1220648-2024-11097 | 19300007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19300007&pc=OZD | 05/08/2024 | 06/24/2024 |  | 05/13/2024 |  |  | 05/08/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025487500 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The complainant reported a 49-year-old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support and to allow for percutaneous coronary intervention. the cp pump was explanted after approximately a day of support due to signs of hemolysis. the explant was performed successfully, and an intra-aortic balloon pump was placed. the patient survived the event. | [5 empty] | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
26790 | 1220648-2024-11098 | 19300075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19300075&pc=OZD | 05/08/2024 | 06/24/2024 |  | 05/13/2024 |  |  | 05/08/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025468700 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The user facility reported an 85-year-old male undergoing a cryoablation was implanted with an impella cp device for mechanical circulatory support. it was reported that after cryoablation started the cp flows dropped abruptly. troubleshooting was unsuccessful. the physician removed the impella, leaving 14/13 sheath and proceeded with cryoablation without replacing the cp but with a backup available. the patient is stable. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned for investigation. the returned pump was visually inspected, and a cannula kink was observed. minimal biomaterial was observed in the cannula but not on the impeller. the cause of the issue was a kinked cannula. pump in ventricle alarms at the time of the suction and low pump flow which may have contributed to the cannula kink. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26791 | 1220648-2024-11099 | 19300358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19300358&pc=OZD | 05/08/2024 | 06/25/2024 |  | 05/13/2024 |  |  | 05/08/2024 | 05/13/2024 |  |  |  |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025460418 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2024 |  |  |  |  | The complainant reported a 49-year-old female patient with a high risk of percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. the complainant reported that the patient developed a hematoma and oozing at the access site. medical staff obtained hemostasis with manual pressure and a femoral stop. medical staff also withheld heparin. the patient survived the event and was successfully weaned off the cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) limb¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the root cause of the access site bleeding was most likely due to patient movement since the hematoma formed due to excessive leg movements. g1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h6 code 925 was reported incorrectly on the previously submitted report.
26792 | 1220648-2024-11100 | 19300412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19300412&pc=OZD | 05/11/2022 | 05/16/2025 |  | 05/13/2024 |  |  | 05/11/2022 | 05/13/2024 |  |  | 12/31/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023163500 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced intermittent temporary pump stoppage that required the impella to be explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: pump stop: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26793 | 1220648-2024-11101 | 19300587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19300587&pc=OZD | 02/12/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 02/12/2022 | 05/13/2024 |  |  | 09/30/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142332 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2022 |  |  |  |  | The user facility reported a female patient needing mechanical circulatory support. it was reported while being on the impella cp the patient had bleeding at the access site. no retroperitoneal (rp) was observed on the computed tomography (ct) scan. it was stated the patient¿s urine was red and the liver function tests (lfts) were elevated. repositioning was recommended but not initiated. weaning was successful, the device was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26794 | 1220648-2024-11102 | 19301432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19301432&pc=OZD | 04/02/2024 | 05/13/2024 |  | 05/13/2024 |  |  | 04/02/2024 | 05/13/2024 |  |  | 12/31/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172954 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/02/2024 |  |  |  |  | Th complainant reported a patient was placed onto impella cp support for cardiomyopathy. while on support, the patient experienced access site bleeding that was resolved with discontinuation of heparin. patient's condition was stable post procedure. | [5 empty] | The investigation for the reported access site bleeding has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. the root cause the access site bleeding was most likely due to anticoagulation management since the bleeding was resolved by stopping the heparin. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26795 | 1220648-2024-11103 | 19306623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19306623&pc=OZD | 05/13/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 05/13/2022 | 05/13/2024 |  |  | 07/05/2024 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022069638 |  | 0048-3075 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/13/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support due to mi (myocardial infarction), stroke, arrhythmia including vf (ventricular fibrillation), severe left main, and occluded distal lad (left anterior descending). while on support problems with axillary insertion to tortuosity of axillary artery was noted. issue resolved by using a longer introducer sheath 14fr x 25 cm instead of the 23fr introducer sheath from axillary kit. patient remained on support and survived the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26796 | 1220648-2024-11104 | 19306782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19306782&pc=OZD | 08/27/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 08/27/2022 | 05/13/2024 |  |  | 05/31/2024 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230725 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2022 |  |  |  |  | The user facility reported a 73-year-old male patient was implanted with an impella cp for mechanical circulatory support due to acute myocardial infarction/cardiogenic shock. while on support, the patient experienced access site bleeding that was resolved by withholding heparin. the patient was reported to be in stable condition. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26797 | 1220648-2024-11105 | 19306804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19306804&pc=OZD | 08/31/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 08/31/2022 | 05/13/2024 |  |  | 10/31/2023 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152045 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced suction and was repositioned with no issues. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26798 | 1220648-2024-11106 | 19306810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19306810&pc=OZD | 09/01/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 09/01/2022 | 05/13/2024 |  |  | 03/31/2024 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214131 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported a 74-year-old female on impella cp for mechanical circulatory support. it was reported the the impella cp device was weaned and removed from the patient. perclose failed x4. manual pressure was held for 30 min partnered with 12mm balloon inflation. hemostasis was successfully achieved. a fem stop was placed for light pressure. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26799 | 1220648-2024-11107 | 19306832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19306832&pc=OZD | 09/01/2022 | 05/15/2025 |  | 05/13/2024 |  |  | 09/01/2022 | 05/13/2024 |  |  | 01/31/2024 |  | 05/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187079 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported a 70-year-old male was implanted with an impella cp for mechanical circulatory support. after implantation, the dressing was changed on the right femoral impella site with slight sanguineous drainage was noted. angle matching was performed, and hub supported with gauze. the patient also received packed red blood cells due to hemoglobin dropping. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26800 | 1220648-2024-10994 | 19285324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19285324&pc=OZD | 02/16/2022 | 05/16/2025 |  | 05/10/2024 |  |  | 02/16/2022 | 05/10/2024 |  |  | 08/31/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022129324 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site oozing and a small hematoma at the insertion site. the patient received an unknown amount of packed red blood cells that were given for hemoglobin of 6.8. additionally, it was reported heparin was not started in the purge. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26801 | 1220648-2024-10996 | 19285514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19285514&pc=OZD | 02/24/2022 | 05/15/2025 |  | 05/10/2024 |  |  | 02/24/2022 | 05/10/2024 |  |  | 11/30/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162967 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2022 |  |  |  |  | The user facility reported a patient noted as high-risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. the impella cp was unable to advance past the common iliac and the left femoral artery were occluded. the impella cp insertion was aborted due to the patient's vascular anatomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26802 | 1220648-2024-10997 | 19285552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19285552&pc=OZD | 05/03/2024 | 06/24/2024 |  | 05/10/2024 |  |  | 05/03/2024 | 05/10/2024 |  |  | 09/30/2025 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402864 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2024 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during the impella cp insertion that purge tubing and fluid was not connected to impella purge sidearm. the cp was removed due to likely blood in motor. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the reported purge leak issue has been completed. data logs nor the device were returned for evaluation. however, clinical details stated that there was improper priming of initial implanted impella and upon insertion, the pump immediately had a leak and was therefore removed. the root cause of the leaking issue was most likely use related as there was no priming of the pump that was inserted and it was not connected to the purge tubing at the pump sidearm. the instructions for use referenced in the initial report is for the impella cp impella cp with smartassist for use during cardiogenic shock and high risk cpi. added- h6 impact code 4629 d4-corrected model number, corrected udi #.
26803 | 1220648-2024-10998 | 19285614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19285614&pc=OZD | 11/04/2021 | 05/14/2024 |  | 05/10/2024 |  |  | 05/02/2024 | 05/10/2024 |  |  | 08/31/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022126020 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The complainant reported a 50-year old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that purge system damage occurred during purge cassette change after an impella 5.5 exchange. no patient harm was reported. | [5 empty] | The investigation for the reported purge leak has been completed. the root cause of the purge leak was unable to be determined as the pump was not returned for investigation. this report is being filed as part of a retrospective review of historical records. corrected information provided in h6 and h10. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
26804 | 1220648-2024-10999 | 19285828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19285828&pc=OZD | 04/12/2024 | 05/31/2024 |  | 05/10/2024 |  |  | 04/12/2024 | 05/10/2024 |  |  |  |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025468584 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported an 81-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on support, the pump displayed "high purge pressure" alarm. the staff noticed high purge pressure and low purge flows and eventually the purge flow became blocked. intravenous sugar solution was started in the purge and this resulted in a marginal change in purge flow. the pump would not run higher than p-1. the pump was removed and exchanged. there was no patient harm. | [5 empty] | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the high purge pressure issue was not determined since product and logs were not returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26805 | 1220648-2024-11000 | 19285872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19285872&pc=OZD | 02/03/2022 | 05/15/2025 |  | 05/10/2024 |  |  | 02/03/2022 | 05/10/2024 |  |  | 10/31/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022150373 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/03/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced sustained ventricular fibrillation. percutaneous coronary insertion was aborted and the impella was removed. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The device passed all post-sterile inspection checks. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant). ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records.
26806 | 1220648-2024-11001 | 19285917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19285917&pc=OZD | 02/07/2022 | 05/10/2024 |  | 05/10/2024 |  |  | 02/07/2022 | 05/10/2024 |  |  | 07/31/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114064 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/07/2022 |  |  |  |  | The complainant reported a patient implanted with an impella cp device for mechanical circulatory support due to st-segment elevation myocardial infarction. while on support, the physician completed an external fem/fem bypass to facilitate blood flow past the sheath. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26807 | 1220648-2024-11002 | 19285961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19285961&pc=OZD | 01/29/2022 | 05/15/2025 |  | 05/10/2024 |  |  | 01/29/2022 | 05/10/2024 |  |  | 09/30/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022136904 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, pulsatile oozing was observed at the access site. it was reported that anticoagulation was reversed, angle matching techniques were performed, and a femstop was placed to resolve this issue. the patient remained on impella support, and the procedure was completed successfully. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26808 | 1220648-2024-11003 | 19286046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19286046&pc=OZD | 02/07/2022 | 05/10/2024 |  | 05/10/2024 |  |  | 02/07/2022 | 05/10/2024 |  |  | 11/30/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163086 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/07/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the patient had hemashield platinum graft with lots of oozing and lots of product given. heparin was stopped for a period during recovery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26809 | 1220648-2024-11004 | 19286192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19286192&pc=OZD | 02/03/2022 | 05/15/2025 |  | 05/10/2024 |  |  | 02/03/2022 | 05/10/2024 |  |  | 11/30/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158386 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/03/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was found that both legs were cool and there were no foot pulses. recommendations were given regarding volume management. additionally, heparin-induced thrombocytopenia (hit) was confirmed. therefore, heparin was withheld and argatroban was increased several times. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The device passed all post-sterile inspection checks. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the cause of the issue was not determined due to insufficient clinical details. this report is being filed as part of a retrospective review of historical records.
26810 | 1220648-2024-11005 | 19286252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19286252&pc=OZD | 02/07/2022 | 05/10/2024 |  | 05/10/2024 |  |  | 02/07/2022 | 05/10/2024 |  |  | 10/31/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022150401 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/07/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while under support, during transport to the cardiovascular intensive care unit, the patient experienced an episode of ventricular tachycardia requiring multiple cycles of cardioversion. weaning was successful, and the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26811 | 1220648-2024-11006 | 19286263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19286263&pc=OZD | 02/03/2022 | 05/15/2025 |  | 05/10/2024 |  |  | 02/03/2022 | 05/10/2024 |  |  | 09/30/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022128447 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/03/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella support, the patient was found to have sepsis. it is unknown how the sepsis was resolved, but the patient remained on impella support and survived. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | An infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: patient symptoms and medical history obtained by the treating clinician. physical examination findings. results of special investigations: laboratory test results and imaging studies. microbiological culture and sensitivity results. response to previous treatments and therapies. any relevant epidemiological information or exposure history. concomitant devices in use. preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. care of access site. because this evidence has not been provided, the root cause of the infection cannot be determined. the device passed all post-sterile inspection checks. this report is being filed as part of a retrospective review of historical records.
26812 | 1220648-2024-11007 | 19286588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19286588&pc=OZD | 02/28/2022 | 05/10/2024 |  | 05/10/2024 |  |  | 02/28/2022 | 05/10/2024 |  |  | 01/31/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022047635 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Unknown |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 02/25/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced a gastrointestinal (gi) bleed and heparin was changed from systematic to purge only. the gi bleed resolved and the patient was put back on systematic heparin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26813 | 1220648-2024-11008 | 19286642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19286642&pc=OZD | 02/08/2022 | 05/10/2024 |  | 05/10/2024 |  |  | 02/08/2022 | 05/10/2024 |  |  | 09/30/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142344 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/08/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support after a segment elevation myocardial infarction and the patient decompensated. while on support, the perclose sutures were tightened for oozing and bleeding stopped, however 2 units of blood products were given. in addition, it was noted that the echocardiogram was measuring the impella at 1.0 cm. it was later reported that the patient was doing well, and drips were weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26814 | 1220648-2024-11009 | 19286646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19286646&pc=OZD | 04/22/2024 | 11/22/2024 |  | 05/10/2024 |  |  | 04/14/2024 | 05/10/2024 |  |  |  |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451702 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 43 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Sepsis | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/14/2024 |  |  |  |  | The user facility reported a 43-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support and had hemolysis with increasing inotropic support and low cardiac power output. the physician removed the impella cp and escalated to an impella 5.5. | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured ventricular tachycardia with a "probable" relationship to the impella device used: ¿on (b)(6) 2024, during the pci procedure the patient had recurrent episodes of ventricular tachycardia requiring defibrillation. the patient had multiple episodes throughout the procedure with approximately 20 defibrillation episodes.¿ the patient recovered with sequelae. the complaints team received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured cerebrovascular accident with a "probable" relationship to the impella device used: ¿at 0805 on (b)(6) 2024, the nurse caring for patient documented that the patient had a fixed deviated gaze to the left with nystagmus being present. patient was unable to look to the right at all. he was able commands in left upper and lower extremities. no movement noted and does not withdraw to pain in right upper or lower extremities. md was notified and at bedside. code stroke called. the neurologist later noted that patient developed new right hemiplegia with left pca stroke which is at least 48 hrs old and left mca stroke on (b)(6)¿. the patient recovered with sequelae. the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured impella site infection with a "possible" relationship to the impella device used: ¿subject was admitted on (b)(6) 2024 for sever sepsis. a debridement, wound exploration, right chest impella site, incision and drainage, abscess was preformed¿. the patient recovered with no sequelae. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ the following have been reported incorrectly or missing from manufacturer device report: 1220648-2024-11009: b2: updated. b3: date of event. b5: updated with additional information. d6b: inadvertently reported incorrect date. f6/f8: should have been left blank. g1: reporting contact fax number missing. g2: report source; study missing. h6: updated to reflect additional failure modes. h6: component code revised from the initial submission in accordance with updated procedures. h10: instructions for use added to reflect additional failure modes. | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
26815 | 1220648-2024-11010 | 19286820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19286820&pc=OZD | 02/11/2022 | 05/15/2025 |  | 05/10/2024 |  |  | 02/11/2022 | 05/10/2024 |  |  | 09/30/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142991 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/11/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, it was noted that there was calcium in common femoral artery. therefore, the impella could not advance the sheath all the way inside the artery. a shockwave balloon was used to expand arterial wall to pass the catheter. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26816 | 1220648-2024-11011 | 19286842 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19286842&pc=OZD | 04/28/2024 | 06/10/2024 |  | 05/10/2024 |  |  | 04/28/2024 | 05/10/2024 |  |  |  |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025454510 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Heart Block | [12 empty] | Required Intervention | [5 empty] | T-PACER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/28/2024 |  |  |  |  | The user facility reported a 63-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after the implant, the impella appeared deep. the impella was pulled back six centimeters and the patient went into complete heart block. the patient had no rhythm issues prior. a t-pacer was placed in left femoral vein. a pericardial drain was needed due to the t-pacer implant. the patient was reported as stable. | [5 empty] | The investigation into the vf/vt/af/at (complete heart block) issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26817 | 1220648-2024-11012 | 19287129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19287129&pc=OZD | 02/11/2022 | 05/15/2025 |  | 05/10/2024 |  |  | 02/11/2022 | 05/10/2024 |  |  | 10/31/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022155899 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/11/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hemolysis with no specified resolution. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26818 | 1220648-2024-11013 | 19287238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19287238&pc=OZD | 02/11/2022 | 04/16/2025 |  | 05/10/2024 |  |  | 02/11/2022 | 05/10/2024 |  |  | 10/31/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157856 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/11/2022 |  |  |  |  | The user facility reported a 68-year-old patient needing mechanical circulatory support. it was reported that the patient was on impella cp support and the patient had too much calcium build up and the device was unable to be placed. the physician switched sides to right femoral artery (rfa) and ultrasound guided micropuncture was obtained. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26819 | 1220648-2024-11014 | 19287289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19287289&pc=OZD | 04/30/2024 | 06/13/2024 |  | 05/10/2024 |  |  | 04/30/2024 | 05/10/2024 |  |  |  |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492924 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | The user facility reported an 86-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had an impella cp implant and there was oozing and a hematoma at the access site. manual pressure was applied and the hematoma was expressed. the site was redressed utilizing an angle match dressing and after the site was no longer oozing. heparin was temporarily discontinued. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was anticoagulation management due to high patient act values, heparin was stopped to achieve hemostasis as per the clinical description.
26820 | 1220648-2024-11015 | 19287358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19287358&pc=OZD | 02/11/2022 | 05/10/2024 |  | 05/10/2024 |  |  | 02/11/2022 | 05/10/2024 |  |  | 11/30/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023162954 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/11/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the iliac site required clot aspiration and surgical wound closure. it was reported the patient recovered very well and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves as noted in the impella ifu ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
26821 | 1220648-2024-11016 | 19287455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19287455&pc=OZD | 02/12/2022 | 05/10/2024 |  | 05/10/2024 |  |  | 02/12/2022 | 05/10/2024 |  |  | 08/31/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022137434 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/12/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced oozing at the groin site. to resolve the oozing the dressing was removed, the sheath advanced, and the site was sutured and redressed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26822 | 1220648-2024-11017 | 19287619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19287619&pc=OZD | 02/13/2022 | 05/10/2024 |  | 05/10/2024 |  |  | 02/13/2022 | 05/10/2024 |  |  | 10/31/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149206 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/13/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding, the physician was unable to stop the bleeding which led to the impella being removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26823 | 1220648-2024-11018 | 19287808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19287808&pc=OZD | 05/05/2024 | 06/19/2024 |  | 05/10/2024 |  |  | 05/05/2024 | 05/10/2024 |  |  |  |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025474228 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/05/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. the patient's vasculature was too tight. the long sheath was used and the physician attempted three times however, they were unsuccessful. the implant was aborted. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition since the pump could not be placed due to vascular being too tight to pass pump along with axillary access found to have several large lung masses. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
26824 | 1220648-2024-11019 | 19287833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19287833&pc=OZD | 05/05/2024 | 06/26/2024 |  | 05/10/2024 |  |  | 05/05/2024 | 05/10/2024 |  |  |  |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025444367 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/05/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as escalation of therapy from an impella cp. while on support, the patient developed oozing at the access site. the patient was given packed red blood cells, platelets, fresh frozen plasma, and protamine. the patient remains on impella 5.5 support on the date of this report. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ potential adverse events: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient information is provided in the clinical description.
26825 | 1220648-2024-11020 | 19287848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19287848&pc=OZD | 04/29/2024 | 05/10/2024 |  | 05/10/2024 |  |  | 05/07/2024 | 05/10/2024 |  |  | 05/31/2025 |  | 05/10/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024350496 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | Component Incompatible | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 05/07/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported that an impella controller generated an error "please connect purge cassette". the purge cassette was replaced to troubleshoot, but the issue was resolved upon replacement of the controller. no harm to the patient was reported. | [5 empty] | The impella device was received from the customer on 08-may-2024. data review: after reviewing the imc logs, the issue with the user getting stuck at the ¿click purge disc, then close purge door¿ prompt was confirmed. there were numerous purge disc not detected alarms (event set #402) and complete the procedure (user interaction stopped) alarms found after pc_display_purge_disc_help_screen in the logs. device analysis: console was tested in collaboration with field service. clinical staff reported that a please connect purge housing screen was displayed but it is likely that they meant the ¿click purge disc, then close purge door¿ prompt. after examination of the console, it was discovered that the purge disc flag was broken (missing at blue retainer). this broken purge flag is what caused the user to get stuck on the please insert the purge disc prompt because it was unable to detect that the purge disc was inserted. before sending this console back to the customer, field service was instructed to replace the purge disc detect flag [3001198] and purge flag spring [3001196], validate sensor accuracy via section 12 of pm procedure [0042-7309], and perform a full functional check on the console and optical system [0042-7316].
26826 | 1220648-2024-11021 | 19287935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19287935&pc=OZD | 05/07/2024 | 06/24/2024 |  | 05/10/2024 |  |  | 05/07/2024 | 05/10/2024 |  |  |  |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025494751 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  |  |  | Malfunction | No | Yes | Difficult or Delayed Positioning | Difficult to Remove | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient planned for high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. however, the pump appeared shallow. therefore, the physician advanced the pump into the proper position and flows dropped with suction alarms. it was noted the pump was in a good position; however, the patient became hemodynamically unstable. the physician, therefore, chose to remove the impella pump. the pump turned off when pulled back into aorta. the physician attempted to remove the pump through the peel away but still had the single access wire in the sheath so the pump became caught at the distal end of the sheath. the pump and peel away sheath were eventually removed as one piece. a new impella cp device was placed, and the team proceeded with the percutaneous coronary intervention. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow and the removal issues has been completed since the original report was filed. the pump was returned for investigation. no biomaterial was found around the impeller. there are no other visible abnormalities. the root cause of the low pump flow issue was most likely due to a kinked cannula. the root cause of the removal issues was most likely use related as they attempted to remove the pump through the peel away with the single access wire still inside and the pump got caught on it; the peel away and pump had to be removed as one piece.
26827 | 1220648-2024-11022 | 19287973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19287973&pc=OZD | 02/16/2022 | 05/14/2025 |  | 05/10/2024 |  |  | 02/16/2022 | 05/10/2024 |  |  | 09/30/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022146555 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was escalated to an impella 5.5 device to evaluate the patient for aortic valve replacement and mitral valve replacement versus left ventricular assist device for transplant workup. while on support the patient experienced bleeding, to address the bleeding systemic heparin was withheld. it was reported that the patient went into atrial fibrillation with rapid ventricular response. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26828 | 1220648-2024-11023 | 19288148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19288148&pc=OZD | 05/10/2022 | 05/15/2025 |  | 05/10/2024 |  |  | 05/10/2022 | 05/10/2024 |  |  | 01/31/2024 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192468 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2022 |  |  |  |  | The user facility reported a 77-year-old female on impella cp for mechanical circulatory support. it was reported the patient¿s pulses in the left leg (impella cp) were more diminished, and that leg was cooler to the touch than days prior. it was decided to implant the impella rp for bipella support. after the impella rp was advanced, it was ramped up to p-9 and the impella cp was able to be increased in support. it was also reported the patient¿s platelet count dropped. the decision was made to explant the cp and leave the impella rp in as support. heparin was withheld and blood products were given to the patient. manual pressure was also applied, but hemostasis was not achieved. they physician made the decision to place the fem-stop. with much difficulty, the fem-stop was successfully placed and inflated. hemostasis was achieved and pulse by doppler in foot. the patient also developed a hemotoma in which a an angiogram was performed and a fem-stop was applied. an acute bleed at the left femoral head was also noted. neo drip was increased, 2 units of blood was given and a stent was deployed. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26829 | 1220648-2024-11024 | 19288286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19288286&pc=OZD | 02/26/2022 | 05/10/2024 |  | 05/10/2024 |  |  | 02/26/2022 | 05/10/2024 |  |  | 12/31/2023 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173619 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/26/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced a large hematoma on the left groin. a vascular cut down and repair was required to resolve the issue. the patient was reported as stable post procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26830 | 1220648-2024-11025 | 19288750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19288750&pc=OZD | 09/02/2022 | 05/15/2025 |  | 05/10/2024 |  |  | 09/02/2022 | 05/10/2024 |  |  | 06/30/2024 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231314 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/02/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on support, vascular dissection in the iliac was noted. vascular surgery was performed, and stent was placed in the right iliac. impella cp was successfully weaned and explanted. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26831 | 1220648-2024-11026 | 19289056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19289056&pc=OZD | 09/02/2022 | 05/15/2025 |  | 05/10/2024 |  |  | 09/02/2022 | 05/10/2024 |  |  | 03/31/2024 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211487 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/02/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, access site bleeding and hemolysis was noted. manual compression was done to the site to resolve the bleeding. hemostasis was achieved. impella was successfully weaned and explanted. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event (access site bleeding) are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system instructions for use for the related event (hemolysis) are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | This complaint has been identified as a part of retrospective review. the device passed all post-sterile inspection checks. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined.
26832 | 1220648-2024-11027 | 19289295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19289295&pc=OZD | 09/04/2022 | 05/15/2025 |  | 05/10/2024 |  |  | 09/04/2022 | 05/10/2024 |  |  | 03/31/2024 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204907 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed oozing from the access site, nose and mouth. purge heparin was decreased to 25 u/ml and there was no systematic heparin. vasopressors were weaned, activated clotting time (act) was down in low 200s and the bleeding was slowed. additionally, the patient developed gastrointestinal bleeding. plasma free hemoglobin was recommended. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event (access site bleeding) are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system instructions for use for the related event (vascular dissection) are as follows: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26833 | 1220648-2024-11028 | 19289487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19289487&pc=OZD | 09/06/2022 | 05/15/2025 |  | 05/10/2024 |  |  | 09/06/2022 | 05/10/2024 |  |  | 04/30/2024 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216198 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/06/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was noted that the patient experienced bleeding and was transitioned to d5w/sodium bicarb purge. the impella was successfully weaned and explanted. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26834 | 1220648-2024-11029 | 19289745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19289745&pc=OZD | 09/06/2022 | 05/15/2025 |  | 05/10/2024 |  |  | 09/06/2022 | 05/10/2024 |  |  | 05/31/2024 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228729 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/06/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, patient was found to have aortic root dissection. impella was successfully weaned and explanted. decision was made to monitor and treat the dissection moving forward. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | This complaint has been identified as part of a retrospective review. the device passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26835 | 1220648-2024-11030 | 19290057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19290057&pc=PYX | 04/12/2024 | 05/28/2024 |  | 05/10/2024 |  |  | 04/12/2024 | 05/10/2024 |  |  |  |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025448706 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella rp flex device for mechanical circulatory support. it was reported that the impella rp flex stopped and would not get flows above 0.5. the impella rp flex was exchanged. the patient expired. there is not enough information to exclude use of the device with the patient's outcome. | [5 empty] | The investigation has been completed. no product was returned. the data logs were reviewed, motor current spike greater than 1650 ma with temporary pump stop. the pump was able to be restarted later, but flows were observed to be lower. no evident biomaterial interaction. cannula kinks were observed. all other clinical information was unknown. the cause of the temporary pump stop issue was not determined since the product was not returned and due to inconclusive evidence from log review. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp flex with smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
26836 | 1220648-2024-11031 | 19290106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19290106&pc=OZD | 09/14/2022 | 05/30/2025 |  | 05/10/2024 |  |  | 09/14/2022 | 05/10/2024 |  |  | 06/30/2024 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236861 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2022 |  |  |  |  | The complainant reported a patient with a history of recurrent ventricular tachycardia was implanted with an impella cp device for mechanical circulatory support for a ventricular tachycardia ablation. while on support, the pump experienced suction alarms that were resolved by infusing the patient with albumin. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26837 | 1220648-2024-11034 | 19292470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19292470&pc=OZD | 04/13/2024 | 05/30/2024 |  | 05/10/2024 |  |  | 04/13/2024 | 05/10/2024 |  |  |  |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446454 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was confused and worked out of the restraints, grabbed the catheter and pulled hard enough to separate the catheter from the red plug. subsequently, the patient expired. it was determined that the death was associated to the event and that there is not enough evidence to say that our device did not contribute to the patient's outcome. | [5 empty] | The investigation for the reported product damage has been completed. the cp pump was not returned for investigation. data logs were reviewed, and alarm 115 pump stop with motor current spike to greater than 30000 ma was observed. the pump was unable to be restarted and no other abnormalities were found. as per clinical, patient was confused and pulled hard enough to separate the catheter from the red plug. the cause of the pump stop issue was most likely an open electrical line due to patient movement per log and clinical review. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, section h10 has been revised from the initial submission. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26838 | 1220648-2024-11032 | 19292637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19292637&pc=OZD | 04/12/2024 | 06/04/2024 |  | 05/10/2024 |  |  | 04/12/2024 | 05/10/2024 |  |  | 06/30/2025 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369987 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/12/2024 |  |  |  |  | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured thrombocytopenia with a "definite" relationship to the impella device used: ¿the patient developed thrombocytopenia. additionally, patient also endured hemolysis with a "definite" relationship to the impella device used: the patient developed hemolysis. care was withdrawn and the patient expired. | The user facility reported a 56-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support and developed ischemia. an antegrade stick was used to provide extra flow. care was withdrawn and the patient expired. there is not enough information to associate the use of the device with the patient's outcome. | The patient expired after the family withdrew care. the patient 's outcome was reviewed by abiomed's medical safety officer (mso). it was determined that there was not enough information to associate use of the device with patient outcome. decision already made from initial. |  |  |  | B5, g2 and h6 has been updated as per additional information received. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia, thrombocytopenia, and hemolysis has been completed. the product was not returned. the root cause of the thrombocytopenia was not determined as limited clinical details were provided. data log analysis was conducted. there is one instance of a ¿impella position unknown¿ alarm. there are no suction events, motor current spikes, or a trend in the placement signal. it is unknown when the hemolysis occurred or resolved. the root cause of the hemolysis cannot be determined. without additional clinical details, the root cause of the limb ischemia could not be determined. the instructions for use for the impella cp with smartassist system in section: use of echocardiography for positioning of the impella catheter states ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
26839 | 1220648-2024-11033 | 19293203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19293203&pc=OZD | 04/11/2024 | 06/20/2024 |  | 05/10/2024 |  |  | 04/11/2024 | 05/10/2024 |  |  | 09/30/2025 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024400266 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support, a 57-year-old male patient experienced limb ischemia. the patient's leg was noted to be marbled and changed color. the patient expired on impella support. there is not enough information to exclude use of the device with the patient's outcome. | [5 empty] | The investigation into the limb ischemia - reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26840 | 1220648-2024-11035 | 19293457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19293457&pc=OZD | 05/03/2024 | 07/19/2024 |  | 05/10/2024 |  |  | 05/07/2024 | 05/10/2024 |  |  |  |  | 05/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1359283 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/07/2024 |  |  |  |  | Complainant in japan reported the automated impella controller (aic) was turned on and when the staff tried to turn the power on the aic off, it froze. the sound was still playing when it froze. the aic was forcefully shutdown and an aic failure alarm appeared. the on and off buttons did not click when pressed. no patient harm or use was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. e1 was updated to unknown as information was not currently available. | After reviewing the logs, the reported issue with freezing was not confirmed. the console was tested in collaboration with field service. the reported issue was unable to be reproduced through testing. the root cause of the reported frozen automation impella controller (aic) could not determined due to the inability to reproduce. added d9 date device returned to manufacturer d2 common device name f6/f8-should have been left blank in the initial report.
26841 | 1220648-2024-11036 | 19293510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19293510&pc=OZD | 04/13/2024 | 06/21/2024 |  | 05/10/2024 |  |  | 04/13/2024 | 05/10/2024 |  |  |  |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2025459067 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that during impella 5.5 support for a 66-year-old male, the purge pressure was too high and vented several times. there were no kinks observed on the line. the purge pressure rose to approximately 820 and then dropped to over 1000. pump was exchange was recommended, and lysis protocol was sent. tissue plasminogen activator was added to the purge and the motor current was monitored. there was no reported patient harm. the patient expired on support after care was withdrawn. there is not enough evidence to say not our device. | [5 empty] | The investigation for the reported high purge pressure has been completed. clinical details state that tissue plasminogen activator (tpa) was added to the purge. high purge pressure remained before pump stop the following day. the data logs were reviewed which confirmed high purge pressure and low purge flow that did not resolve. there were also purge system blocked alarms before the pump stop that occurred due to high motor current. there was a gradual rise in purge pressure before alarm was triggered. the root cause of pump stop was most likely high purge pressure but the cause of the high purge pressure was unknown because no product was returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
26842 | 1220648-2024-11037 | 19293736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19293736&pc=OZD | 04/11/2024 | 06/26/2024 |  | 05/10/2024 |  |  | 04/12/2024 | 05/10/2024 |  |  | 10/31/2025 |  | 05/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414243 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 75-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump was placed for six days when it was explanted for concerns of hemolysis. the team explanted and escalated to the impella 5.5 pump. the patient expired. there is not enough information to exclude use of the device with the patient's outcome. | [5 empty] | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. b2-should have been left blank for death, remove date of patient death f6/f8-should have been left blank in the initial report. h1-changed to serious injury h6 impact code-remove 1802. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
26843 | 1220648-2024-10925 | 19275767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19275767&pc=OZD | 03/01/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 03/01/2022 | 05/09/2024 |  |  | 12/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173054 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/01/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion (hrpci) was implanted with an impella cp device for mechanical circulatory support. placement of the device was not successful in the right femoral region. the sheath was only advancing antegrade towards the foot. patient became hypotensive requiring blood and pressors after multiple attempts to place the cp through the oscor sheath. additionally, there was bleeding through the 14f dry sheath. the pump and sheath were removed. preclose closed. pci was aborted and the patient was sent to the intensive care unit (icu). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system instructions for use for the related event (delivery issues) are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26844 | 1220648-2024-10927 | 19276407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19276407&pc=OZD | 05/01/2024 | 07/03/2024 |  | 05/09/2024 |  |  | 05/01/2024 | 05/09/2024 |  |  | 07/31/2025 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435390 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2024 |  |  |  |  | The user facility reported a 58-year-old male with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the patient developed a hematoma and oozing at the access site. the hematoma healed quickly with manual pressure but skin tract bleeding continued and worsen. the medical staff resolved the bleeding with tranexamic acid soaked gauze and additional forward tension sutures. the next day, a thrombus was evident on echocardiogram. the plans were made to treat the thrombus with anticoagulants and to explant the impella. the patient remains on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the access site bleed and thrombus has been completed. the investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, the bleeding was resolved with tranexamic acid gauze and additional forward tension sutures as per the clinical description. as the necessary information was not provided, the root cause of the thrombus was not determined. additional information d6b explant date f6/f8-should have been left blank in the initial report.
26845 | 1220648-2024-10928 | 19276607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19276607&pc=OZD | 04/11/2024 | 06/05/2024 |  | 05/09/2024 |  |  | 04/11/2024 | 05/09/2024 |  |  | 10/31/2025 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424968 |  | 048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported, a patient in acute myocardial infarction/cardiogenic shock. was implanted with an impella cp device, for mechanical circulatory support. after turning on the catheter, "impella stopped retrograde flow" alarm presented. it was noted, all possible troubleshooting steps were completed. however, the alarm condition did not resolve. the decision was made to remove and replace the impella pump. however, the patient decompensated and expired. there is not enough information to exclude, the use of the device with patient's outcome. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the low flow was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received, from the customer. and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26846 | 1220648-2024-10929 | 19276873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19276873&pc=OZD | 10/02/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 10/02/2022 | 05/09/2024 |  |  | 02/29/2024 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203086 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/02/2022 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient was transferred to her bed and immediately following transfer, the impella alarmed for "impella in aorta". the medical team confirmed tuohy-borst was still locked, and concluded the impella was pulled back into the aorta during patient transfer. the physician advanced the impella back into the left ventricle and resolved the alarm. it was also reported the patient developed hematuria and was noted to have red urine at the time of explantation. the medical team noted that the two perclose device sutures had been cut after deployment, and the two perclose devices failed. the access site began bleeding, and the obturator and 11 french sheath were replaced but hemostasis was not achieved. fluoroscopy was performed to observe the sheath was extravascular and was it removed and reinserted successfully. manual pressure was held to achieve hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26847 | 1220648-2024-10930 | 19277274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277274&pc=OZD | 05/06/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 05/06/2022 | 05/09/2024 |  |  | 06/30/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022099172 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/06/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was escalated from the impella cp to the impella 5.5 pump. it was reported while on impella 5.5 support, blood and clots were lavaged from an orogastric (og) tube. heparin was withheld, and patient observed with no further events noted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. analysis summary the failure mode of "thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
26848 | 1220648-2024-10931 | 19277337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277337&pc=OZD | 05/01/2024 | 07/15/2024 |  | 05/09/2024 |  |  | 05/03/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467765 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Device Handling Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2024 |  |  |  |  | It was further reported that during support, the patient was given a "do not resuscitate" order and care was withdrawn. the patient subsequently expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there is not enough information to exclude the impella as an associated factor in the patient's outcome. | The user facility reported a 79-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella 5.5 pump perforated the patient's left ventricle during a ventricular septal defect repair. it was stated that the perforation was due to the friable tissue in conjunction with the surgeon manipulating the heart. the site was surgically managed, and the scheduled repair was completed. support with the pump was maintained following the event. |  |  |  |  | B.2, b.5, d.6b and h.6 health effect - impact code: additional information was received and abiomed's medical safety office review was completed. the investigation for the reported ventricle perforation has been completed. the impella was not returned for evaluation. perforated during ventricular septal defect repair due to very friable tissue while surgeon manipulated heart was reported. the root cause of the ventricle perforation was most likely patient condition related due to friable left ventricle tissue and heart manipulation. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ b.2 added required intervention as it was inadvertently omitted from manufacturer device report 1220648-2024-10931. b.7 revised as this information was inadvertently omitted from manufacturer device report 1220648-2024-10931. d.4 revised model number from manufacturer device report 1220648-2024-10931 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-10931 and should not have been. g.1 revised reporting contact facility name and reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-10931 was submitted. h.1 was revised from manufacturer device report 1220648-2024-10931 due to medical safety officer review. h.2 additional information was selected incorrected on manufacturer device report 1220648-2024-10931. there should not of been anything selected in h.2 as that was the original report. h.6 codes 3265 and 463 were reported incorrectly on manufacturer device report 1220648-2024-10931. h.10 additional manufacturer narrative instructions for use were revised from manufacturer device report 1220648-2024-10931 in accordance with updated procedures. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
26849 | 1220648-2024-10932 | 19277349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277349&pc=OZD | 05/01/2024 | 06/13/2024 |  | 05/09/2024 |  |  | 05/01/2024 | 05/09/2024 |  |  | 08/31/2025 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024405213 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention experienced an unsuccessful implant with the first impella cp device for mechanical circulatory support. the impella cp pump was inserted and tracked as far as the abdominal aorta before stopping and could not be advanced any further. there were no indications of what was causing the issue on fluoroscopy, and the impella cp pump had exited the sheath. the decision was made to remove this impella cp pump and change out impella 14fr x 13 sheath for the impella 14fr x 25 sheath. thereafter, the physician attempted implant again with a new impella cp pump, which was successfully placed. the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition since the insertion was aborted due to patient anatomy.
26850 | 1220648-2024-10933 | 19277357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277357&pc=OZD | 05/02/2024 | 06/25/2024 |  | 05/09/2024 |  |  | 05/02/2024 | 05/09/2024 |  |  | 06/30/2025 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024370029 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The complainant reported an 63-year old male patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient developed a hematoma at the access site. medical staff held pressure at the site for 20 minutes and placed a femoral stop. the patient then developed right limb ischemia, which prompted medical staff to remove the femoral stop. the hematoma reformed, which prompted medical staff to escalate the patient to an impella 5.5. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia and access site bleed has been completed. the device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. the root cause of the access site bleeding was most likely due to patient movement since the hematoma formed after the patient transfer. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures.h.6 codes 1888, 4627, 4642, 4625, 925 and 3038 were reported incorrectly on the previously submitted report.
26851 | 1220648-2024-10934 | 19277365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277365&pc=OZD | 05/02/2024 | 07/11/2024 |  | 05/09/2024 |  |  | 05/02/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446509 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The complainant reported an 82-year old male patient with a high risk of percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. the complainant reported oozing from the patient¿s access site. medical staff controlled the bleeding with a femoral stop and heparin was temporarily withheld. the patient was eventually weaned, with the cp eventually explanted, and remains stable | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. f6 and f8 should have been blank. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26852 | 1220648-2024-10935 | 19277403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277403&pc=OZD | 05/02/2024 | 07/08/2024 |  | 05/09/2024 |  |  | 05/04/2024 | 05/09/2024 |  |  | 02/28/2025 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024320840 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2024 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. an impella cp "purge liquid low" alarm occurred on the sixth day of use. the connection was retightened and the yellow connector broke and it started leaking. a purge liquid alarm appeared so the impella cp was exchanged. there was no harm to the patient. it was noted that sodium bicarbonate was not used. | [5 empty] | The investigation for the reported purge leak has been completed. the pump was returned for evaluation. during evaluation, there were several cracks along the yellow luer characteristic of excessive force which align with the clinical details. clinical details states that the yellow luer was damaged during re-tightening efforts after purge fluid leak was observed. the root cause of the purge leak-pump was determined to be a damaged polycarbonate yellow luer due to excessive force. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high risk cpi in section: cleaning. added- d9 device return date. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
26853 | 1220648-2024-10936 | 19277466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277466&pc=OZD | 05/03/2024 | 06/26/2024 |  | 05/09/2024 |  |  | 05/03/2024 | 05/09/2024 |  |  | 11/30/2025 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024423606 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support, and during support, the pump position issue was encountered along with hemolysis. it was recommended to reposition the pump. however, it was determined the pump came out of its position, and the decision was made to explant the impella cp pump. | [5 empty] | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
26854 | 1220648-2024-10938 | 19277487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277487&pc=OZD | 05/04/2024 | 09/22/2025 |  | 05/09/2024 |  |  | 05/01/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447063 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/01/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a drop in hemoglobin, hematocrit and platelets. two units of blood products and one platelet were provided and heparin was withheld. the patient remained on impella cp support and was successfully weaned four days after implant. | Additional information indicated that, during patient support, the pump experienced placement signal unreliable alarm. |  |  |  |  | The device was not returned for evaluation. at the conclusion of the investigation, a supplemental report will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the optical signal issue and the thrombocytopenia issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the false position in aorta alarms was most likely patient condition related. the cause of the thrombocytopenia issue was not determined due to insufficient clinical information. | B1 updated with product problem. b5 updated with additional information from the initial manufacturer device report number 1220648-2024-10938. d4 model number updated. f6 and f8 should have been left blank. g1 updated with correct mdr reporting contact email. g1-updated g1 manufacturing site name and address. h6 medical device problem code revised to include 2917 device sensing problem and component code 925 pump from the initial manufacturer device report number 1220648-2024-10938.
26855 | 1220648-2024-10937 | 19277489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277489&pc=OZD | 05/04/2024 | 06/10/2024 |  | 05/09/2024 |  |  | 05/04/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025457945 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Component Incompatible | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2024 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the aic (automated impella controller) was not recognizing the purge disc during cassette change despite trying to two cassettes. the aic was exchanged. there was no patient harm. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the purge disc flag was found to be broken. the root cause of the issue was a broken purge disc flag. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26856 | 1220648-2024-10939 | 19277535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277535&pc=OZD | 05/01/2024 | 06/10/2024 |  | 05/09/2024 |  |  | 05/01/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025464470 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and began to experience limb ischemia. it was noted that the patient femoral artery was small prior to implant and likely contributed to the development of the limb ischemia. insertion of a reperfusion sheath was attempted but eventually aborted when distal flow was observed on angiogram. despite the confirmation of flow, the decision was ultimately made to explant the pump to fully resolve the ischemia. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system-section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
26857 | 1220648-2024-10940 | 19277636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277636&pc=OZD | 05/02/2024 | 07/11/2024 |  | 05/09/2024 |  |  | 05/02/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025453666 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The user facility reported an 80-year-old female patient undergoing cardiac proceudre was implanted with an impella cp device for mechanical circulatory support. it was reported that the physician didn¿t like the position of the impella cp as it appeared to be in the aorta. repositioning attempts were made but unsuccessful. the cp was removed and a new impella was placed. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issues has been completed. the pump was not returned for investigation. the data logs were reviewed and there was a "suction" alarm followed by a "position in aorta" alarm before the pump was explanted. no other abnormalities were observed. therefore, the root cause of the positioning issue was not determined since the product was not returned and insufficient clinical information was provided and data logs were inconclusive. b.7 revised as this information was inadvertently omitted from manufacturer device report number 1220648-2024-10940. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-10940 and should not have been. h.6 code 3009 was reported incorrectly on manufacturer device report number 1220648-2024-10940. a new code has been added to medical device problem codes.
26858 | 1220648-2024-10941 | 19277757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277757&pc=OZD | 05/01/2024 | 06/11/2024 |  | 05/09/2024 |  |  | 05/02/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025445474 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella 5.5 support had a stroke. the physician feels impella could have contributed to the stroke. stroke symptoms improving with the patient able to stand with assistance but deficit still left side with severe weakness in the left arm. plan for ballon pump placement in cath lab and then removal of 5.5 in the operation room. heparin was stopped with stroke code. no impella alarms during stroke code or prior indicating clot on device. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of stroke could not be determined as the device was not returned nor sufficient clinical details. d4-updated model number.
26859 | 1220648-2024-10942 | 19277839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277839&pc=OZD | 05/02/2024 | 09/06/2024 |  | 05/09/2024 |  |  | 05/02/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025464053 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Pleural Effusion | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 05/02/2024 |  |  |  |  | The complainant reported a 34-year old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient decompensated the day before while on an intra-aortic balloon pump. the complainant reported the patient developed a hematoma at the access site during 5.5 support. the patient was taken to the operating room, where medical staff evacuated the access site and placed a jackson-pratt drain to manage the hematoma. the patient remained on support following the intervention. | Additional narrative related to this event was provided via "abiomed's long-term outcome and quality indicator". it was noted that following evacuation of the site, there was diffuse bleeding from the graft anastomosis site. a suture was tightened and the bleeding stopped. a repair suture was applied with 5-0 prolene and the purse-string suture was subsequently tightened. a transesophageal echocardiogram also showed a large right pleural effusion and a 24fr right pleural chest tube was placed to drain the site. the patient recovered from the incident. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | Additional information provided in b5, g2, and h6. should additional relevant information become available, a supplemental report will be submitted.
26860 | 1220648-2024-10943 | 19277868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277868&pc=OZD | 02/23/2022 | 05/19/2025 |  | 05/09/2024 |  |  | 02/23/2022 | 05/09/2024 |  |  | 11/30/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168573 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was stated that the impella device was implanted and pulled across the valve to reposition, but the device was unable to return across the valve, therefore, an additional impella device was required. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26861 | 1220648-2024-10944 | 19277934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277934&pc=OZD | 04/11/2024 | 05/28/2024 |  | 05/09/2024 |  |  | 04/11/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025467774 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2024 |  |  |  |  | The user facility reported a 66-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support there was access site bleeding. a blood transfusion was performed. the staff assessed the vessel and ruled out any damage to vessel. the patient was deemed coagulopathic and no heparin was given. the patient expired after 14.82 hours of impella support. the cause of death is unknown. there is not enough information to exclude use of the device with the patients outcome. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely patient condition since the patient was deemed coagulopathic and continued to bleed. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26862 | 1220648-2024-10945 | 19277976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19277976&pc=OZD | 02/13/2022 | 04/17/2025 |  | 05/09/2024 |  |  | 02/13/2022 | 05/09/2024 |  |  | 11/30/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023159832 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the repo sheath cracked, the physician was contacted for a replacement. additionally the physician repositioned the impella into lv (left ventricle) to 3.5 from aortic annulus. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26863 | 1220648-2024-10946 | 19278070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19278070&pc=OZD | 02/14/2022 | 05/09/2024 |  | 05/09/2024 |  |  | 02/14/2022 | 05/09/2024 |  |  | 11/30/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165786 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/14/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced bleeding at the access site. systemic heparin was withheld and the patient was given packed red blood cells to resolve the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26864 | 1220648-2024-10947 | 19278124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19278124&pc=OZD | 02/13/2022 | 04/07/2025 |  | 05/09/2024 |  |  | 02/15/2022 | 05/09/2024 |  |  | 12/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161613 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/15/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced ventricular tachycardia and the impella was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. section b3: date of event has been corrected.
26865 | 1220648-2024-10948 | 19278175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19278175&pc=OZD | 02/25/2022 | 05/20/2025 |  | 05/09/2024 |  |  | 02/25/2022 | 05/09/2024 |  |  | 03/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022063617 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2022 |  |  |  |  | The user facility reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during explant of the impella, the preclose became entangled within the pigtail and the device was unable to be pulled back into the sheath. a 0.18 wire was placed to maintain access and the impella and sheath were removed. one (1) preclose was unsuccessful and manual pressure was held until vascular surgery could be consulted. the patient was brought to surgery for access closure and emergent aortic aneurysm repair was performed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26866 | 1220648-2024-10949 | 19278182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19278182&pc=OZD | 02/16/2022 | 05/12/2025 |  | 05/09/2024 |  |  | 02/16/2022 | 05/09/2024 |  |  | 07/31/2024 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022121210 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Asystole | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient intermittently went asystole throughout the case with multiple rounds of cardiopulmonary resuscitation (cpr) and defibrillation. a temporary pacer was placed, and the right hearth catheterization was completed. additionally, oozing form the insertion site was noted. the site was re-sutured, angle matching techniques were utilized, and a femstop was placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26867 | 1220648-2024-10950 | 19278253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19278253&pc=OZD | 02/18/2022 | 05/09/2024 |  | 05/09/2024 |  |  | 02/18/2022 | 05/09/2024 |  |  | 11/30/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169164 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/18/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced warm extremities and had no dopplerable pulses. the device was explanted. after the device was explanted, the flow returned and dopplerable pedal pulses were achieved. additionally, there was oozing at puncture site, resolution was not provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26868 | 1220648-2024-10952 | 19278396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19278396&pc=OZD | 04/10/2024 | 05/29/2024 |  | 05/09/2024 |  |  | 04/10/2024 | 05/09/2024 |  |  | 09/30/2025 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444718 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the same day of placement, there was a concern for right limb ischemia with extracorporeal membrane oxygenation escalation. popliteal pulses were felt on right leg (impella side), but their foot was cold and no pulses were felt. the right limb became mottled and the patient was on more vasopressors. the patient continued to decompensate and the decision was made to withdraw care and the patient expired. the cause of death is unknown. there is not enough information to exclude the use of the device with the patients outcome. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the pump was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26869 | 1220648-2024-10951 | 19278398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19278398&pc=OZD | 02/20/2022 | 05/09/2024 |  | 05/09/2024 |  |  | 02/20/2022 | 05/09/2024 |  |  | 08/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022130283 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/20/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced oozing at the groin access site around the sheath when the medication integrilin was stopped, reportedly the site was soft and oozing slowed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26870 | 1220648-2024-10953 | 19278419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19278419&pc=OZD | 02/25/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 02/25/2022 | 05/09/2024 |  |  | 10/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157562 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella cp device for mechanical circulatory support. the pump was unable to cross the common iliac, therefore, the procedure was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26871 | 1220648-2024-10954 | 19278468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19278468&pc=OZD | 02/21/2022 | 05/09/2024 |  | 05/09/2024 |  |  | 02/21/2022 | 05/09/2024 |  |  | 07/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2022114254 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hemolysis, it is unknown how the hemolysis was resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
26872 | 1220648-2024-10955 | 19278557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19278557&pc=OZD | 02/22/2022 | 05/21/2025 |  | 05/09/2024 |  |  | 02/22/2022 | 05/09/2024 |  |  | 11/30/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166563 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Great Vessel Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2022 |  |  |  |  | The complainant reported a patient with ventricular tachycardia (vt) ablation was implanted with an impella cp. it was reported that ischemic vt was induced throughout the case, the physician was able to successfully map and ablate. while the 8fr sheath was in the heart, the patient went into vt on his own requiring defibrillation, when the patient was shocked, the sheath punctured the heart. the physician utilized the impella to assist with ventricular tachycardia ablation to stabilize the patient. the patient was stabilized and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: vascular dissection: impella cp with smart assist for use during cardiogenic shock and high risk pci & impella 5.5 with smart assist for use during section: warnings & cautions: cautions as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: as noted in the impella ifu ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: use of impella with transcatheter aortic valves as noted in the impella ifu ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26873 | 1220648-2024-10956 | 19278587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19278587&pc=OZD | 02/25/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 02/25/2022 | 05/09/2024 |  |  | 10/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151816 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, bleeding and hematoma was noted. two units of blood was transfused, site was redressed multiple times, femstop applied, and heparin was withheld due to bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26874 | 1220648-2024-10957 | 19278607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19278607&pc=OZD | 04/11/2024 | 06/04/2024 |  | 05/09/2024 |  |  | 04/11/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025460544 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2024 |  |  |  |  | The user facility reported an 89-year-old female with high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that when attempting impella cp placement via single access, the physician punctured the peel away with the needle. due to this, the peel away sheath had to be removed and the repo sheath inserted. single access was not able to be used. percutaneous coronary intervention was performed through the radial. the impella cp was weaned and explanted successfully at conclusion of the case. no patient harm was reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported introducer damage has been completed. according to the clinical, before beginning impella cp support while the physician was attempting single access they punctured the peel-away sheath with a needle. the peel-away sheath was then removed and the repo sheath was inserted. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the mechanical introducer damage was use related.
26875 | 1220648-2024-10959 | 19278890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19278890&pc=OZD | 02/26/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 02/26/2022 | 05/09/2024 |  |  | 10/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151381 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia and hematoma. therefore, the impella cp was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26876 | 1220648-2024-10958 | 19278913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19278913&pc=OZD | 04/30/2024 | 06/25/2024 |  | 05/09/2024 |  |  | 04/30/2024 | 05/09/2024 |  |  | 11/30/2025 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024425823 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2024 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support and the patient had hemolysis. there was no known resolution to the hemolysis. additionally, heparin was withheld due to a possible gastrointestinal (gi) bleed. the bleed was never verified. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation into the hemolysis and the vascular damage - peripheral vessel issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the hemolysis and the vascular damage issue was not determined since product was not returned and limited clinical information was provided. d6b, g3.
26877 | 1220648-2024-10960 | 19279050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19279050&pc=OZD | 05/01/2024 | 06/03/2024 |  | 05/09/2024 |  |  | 05/01/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025455052 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2024 |  |  |  |  | The user facility reported a 75-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that oozing developed at the patient's impella cp access site during support. it was noted that the oozing occurred primarily after the patient was moved. three units of packed red blood cells were infused to maintain the baseline lab values in response to the oozing. support was maintained with the implanted pump throughout the event. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to patient movement since the oozing was noted after staff moved patient.
26878 | 1220648-2024-10961 | 19279142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19279142&pc=OZD | 02/26/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 02/26/2022 | 05/09/2024 |  |  | 10/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022155902 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, access site bleeding was noted. as a result, purged was changed from d5w + 50 (u/ml) heparin to d5w. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26879 | 1220648-2024-10962 | 19279173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19279173&pc=PYX | 05/01/2024 | 06/25/2024 |  | 05/09/2024 |  |  | 05/01/2024 | 05/09/2024 |  |  | 06/30/2025 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024371148 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2024 |  |  |  |  | The user facility reported an 82-year-old female of unknown ethnicity with acute right heart failure was implanted with an impella rp device for mechanical circulatory support. it was reported that the impella rp pump migrated during support with outlet positioning resting in the right ventricle. due to the shift in position, the decision was made to replace the impella rp with an impella rp smartasssist device via right femoral venous access. there was no patient harm as a result of the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information provided related to positioning issue.
26880 | 1220648-2024-10963 | 19279321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19279321&pc=OZD | 02/28/2022 | 04/16/2025 |  | 05/09/2024 |  |  | 02/28/2022 | 05/09/2024 |  |  | 11/30/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166534 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. the impella cp was unable to be placed and the case was not completed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26881 | 1220648-2024-10964 | 19279483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19279483&pc=OZD | 04/30/2024 | 06/24/2024 |  | 05/09/2024 |  |  | 05/02/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471086 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The user facility reported a 54-year-old male patient of unknown ethnicity post-operative coronary artery bypass graft for severe mitral valve and coronary artery disease was implanted with an impella 5.5 device for mechanical circulatory support. the patient was supported for 6 days and then weaned to explant. it was reported that after the explant and removal of the intubation, it was noted that the patient had a facial droop. the patient was reported as stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into the removal issues and the stroke/cva has been completed since the original report was submitted. the device was not returned for investigation. based on the clinical information provided, the most likely cause of the removal issues was patient condition related. as the clinical information related to the stroke/cva was not provided, the root cause could not be determined.
26882 | 1220648-2024-10965 | 19279637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19279637&pc=OZD | 02/28/2022 | 04/16/2025 |  | 05/09/2024 |  |  | 02/28/2022 | 05/09/2024 |  |  | 11/30/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169304 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2022 |  |  |  |  | The complainant reported a patient was being implanted with an impella cp device for mechanical circulatory support. during implantation the attempted impella cp implant but aborted related to patient anatomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26883 | 1220648-2024-10966 | 19280152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19280152&pc=OZD | 04/15/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 04/15/2022 | 05/09/2024 |  |  | 06/30/2022 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1485094 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 device for mechanical circulatory support. during placement, there was difficulty advancing the pump past the tortuous innominate artery into the aorta, and the wire moved out of the left ventricle (lv). the impella 5.0 and wire were both removed. another wire was advanced into the aortic valve, and a second wire was then placed in the lv apex for additional support. the impella 5.0 was advanced with some difficulty passing the ostia of the innominate artery, but using a push-pull technique, the physicians were able to advance the pump into the lv. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26884 | 1220648-2024-10967 | 19280708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19280708&pc=OZD | 05/04/2024 | 07/01/2024 |  | 05/09/2024 |  |  | 05/04/2024 | 05/09/2024 |  |  | 09/30/2025 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444602 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient lost pulse in the impella leg. . in response to the condition, the decision was made to remove the pump with surgical wound closure. the patient was stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: : potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ a.5 ethnicity was previously entered incorrectly. the patient's ethnicity is unknown. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
26885 | 1220648-2024-10968 | 19280794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19280794&pc=PYX | 05/05/2024 | 06/24/2024 |  | 05/09/2024 |  |  | 05/05/2024 | 05/09/2024 |  |  | 05/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2022103183 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/05/2024 |  |  |  |  | The user facility reported a 64-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that the impella rp was dropped to p2 to remove an old swan from the left groin, tip of swan near inlet. right after the swan was removed the motor current spiked, flows dropped, and the placement signal went high. clot ingested into the rp was suspected. md called to transfer the patient to another facility as there was no backup rp onsite. the patient was transferred to have a new rp inserted. the impella rp had a pump stop and was removed. a 16fr cook sheath was placed in its place. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings & cautions: to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers. | The investigation into the pump stop issue has been completed since the original report was filed. the pump was returned for investigation. upon visual inspection, biomaterial was observed in the inflow cage. data logs confirmed motor current spike indicating biomaterial ingestion. the root cause of the pump stop was due to biomaterial ingestion.
26886 | 1220648-2024-10969 | 19280907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19280907&pc=OZD | 05/05/2024 | 06/20/2024 |  | 05/09/2024 |  |  | 05/05/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025474750 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Deliver | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/05/2024 |  |  |  |  | The user facility reported a 62-year-old patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp pump was unable to advance through the patient's anatomy during attempted delivery. the implant was subsequently aborted. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the patient anatomy. d4-corrected model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
26887 | 1220648-2024-10970 | 19280927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19280927&pc=OZD | 05/05/2024 | 06/25/2024 |  | 05/09/2024 |  |  | 05/05/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1641550 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 05/05/2024 |  |  |  |  | The complainant reported their automated impella controller (aic) had a controller failure red alarm following an impella cp implant. the aic was replaced, and no patient harm has been reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation into the aic - controller failure (red alarm) has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the controller error was not reproduced while running an optical pump simulator. there was no issue with the console hardware related to the reported failure mode. the front panel optical connector was heavily contaminated as received. the root cause of the controller error was due to an aic software issue.
26888 | 1220648-2024-10971 | 19280939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19280939&pc=OZD | 05/06/2024 | 07/08/2024 |  | 05/09/2024 |  |  | 05/06/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025458547 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information provided that the patient developed hemolysis and subsequently the impella 5.5 was explanted. | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and while delivering support alarms occurred indicating "impella in aorta" and "impella position unknown." the physician believed the alarms were due to the patient's low native contractility. there was no patient harm reported. |  |  |  |  | The investigation for the reported optical issue and hemolysis has been completed. according to the clinical, on the fourteenth day of impella 5.5 support the patient's urine became a dark red color. the following day, multiple impella position unknown and impella in aorta alarms were recorded. the pump position was confirmed 4.9cm across the av. no attempts were made to resolve either issue. product evaluation confirmed the presence of biomaterial in the outlet cage. data log analysis confirmed suction prior to a motor current spike and no change in the 5s parameters. the cause of the optical signal issue was biomaterial. the cause of the hemolysis was biomaterial. including instructions for use for hemolysis section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿
26889 | 1220648-2024-10972 | 19280960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19280960&pc=OZD | 05/06/2024 | 06/28/2024 |  | 05/09/2024 |  |  | 05/06/2024 | 05/09/2024 |  |  | 11/30/2025 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024423535 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Disability | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information: it was reported the patient was in and out of ventricular tachycardia after being made a do not resuscitate. the patient expired. per the abiomed medical safety office, abiomed does not have enough information to associate use of device with outcome. | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support and suffered an embolic stroke. the patient was stated to be awake but not responding appropriately to verbal commands. action was not taken regarding the stroke and support with the implanted pump was maintained. it was noted there was no direct implication of the impella pump as the causative factor. |  |  |  |  | The investigation into the stroke/cerebral vascular accident has been completed since the original report was submitted. the device nor data logs were returned for investigation. as clinical information related to the stroke/cardiovascular accident was not provided, the root cause could not be determined. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock-section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
26890 | 1220648-2024-10973 | 19281054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19281054&pc=OZD | 05/07/2024 | 06/24/2024 |  | 05/09/2024 |  |  | 05/07/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442933 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/07/2024 |  |  |  |  | The complainant reported a 36-year old male patient was electively implanted with an impella cp for mechanical circulatory support. while on support, the patient developed oozing at the access site, and a high purge pressure alarm was generated. medical staff managed the bleeding with additional sutures. medical staff administered two units of fresh frozen plasma, four units of platelets, seven units of packed red blood cells, and two units of cryoprecipitate. the patient remains on support. | [5 empty] | The investigation into the high purge pressure and the access site bleeding ¿ major issues has been completed since the original report was filed. the pump was returned for investigation. product evaluation did not confirm the presence of any abnormalities. additionally, the introducer was not returned for a visual inspection. data log analysis confirmed the presence of purge related alarms and trends that resolve after tpa was added. the cause of the high purge pressure was biomaterial. the cause of the access site bleeding was patient condition related. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
26891 | 1220648-2024-10974 | 19282698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19282698&pc=OZD | 02/13/2022 | 04/17/2025 |  | 05/09/2024 |  |  | 02/13/2022 | 05/09/2024 |  |  | 11/30/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166082 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the peel away sheath cracked and was removed while partially still in the patient's body. however, the device was placed successfully, and the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
26892 | 1220648-2024-10975 | 19282709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19282709&pc=OZD | 02/14/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 02/14/2022 | 05/09/2024 |  |  | 10/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156979 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/14/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved with femstop and a mattress stitch. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26893 | 1220648-2024-10976 | 19282746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19282746&pc=OZD | 02/15/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 02/15/2022 | 05/09/2024 |  |  | 10/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152098 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | Vascular Dissection | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/15/2022 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient then exhibited oozing from the insertion site. the bleeding was initially managed with manual pressure, but the patient eventually required transfusion of one unit of packed red blood cells. the patient was also treated with antibiotics due to the result of blood cultures and increased white blood cell counts. the patient's right lower extremity was cold to touch and lost pulses, and the impella was scheduled for removal. during induction for ct the patient had experienced vt arrest. it was found that the patient had aortic dissection from infrarenal to right common femoral requiring endovascular aneurysm repair. it is reported that the patient survived. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ limb ischemia: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ infection/sepsis: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ vascular dissection: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26894 | 1220648-2024-10977 | 19282805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19282805&pc=PYX | 05/17/2016 | 06/24/2024 |  | 05/09/2024 |  |  | 05/02/2024 | 05/09/2024 |  |  | 11/30/2017 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1226178 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult or Delayed Positioning | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | On 2024-05-02 abiomed complaints team forwarded publication entitled: "imaging modalities for correct positioning of percutaneous right ventricular assist device after left ventricular assist device implantation" which outlined an impella rp support - along with heartware hvad - in which there was reported "the rv dilation was attributed to moderate pr (interval), as a result of leaflet tenting by the impella rp. the device was advanced by 2 cm, assuring parallel alignment with the main pulmonary artery, which resulted in subsequent improvement in the left ventricle filling, return of the interventricular septum position to midline, and noticeable decrease in pr (interval)." | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined.
26895 | 1220648-2024-10978 | 19282816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19282816&pc=PYX | 05/03/2024 | 07/10/2024 |  | 05/09/2024 |  |  | 05/03/2024 | 05/09/2024 |  |  |  |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2025478700 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | IMPELLA CP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2024 |  |  |  |  | It was further reported that care was withdrawn from the patient and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. | The user facility reported a 65-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that the patient had bleeding in their endotracheal tube. an isolated spot of bleeding was found which was resolved after flushing and suction of the area. the physician believes the bleeding was caused by difficulty placing the guidewire while implanting the impella. the patient remains on support and is reported as stable. |  |  |  |  | The investigation for the reported insertion/removal difficulties and the great vessel damage has been completed. there was a patient injury suspected by physician to be caused by the guidewire when placing impella. no product was returned. the root cause of the delivery issue was not determined as no product was returned and limited clinical information was provided. the root cause of the vascular damage - peripheral vessel was not determined as no product was returned and limited clinical information was provided. added- b5 mso review d3-corrected mfr city. f6/f8 should have been left blank in the initial report. g1-mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings & cautions: to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.
26896 | 1220648-2024-10979 | 19282859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19282859&pc=OZD | 02/16/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 02/16/2022 | 05/09/2024 |  |  | 09/30/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022147388 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, there was a bleeding issue due to accidentally moving the sheath. therefore, the impella was removed early. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26897 | 1220648-2024-10981 | 19282895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19282895&pc=OZD | 03/22/2022 | 05/09/2024 |  | 05/09/2024 |  |  | 03/22/2022 | 05/09/2024 |  |  | 12/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176447 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/22/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a run of ventricular tachycardia, which was resolved with cardioversion and cardiopulmonary resuscitation. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26898 | 1220648-2024-10982 | 19282926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19282926&pc=OZD | 04/05/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 04/05/2022 | 05/09/2024 |  |  | 01/01/2024 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187087 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, an attempted axillary access was made, but during the deployment of the perclose device, a small branch was inadvertently perforated. the patient remained stable, and the vascular team exposed the area to address and repair the perforation. the procedure transitioned from percutaneous axillary to a cutdown approach during the axillary dissection. the patient remained on impella support through completion of the procedure. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26899 | 1220648-2024-10983 | 19282946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19282946&pc=OZD | 09/09/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 09/09/2022 | 05/09/2024 |  |  | 05/31/2024 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231014 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2022 |  |  |  |  | It was reported that a 59 year old male patient was implanted with an impella cp for mechanical circulatory support due to acute myocardial infarction and cardiogenic shock (ami/csg). while on support, the physician reported that the patient's bilateral groins were both bleeding. manual pressure was applied and epi-lidocaine was utilized at the site. the impella cp was removed several hours later. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26900 | 1220648-2024-10984 | 19282978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19282978&pc=OZD | 03/01/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 03/01/2022 | 05/09/2024 |  |  | 10/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149275 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/01/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella was removed due to hemolysis and the patient was started on continuous renal replacement therapy. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26901 | 1220648-2024-10985 | 19283000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19283000&pc=OZD | 09/27/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 09/27/2022 | 05/09/2024 |  |  | 08/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127401 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp for mechanical circulatory support. the patient had an intra-aortic balloon pump (iabp) that was replaced with the impella cp. while on support, access site bleeding and a hematoma was observed. a transfusion was given to the patient to resolve the issue. the patient was successfully weaned and the device explanted. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26902 | 1220648-2024-10986 | 19283002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19283002&pc=OZD | 03/01/2022 | 05/30/2025 |  | 05/09/2024 |  |  | 03/01/2022 | 05/09/2024 |  |  | 10/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157575 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/01/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during explant of the device, a thrombus and potential biomaterial were observed. the patient¿s condition was reportedly stable. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) medications or treatments that the patient was receiving at the time of thrombus formation surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
26903 | 1220648-2024-10987 | 19283008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19283008&pc=OZD | 10/14/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 10/14/2022 | 05/09/2024 |  |  | 06/30/2024 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238729 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2022 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced minimal oozing at the groin. the site bleed was addressed by utilizing angle matching techniques and pressure dressing. the impella cp was explanted. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26904 | 1220648-2024-10988 | 19283012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19283012&pc=OZD | 03/02/2022 | 05/20/2025 |  | 05/09/2024 |  |  | 03/02/2022 | 05/09/2024 |  |  | 10/31/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022147583 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 03/02/2022 |  |  |  |  | The user facility reported 67-year-old patient needing mechanical circulatory support. it was reported that the patient¿s urine had a pink tinge and hemolysis was observed as evidenced by hematuria and an uptrend in lactate dehydrogenase (ldh). it was stated the device was repositioned; however, repositioning did not manage this complication. the patient was not a candidate for orthotopic heart transplantation (oht) or any advanced therapies. the patient was weaned, and the device was explanted so the surviving patient could be medically managed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26905 | 1220648-2024-10989 | 19283044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19283044&pc=OZD | 03/02/2022 | 05/16/2025 |  | 05/09/2024 |  |  | 03/02/2022 | 05/09/2024 |  |  | 11/30/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167186 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced hemolysis that was resolved by decreasing the pump from p9 to p5. the patient also experienced ventricular tachycardia that self-resolved. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
26906 | 1220648-2024-10990 | 19283079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19283079&pc=OZD | 03/08/2022 | 05/09/2024 |  | 05/09/2024 |  |  | 03/08/2022 | 05/09/2024 |  |  | 02/28/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022043571 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/08/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support no distal flow was noted with the peel-away in, the decision was made to peel the peel-away and this resolved the limb ischemia. the physician started to peel away the sheath while it was still in the artery, and this caused a brief period of significant bleeding at the access site. this issue was resolved by dressing this site. the patient also experienced an episode of non-sustained ventricular tachycardia and hemolysis while on support. the physician decided to remove the device and it was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26907 | 1220648-2024-10991 | 19283086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19283086&pc=OZD | 03/08/2022 | 05/30/2025 |  | 05/09/2024 |  |  | 03/08/2022 | 05/09/2024 |  |  | 11/30/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023159814 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella support, labs showed that the patient was hemolyzing. an echocardiogram showed the impella was slightly shallow, and the physician attempted to reposition the device, but this was unsuccessful as the device was angled toward the mitral valve. the decision was made to escalate the patient to an impella 5.5. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined.
26908 | 1220648-2024-10992 | 19283104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19283104&pc=OZD | 03/10/2022 | 05/09/2024 |  | 05/09/2024 |  |  | 03/10/2022 | 05/09/2024 |  |  | 11/30/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161056 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/10/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia. it was reported the patient lost distal pulses. the impella was repositioned resolving the lost pulse issue. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event is as follows: impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26909 | 1220648-2024-10993 | 19283112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19283112&pc=OZD | 03/12/2022 | 05/09/2024 |  | 05/09/2024 |  |  | 03/12/2022 | 05/09/2024 |  |  | 11/30/2023 |  | 05/09/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169242 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/10/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced arrhythmia and was returned to the cardiac catheterization lab for swan placement. the next day it was reported the patient was cardioverted again and continued to have ectopy. it was reported the swan was removed and the patient was stable. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26910 | 1220648-2024-10874 | 19267152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19267152&pc=OZD | 04/22/2022 | 05/08/2024 |  | 05/08/2024 |  |  | 04/22/2022 | 05/08/2024 |  |  | 11/30/2023 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168114 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/22/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was being placed on an impella cp device for mechanical circulatory support. it was reported that the patient¿s vasculature (vessels were too small) made it impossible to implant the impella device. the procedure was aborted, | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
26911 | 1220648-2024-10875 | 19267363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19267363&pc=PYX | 04/08/2024 | 06/05/2024 |  | 05/08/2024 |  |  | 04/08/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025478612 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that upon changing out the peel-away sheath, the impella rp flex motor current spiked and pump flows dropped. the physician was notified of possible clot ingestion, but no action was initially taken in response to the issue. when the flows dropped yet again at p9 support levels, the decision was made to replace the pump. a fibrous clot was subsequently found at the inlet upon explant. there was no patient harm resulting from the issue. | [5 empty] | The investigation for the low pump flow has been completed. the product was not returned for review. the data logs confirmed the failure mode and clinical narrative. the logs showed after the pump started for about 10 minutes, motor current spiked followed low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial ingestion. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. h.6 code 2954 was reported incorrectly on the previously submitted report. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
26912 | 1220648-2024-10758 | 19267536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19267536&pc=OZD | 04/28/2024 | 06/19/2024 |  | 05/08/2024 |  |  | 04/28/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025463805 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Difficult to Insert | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/28/2024 |  |  |  |  | The complainant reported a male patient of unknown age with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. delivery failed due to the patient¿s vascular anatomy ¿ there was trauma at the anastomosis that had to be corrected. the anastomosis was bleeding and the pump had become kinked. the implant procedure was aborted. | [5 empty] | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the clinical states that there were delivery issues due to patient anatomy requiring numerous attempts to deliver the pump likely leading to the vascular damage and bleeding requiring blood products and vascular repair surgery. the root cause of the delivery issue was most likely patient condition related. the root cause of the vascular damage peripheral vessel issue was most likely use related. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
26913 | 1220648-2024-10876 | 19267599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19267599&pc=OZD | 04/30/2024 | 09/22/2025 |  | 05/08/2024 |  |  | 04/30/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025455090 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Difficult to Insert | Patient-Device Incompatibility |  |  |  |  | Arrhythmia | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/30/2024 |  |  |  |  | The complainant reported a 66-year old male patient with post-cardiotomy cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that initial placement of the 5.5 was difficult, however, the 5.5 was flowing. the pump was thought to be too deep and was subsequently pulled out of the left ventricle (lv). the 5.5 was removed, and lv was again accessed using another pigtail and a .035 wire. the 5.5 was placed again with some difficulty. initially there was a placement signal but after the wire was removed, the placement signal failed. medical staff discussed the issue and decided to leave the 5.5 in place because the patient's mean arterial pressure was low. the medical staff¿s physician did not want to remove or replace the impella due to the complexity of the implant. after approximately 15 minutes, the patient lost pulsatility. the physician thought it was due to the thickened septum from years of heart failure leaving inadequate space, which in turn was causing the patient to experience arrhythmia. a trans-esophageal echocardiogram did reveal the lv to be empty. the physician then decided to remove the 5.5 because it was causing arrhythmia. the patient remains stable. | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. corrected model number. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump was not returned. data logs were reviewed, and placement signal not reliable observed immediately after pump insertion for second time. the optical 5s parameters signal-to-noise ratio (snr) and contrast were found to be decreasing while the gain and lamp increased. no other abnormalities were found. the cause of the optical signal failure was most likely an air gap between the automated impella controller and the patient cable plug per clinical and log review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. corrections that were made from the initial manufacturer device report number 1220648-2024-10876. d10 concomitant medical products and therapy dates updated. f6 and f8 should not have had entries in the initial report. g1 reporting contact email updated.
26914 | 1220648-2024-10877 | 19267636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19267636&pc=OZD | 04/23/2024 | 06/24/2024 |  | 05/08/2024 |  |  | 04/30/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461141 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Twisted/Bent | [6 empty] | Embolism/Embolus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | The user facility reported a 62-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that there were issues with the kinking of the guidewire upon implanting the pump. two 0.018" wires were used to insert the pump. however, both kinked while the pump was going across the aortic valve. the wires were aborted, and it was restarted with an al1, and .035" j glide used to access the left ventricle, and a 0.018" platinum plus was used to insert the impella over. while on support, the patient required an embolectomy of the right brachial artery. the impella was removed and replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the reported thrombus and the product damage has been completed. according to the clinical details, while attempting to insert the guidewire across the aortic valve, the wire kinked twice. on the third attempt optimal position was achieved. seven days later the patient required an embolectomy and the pump was removed. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of product damage (kink issue) was related to use related, as it took 3 tries to successfully deliver the device to the optimal position. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
26915 | 1220648-2024-10878 | 19267653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19267653&pc=OZD | 05/01/2024 | 06/06/2024 |  | 05/08/2024 |  |  | 05/01/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438860 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Patient-Device Incompatibility | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2024 |  |  |  |  | The complainant reported a 58-year old female patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported that the patient¿s femoral arteries were too small to facilitate impella cp placement, so the physician inserted the cp axillary. the upper extremity, distal to the cp access site, became ischemic, so an external bypass was placed. the limb perfused well with the external bypass, and the patient remains stable. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. d4-updated model number. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
26916 | 1220648-2024-10879 | 19267770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19267770&pc=OZD | 05/01/2024 | 06/19/2024 |  | 05/08/2024 |  |  | 05/01/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025453569 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Migration or Expulsion of Device | Activation, Positioning or Separation Problem | Positioning Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2024 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella cp pump was inadvertently pulled back into the aorta while the patient was being moved in their bed. hemodynamics were maintained utilizing vasopressors and inotropes, and the patient was subsequently escalated to impella 5.5 support. there was no patient harm. | [5 empty] | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. based on clinical description, the root cause of positioning issue was most likely patient movement. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
26917 | 1220648-2024-10880 | 19267776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19267776&pc=OZD | 04/23/2024 | 05/23/2024 |  | 05/08/2024 |  |  | 04/23/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED,INC |  |  | US |  | 1125337 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | The user facility biomedical department reported that the loaner console had a battery failure observed upon setup. the automated impella controller (aic) was removed from service per instructions for use a backup will be used for any patient use. no patient harm or involvement was reported. | [5 empty] | The investigation for the reported console (aic) battery failure issue has been completed. the console was returned for evaluation. upon initial boot up with the ac power, the console displayed the battery failure alarm, and the battery 1 lcd screen was partially blank. the batttest result shows a cell imbalance in cell 4 of battery 1. upon removing the original defective battery set and connecting the known-good batteries, aic was able to charge the batteries without any failure. data logs were reviewed finding the console was booted up and shut down nine times, and every time, there was a persistent ¿battery failure" alarm. this confirms the reported failure mode. the root cause of the battery failure alarm was cell imbalance in cell 4 of battery 1. added- d9 device return date d4-updated common device name | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26918 | 1220648-2024-10882 | 19267818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19267818&pc=OZD | 05/01/2024 | 06/26/2024 |  | 05/08/2024 |  |  | 05/01/2024 | 05/08/2024 |  |  | 03/31/2025 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024327106 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the patient developed ischemia in the right lower leg in response to impella cp placement. a thrombus was confirmed in the leg and pump removal was reported to alleviate the issue. support for the patient continued with impella 5.5 support. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system - section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation for the reported limb ischemia and thrombus has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and thrombus was not determined. as noted in the impella instructions for use: impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." d.4 revised model number from the initial submission in accordance with updated procedures. d.4 unique identifier (udi) # was revised as was previously entered incorrectly. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use - a portion was missing from the previously submitted report and was added to this report.
26919 | 1220648-2024-10881 | 19267824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19267824&pc=OZD | 04/25/2024 | 06/04/2024 |  | 05/08/2024 |  |  | 04/25/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED,INC |  |  | US |  | 2025458119 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The user facility reported a 69-year-old female patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that patient was in profound metabolic shock and the limb accessed suffered limb ischemia and so the pump was explanted. extracorporeal membrane oxygenation (ecmo) was placed instead, and patient survived. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
26920 | 1220648-2024-10883 | 19267846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19267846&pc=OZD | 04/27/2024 | 06/04/2024 |  | 05/08/2024 |  |  | 04/27/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025492149 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | VA-ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2024 |  |  |  |  | The user facility reported a 65-year-old female patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed ischemia in their lower left extremity during impella cp support. in response to the condition, the pump was explanted and a new impella cp pump was implanted via right axillary artery. the patient was reported as stable. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use referenced in the initial report is for the cp with smartassist. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
26921 | 1220648-2024-10884 | 19267911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19267911&pc=OZD | 04/30/2024 | 06/13/2024 |  | 05/08/2024 |  |  | 04/27/2024 | 05/08/2024 |  |  | 11/30/2025 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025471305 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2024 |  |  |  |  | The user facility reported a 48-year-old male patient was implanted with an impella 5.5 as a bridge to transplant. it was reported the patient experienced frequent ectopy/ventricular tachycardia (vt) during impella 5.5 support. to treat the condition, the pump was repositioned, and the frequency notably decreased. it was additionally stated that the patient experienced some oozing from their access site following a dressing change. three sutures were placed at the site to achieve hemostasis. support with the implanted pump was maintained despite the events. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be use issue because the issue was resolved by adding additional sutures at the site. as the necessary clinical information was not provided, the root cause of the vt/vf was not determined.
26922 | 1220648-2024-10885 | 19267980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19267980&pc=OZD | 05/01/2024 | 06/11/2024 |  | 05/08/2024 |  |  | 04/15/2024 | 05/08/2024 |  |  | 08/31/2025 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025443681 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2024 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and there was a thrombus noted after the impella was explanted. a small thrombus was noted post closure, and the decision was made to leave the small nonflowing thrombus without intervention. however, the thrombus showered down the right leg causing full occlusion. a five-hour thrombectomy was performed and low was restored to distal extremities. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus and the limb ischemia was not determined. the root cause of the thrombocytopenia cannot be determined due to insufficient clinical details. in accordance with updated procedures, sections d6 have been revised from the initial submission.
26923 | 1220648-2024-10886 | 19267982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19267982&pc=OZD | 04/28/2024 | 06/07/2024 |  | 05/08/2024 |  |  | 04/28/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025473936 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/28/2024 |  |  |  |  | The user facility reported a male patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella 5.5 pump was unable to advance beyond the innominate-aortic junction during attempted delivery. after multiple failed attempts, the catheter began to kink and the decision was made to attempt placement with another impella 5.5 pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue is determined to be patient condition because of the pump was unable to advance and the catheter was kinked due to vascular anatomy of the patient.
26924 | 1220648-2024-10887 | 19268040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19268040&pc=OZD | 04/29/2024 | 06/21/2024 |  | 05/08/2024 |  |  | 04/29/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025455125 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  |  |  | Malfunction | No | Yes | Crack | Restricted Flow rate | Fluid/Blood Leak |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2024 |  |  |  |  | The complainant reported a 58-year old female patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. medical staff noticed a crack on the purge sidearm of the cp, and the purge pressure was low. the complainant did not share how the issue was resolved, and no patient harm has been reported. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist for use during cardiogenic shock and high risk cpi. section: cleaning . ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The investigation for the purge leak has been completed. the impella was not returned for evaluation. as per clinical, crack in purge sidearm observed with low purge pressure alarm. the cause of the purge leak issue was a leak from the sidearm but the exact location of the leak was unknown. d1-common device name. d4-corrected model number. d2-corrected common device name.
26925 | 1220648-2024-10888 | 19268064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19268064&pc=PYX | 04/29/2024 | 06/21/2024 |  | 05/08/2024 |  |  | 04/29/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025443694 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 and an impella rp for mechanical circulatory support. during support, the patient developed red urine. the p-levels were adjusted. a ct scan showed the impella rp outlet was shallow and close to the pulmonic valve. the physician adjusted to adequate position. the patient's urine lightened in color but was still red. the decision was made to explant impella rp flex and change patient to status one for kidney and heart transplant. the care team attributes hemolysis to impella rp flex. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ warnings and cautions: ¿patients with tricuspid or pulmonary valve stenosis or insufficiency, and patients with prosthetic tricuspid or pulmonary valves, may be compromised by the use of the impella rp flex with smartassist system catheter.¿ | The investigation into the hemolysis issue has been completed since the original report was filed. the pump was returned for investigation. visual inspection found biomaterial inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of the hemolysis was biomaterial ingestion.
26926 | 1220648-2024-10889 | 19268235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19268235&pc=OZD | 09/07/2022 | 03/26/2025 |  | 05/08/2024 |  |  | 09/07/2022 | 05/08/2024 |  |  | 05/31/2024 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231619 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump dislodged form its location. impella cp pump was explanted. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26927 | 1220648-2024-10890 | 19268240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19268240&pc=OZD | 03/02/2022 | 03/26/2025 |  | 05/08/2024 |  |  | 03/02/2022 | 05/08/2024 |  |  | 11/30/2023 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168487 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. the physician attempted to advance the impella cp but the patient had ectopy and suction as pigtail was behind papillary structure. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26928 | 1220648-2024-10891 | 19268255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19268255&pc=PYX | 05/01/2024 | 06/24/2024 |  | 05/08/2024 |  |  | 05/01/2024 | 05/08/2024 |  |  | 10/31/2025 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024423948 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2024 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that at the start of impella rp procedure the 0.027 guidewire to gain access to the pulmonary artery kinked. it was a difficult insertion of the rp as there were some anatomical challenges accompanied by a large right ventricle. another 0.027 guidewire was utilized to successfully advance the impella. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (guidewire damage - peripheral vessels) has been completed since the original report was submitted. the device was not returned for investigation. the most likely cause of the guidewire damage was use related.
26929 | 1220648-2024-10892 | 19268332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19268332&pc=OZD | 04/10/2024 | 05/22/2024 |  | 05/08/2024 |  |  | 04/10/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025457983 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | The complainant reported a patient in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the pump was successful for the duration of the support however the patient had an access site ooze. the ooze was not of a severity that needed any intervention/blood/vascular repair. the pump support continued for a total of 190 hours until the family made choice to withdraw care. the team discontinued all systemic heparin due to the access site bleeding. the purge was not utilizing heparin, but rather sodium bicarbonate. the patient's family withdrew care . impella explanted and the patient expired. there is not enough information to associate use of the device with patient outcome. pump was successfully implanted and explanted. | [5 empty] | The investigation into the patient's access site bleeding has been completed. no impella products were returned for evaluation. based on the clinical review the root cause of the access site bleeding was most likely due to anticoagulation management since the heparin was discontinued due to the bleeding and patient condition. instructions for use for this event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ potential adverse events :¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ h.3 revised as previously entered incorrectly. the device was not evaluated by the manufacturer and evaluation summary was not attached. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
26930 | 1220648-2024-10893 | 19268440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19268440&pc=OZD | 05/01/2024 | 06/28/2024 |  | 05/08/2024 |  |  | 05/02/2024 | 05/08/2024 |  |  | 08/31/2025 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024391529 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in japan reported the patient was on impella cp smart assist and after removal a thrombus was noted. the patient had limb ischemia and a blood clot was found in the iliac artery. the blood clot was removed using a fogarty catheter. it was determined that the blood clot was likely related to the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation into the thrombus and the limb ischemia ¿ reversible issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary information was not provided, the root cause of the thrombus and the limb ischemia was not determined.
26931 | 1220648-2024-10894 | 19268540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19268540&pc=OZD | 04/29/2024 | 06/10/2024 |  | 05/08/2024 |  |  | 04/30/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489124 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | The user facility reported a 73-year-old female patient of unknown race and ethnicity with cardiomyopathy was to be implanted with an impella cp for mechanical circulatory support. it was reported that the patient¿s vessel tortuosity would not allow the sheath placement and the case was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition.
26932 | 1220648-2024-10895 | 19268544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19268544&pc=OZD | 05/02/2024 | 07/03/2024 |  | 05/08/2024 |  |  | 05/02/2024 | 05/08/2024 |  |  | 07/31/2025 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025441432 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. initial heparin given was 12k, act was 196 at that time, but additional 4k heparin was administered. it was noticed the left ventricle waveform was way higher than the aortic signal. the impella cp flows became saturated, and motor current observed to be in the 1100s. there was concern for clot ingestion, so impella was discontinued and removed. new impella cp successfully implanted. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported saturated flow has been completed. according to the clinical, shortly after beginning impella cp support pump flows became saturated and motor current rose to the 1100s. due to concerns of clot ingestion, the decision was made to remove the pump and replace it with another impella cp device. product evaluation confirmed the presence of biomaterial wrapped around the impeller along the purge gap. data log analysis also confirmed an increasing trend in the motor current shortly after beginning support. during this time the pump was also at p9 with a flow of 4 l/min. the cause of the saturated flow was biomaterial. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk cpi in section: using the automated impella controller with the impella catheter states: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ f6/f8 should have been left blank.
26933 | 1220648-2024-10896 | 19268546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19268546&pc=OZD | 03/04/2022 | 05/23/2025 |  | 05/08/2024 |  |  | 03/04/2022 | 05/08/2024 |  |  | 12/31/2023 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174250 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a femoral exploration was performed for continuous bleeding around the sheath. repo sheath was removed from vessel and purse string placed to secure catheter in artery. the physician placed two sutures and a pressure dressing to the site. the patient also developed hemolysis that was diagnosed visually by seeing tinged foley/urine. the staff discussed the removal of the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26934 | 1220648-2024-10898 | 19268730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19268730&pc=OZD | 03/05/2022 | 03/26/2025 |  | 05/08/2024 |  |  | 03/05/2022 | 05/08/2024 |  |  | 12/31/2023 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171424 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the device was removed due to the physician's concern for bleeding. the patient went pulseless electrical activity. advanced cardiac life support protocol was initiated, and the patient was resuscitated. the vascular surgeon was called in to assess impella cp access site and right femoral artery where two preclose were deployed. homeostasis was achieved via surgical closure of arteriotomy. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26935 | 1220648-2024-10899 | 19268829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19268829&pc=OZD | 05/03/2024 | 06/25/2024 |  | 05/08/2024 |  |  | 05/03/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492962 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2024 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the pump exhibited low flow. the device was removed, cleaned, re-implanted and still exhibited low flow. the pump was replaced with another impella cp to resolve the issue. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned for investigation. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue under bench test. the root cause of the low flow was most likely use related as no issue was found on the pump and data. d9, g3. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
26936 | 1220648-2024-10900 | 19268929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19268929&pc=OZD | 04/08/2024 | 06/06/2024 |  | 05/08/2024 |  |  | 04/08/2024 | 05/08/2024 |  |  | 08/31/2025 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398288 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2024 |  |  |  |  | The user facility reported a 70-year-old male patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the medical doctor (md) crossed the aortic valve with a 0.035 pigtail without issues and then exchanged the wire for a 0.018 guidewire. the impella cp was then backloaded via an easy guide lumen and was able to advance, albeit shallow. the md was unable to position the inlet in the mid-ventricular space and reported it felt stiff. the md then removed everything and retried with a bs v18 wire. the impella cp was then implanted, and a catheter lab nurse noted a bend in the 0.018 wire. no patient harm was reported. the patient eventually expired. there is not enough information to associate the use of the device with the patient outcome. the pump was successfully implanted. | [5 empty] | The investigation into the product damage (guidewire damage - peripheral vessels) issue has been completed since the original report was filed. the guidewire was returned for investigation. there were two kinks observed on the guidewire. as per the clinical information provided, the patient was obese and had calcified vessels making it difficult for delivery and hence excessive force was applied during delivery. per the clinical and product evaluation, the root cause of the delivery issue was patient condition related to diseased vessels. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26937 | 1220648-2024-10901 | 19268956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19268956&pc=OZD | 03/05/2022 | 03/26/2025 |  | 05/08/2024 |  |  | 03/05/2022 | 05/08/2024 |  |  | 07/31/2023 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022111214 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2022 |  |  |  |  | The user facility reported a patient on impella cp for hemodynamic support. while on support bleeding and hematoma was noted. a blood transfusion was required. no additional information was provided. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26938 | 1220648-2024-10902 | 19269101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19269101&pc=OZD | 03/07/2022 | 03/26/2025 |  | 05/08/2024 |  |  | 03/07/2022 | 05/08/2024 |  |  | 10/31/2023 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022153489 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was noted that the impella cp was measured 4.5 centimeters in the left ventricle and appeared to be caught on the papillary muscle. also the patient¿s urine was tea colored, and the partial thromboplastin time was therapeutic. the patient was weaned, and the device was explanted as a balloon tamponade and preclose were utilized to achieve hemostasis. post closure, a thrombus was identified at the impella access site. aspiration was attempted without success. an ekosonic endovascular system catheter was placed with tissue plasminogen activator infusion to remedy the issue. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26939 | 1220648-2024-10903 | 19269406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19269406&pc=OZD | 03/08/2022 | 03/26/2025 |  | 05/08/2024 |  |  | 03/08/2022 | 05/08/2024 |  |  | 09/30/2023 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142997 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support due to st elevation myocardial infarction. while on support, the pump would not increase from p0. multiple attempts were made to restart the pump again but were not successful and the motor current high alarm persisted. the physician was notified that a pump stoppage most likely being caused by a clot ingestion, causing the pump to unable to be restarted. it was recommended that the physician pull the impella back across the valve, however, the physician was uncomfortable repositioning. the patient was transferred with a non-working pump and the pump was surgically explanted once the patient arrived at the new facility. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26940 | 1220648-2024-10904 | 19269620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19269620&pc=OZD | 03/17/2022 | 05/08/2024 |  | 05/08/2024 |  |  | 03/17/2022 | 05/08/2024 |  |  | 11/30/2023 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023169702 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/17/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support the patient experienced reddish-amber urine output and a diagnosis of hemolysis, resolution unknown. additionally, the patient experienced oozing at the access site, the nurse changed the dressing and reported that this resolved the oozing. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
26941 | 1220648-2024-10905 | 19270248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19270248&pc=OZD | 05/03/2024 | 07/11/2024 |  | 05/08/2024 |  |  | 05/03/2024 | 05/08/2024 |  |  | 11/30/2025 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427749 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Sense | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/03/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. after the impella was inserted and turned on, the placement signal not reliable alarm occurred. despite troubleshooting, there was no placement signal, and the decision was made by the physician to remove it and replace it with a new 5.5. the new impella was inserted successfully and had adequate placement signal and position. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue has been completed. the device was not returned for investigation. data logs were reviewed and 5s parameters of the pump during run are consistent with a break of the optical fiber due to torque. the data logs showed that the break occurred before the pump was turned on. clinical details noted that insertion of the pump was per instructions for use guidelines. the root cause of the optical signal issue is due to a damaged fiber line as a result of a torqued fiber. b.7 revised as this information was inadvertently omitted from manufacturer device report number 1220648-2024-10905. d.6b explant date was revised as it was entered incorrectly on manufacturer device report number 1220648-2024-10905. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-10905 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-10905 was submitted. h.6 codes 2199 and 1559 were reported incorrectly on manufacturer device report 1220648-2024-10905. new codes have been added to medical device problem codes and health effect - impact codes.
26942 | 1220648-2024-10906 | 19270286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19270286&pc=OZD | 05/04/2024 | 07/16/2024 |  | 05/08/2024 |  |  | 05/04/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025467930 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2024 |  |  |  |  | The user facility reported a 52-year-old female patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that high purge pressure was encountered during impella 5.5 support. tissue plasminogen activator was added to the purge line and the high purge pressure / purge system blocked alarm resolved. there was no patient harm resulting from the noted issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the reported high purge pressure has been completed. no product was returned for evaluation. data logs were reviewed and the left logs showed a slow buildup of purge pressure over a period of 32 hours before "high purge pressure" alarm was triggered. high purge pressure was resolved over a period of 19 hours. clinical details noted that no kinks were observed in purge line. tissue plasminogen activator (tpa) was added to the purge and alarms stopped 3 hours after giving tpa. the root cause of the high purge pressure is most likely due to a gradual buildup of biomaterial that was resolved with tpa. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist in section: using the automated impella controller with the impella catheter. added-d6b. f6/f8-should have been left blank in the initial report.
26943 | 1220648-2024-10907 | 19270396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19270396&pc=OZD | 05/05/2024 | 07/17/2024 |  | 05/08/2024 |  |  | 05/05/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025441737 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | Insufficient Flow or Under Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/05/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the device exhibited low pump flow and increased purge pressure. the device was replaced to resolve the issue, and there was no patient harm reported. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure (hpp) has been completed. data logs show that 18 hours into support the purge pressure (pp) begins to rise. this starts when the p-level is turned up and there are suction alarms. the pp continues to rise with a few periods of stabilization, until the last hour of the case where it increases from 800 to ~1000 mmhg over 1 hour. no purge alarms are triggered. the motor current (mc) and flow are slightly trending upwards during this time with slightly saturated flows. no purge line blood ingress was found on the returned product. no biomaterial was found on the returned product and hpp was not reproduced in a bucket in the lab. the root cause of the high purge pressure was likely biomaterial. added- d9 device return date, h6 impact code 4629 d4-corrected model number f6/f8 should have been left blank in the initial report.
26944 | 1220648-2024-10908 | 19270768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19270768&pc=OZD | 04/10/2024 | 06/03/2024 |  | 05/08/2024 |  |  | 04/10/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025447025 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Sepsis | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | The user facility reported a 66 year-old male patient with acute myocardial infarction and cardiogenic shock was implanted with an impella device for mechanical circulatory support. while on support, the patient developed right limb ischemia. consultation with vascular surgery determined that some degree of amputation was inevitable due to the right leg becoming septic as well. the patient was placed on ecpella, and the right leg's access site was placed on ice. due to the patient being on ecpella, vascular surgery was no longer able to intervene. the patient eventually expired. there is not enough information to exclude the use of the device with the patient outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into the limb ischemia ¿ irreversible, infection/sepsis, and the renal failure issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. because this evidence has not been provided, the root cause of the infection cannot be determined. the root cause of the renal failure cannot be determined since the product was not returned and insufficient clinical details were provided. impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections g3 and h10 have been revised from the initial submission.
26945 | 1220648-2024-10909 | 19270854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19270854&pc=OZD | 04/30/2024 | 06/07/2024 |  | 05/08/2024 |  |  | 04/30/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1094708 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an automated impella controller was noted to have had a broken purge cassette holder, and the biomedical engineer noted the door was unable to be opened as a result of the malfunction. there was no patient involvement. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. added- h6 med dev prob code 1069, h6 component code 838 d2-updated common device name d4-updated model number. | The automated impella controller (aic) has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
26946 | 1220648-2024-10910 | 19271049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19271049&pc=OZD | 04/30/2024 | 07/10/2024 |  | 05/08/2024 |  |  | 04/30/2024 | 05/08/2024 |  |  | 02/28/2025 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024340999 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | The complainant reported a 78-year old female patient with post-cardiotomy cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during the implant, there was no flow past the cp sheath after removal of the peel away sheath and placement of the repositioning sheath, and the patient¿s leg was dusky and cold. a reperfusion sheath was placed with an external bypass from the patient¿s right radial artery. the patient remains stable. | [5 empty] | The investigation into the patient's limb ischemia has been completed. the impella cp was returned for evaluation and passed all post sterile tests. data logs were not received for evaluation. as the necessary clinical information was not provided, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
26947 | 1220648-2024-10911 | 19271111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19271111&pc=OZD | 04/30/2024 | 07/12/2024 |  | 05/08/2024 |  |  | 04/30/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024435529 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Infusion or Flow Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | The complainant reported a 59-year old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that hemolysis and pump saturation occurred. the patient had been on impella 5.5 for over 30 days. the decision was made to explant the 5.5 and re-implant another 5.5. the patient remains stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported hemolysis and saturated flow has been completed. the cut pump was returned for investigation. biomaterial was observed on the inlet cage. no other signs of physical evidence were reported. no further investigation was performed due to the cut catheter. data logs show a motor current spike while the pump was running on steady p-level 6. there was no trend in the placement signal during saturated flow and hemolysis issue. the saturated pump flow returned to normal at the end of the case run. the cause of the hemolysis issue was most likely biomaterial. the biomaterial could also lead to saturated pump flow. added- d9 device return date d4-corrected model number f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing h6 health imapct code omit 4642 additional device required, 4624 surgical intervention, 4641 unexpected medical intervention, and 4627 device explantation. coded incorrectly.
26948 | 1220648-2024-10912 | 19271463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19271463&pc=OZD | 04/30/2024 | 07/29/2024 |  | 05/08/2024 |  |  | 04/30/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095482 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support, and while on support, the automated impella controller (aic) had a "controller error" alarm. the aic plug was pushed in, and the alarm condition resolved. however, the team was uncertain if the resolution was coincidental, and consequently, replaced the console. there was no report of harm to the patient. | [5 empty] | The automated impella controller (aic) was not received from the customer; and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) red alarm has been completed. the console was returned for evaluation. during functional inspection, the optical bench was found to have low contrast. date logs were reviewed which confirmed that the console displayed ¿optical sensor system error: disconnect cable.¿ most likely, the pump catheter plug might not be inserted completely into the front panel optical connector. after reinserting the pump, there was no ¿optical sensor system error,¿ and the pump was started. this aligns with the claim that the ¿alarm seemed to resolve when impella plug pushed in.¿ additional failures found in the log were ¿controller error (opm comm crc invalid) [optical pump connected] - alarm,¿ event #1045, during the clinical procedure. this confirms the reported failure mode. real-time (rt) logs confirmed that all signals received from the optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. the root cause of the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. added-d9 device return date d2-corrected common device name. d4-corrected model number.
26949 | 1220648-2024-10913 | 19271619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19271619&pc=OZD | 09/12/2022 | 03/26/2025 |  | 05/08/2024 |  |  | 09/12/2022 | 05/08/2024 |  |  | 04/30/2024 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216393 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient experienced bleeding and hemolysis. a blood transfusion was required. no additional information was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event (access site bleeding) are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system instructions for use for the related event (hemolysis) are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26950 | 1220648-2024-10914 | 19271667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19271667&pc=OZD | 04/11/2024 | 06/05/2024 |  | 05/08/2024 |  |  | 04/11/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438138 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2024 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient was on impella 5.5 support due to having mitral valve failure. while on support, the patient pulled on the catheter. the 3-point fixation came off and the pump was pulled back six centimeters. the staff was eventually able to advance the impella back into position and the flows were worked up. the pump access site was recleaned, dressed and secured. the patient expired while on support. there is not enough information to associate the use of the device with the patient's outcome. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. the data logs confirmed the pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle/aorta. no other issues were found on the logs. based on clinical description and data analysis, the root cause of positioning issue was most likely use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26951 | 1220648-2024-10915 | 19272339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19272339&pc=OZD | 05/02/2024 | 03/20/2025 |  | 05/08/2024 |  |  | 05/02/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025451698 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Aortic Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experiencing postcardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 device. during support, the patient was having hemolysis and aortic valve insufficiency. the physician repositioned the impella pump, but the issues remained. due to unresolved aortic valve insufficiency and hemolysis issues, the decision was made to explant the impella pump, and after the pump was removed, there was no aortic valve insufficiency. the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: axillary insertion of impella 5.5 catheter: ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the positioning issue, hemolysis, and aortic valve insufficiency/heart valve damage has been completed. according to the clinical, on the second day of impella 5.5 support there were concerns for hemolysis due to the patient producing red urine. the team believed two long pump times and an underfilled lv post op could have contributed to hemolysis. later the same day the patient was also suffering from aortic valve insufficiency. although the pump was confirmed to be in good position, the decision was made to reposition the pump in hopes of resolving the issue. however, by the next day the hemolysis remained and the decision was made to explant the pump. product evaluation confirmed no abnormalities and the pump later passed hemolysis testing. no data logs were returned for analysis. the most likely cause of the hemolysis was patient condition related. the cause of the heart valve damage was not determined because not enough clinical information was given. d9 date device returned to manufacturer added. d4 model number was inadvertently reported incorrectly on the initial manufacturer device report number. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number. h6 health effect impact code 4628 was inadvertently omitted on the initial manufacturer device report number.
26952 | 1220648-2024-10916 | 19272399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19272399&pc=OZD | 05/01/2024 | 03/19/2025 |  | 05/08/2024 |  |  | 05/01/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1636031 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported ongoing support with an impella 5.5 device for a patient admitted with multiple cardiac and systemic comorbidities. the impella pump was placed for coronary artery bypass graft and mitral valve surgeries and left atrial appendage clipping. on day two of support, the automated impella controller (aic) had controller alarms and consequently was replaced without harm or injury from the interruption. the impella pump support continued with the new aic. | [5 empty] | The automated impella controller (aic) was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. based on clinical details and data logs, the root cause of the aic issue was a software issue was optical bench fw communication protocol anomaly. d9 date device returned to manufacturer added. d2 common device name revised on manufacturer device report in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number.
26953 | 1220648-2024-10917 | 19272419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19272419&pc=OZD | 05/13/2022 | 05/08/2024 |  | 05/08/2024 |  |  | 05/17/2022 | 05/08/2024 |  |  | 12/31/2023 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165694 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Sepsis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/17/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support the patient experienced bleeding including a hematoma. the patient required a blood transfusion as a result of the bleeding. it was also reported that the patient also experienced an ischemic cerebrovascular accident on (b)(6) 2022, after the impella was explanted, and sepsis with an unknown date of occurence. no additional information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
26954 | 1220648-2024-10918 | 19272456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19272456&pc=OZD | 04/28/2024 | 03/25/2025 |  | 05/08/2024 |  |  | 04/28/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 161550 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/28/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was transferred and the cord for the automated impella controller (aic) was damaged. the aic was replaced. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - cabling issue has been completed. the data logs were examined for this complaint but there was no evidence of difficulty charging. device analysis: the console was returned but performed as expected including charging and there was no damage to the power cord. root cause: there was no issue found during the case. the console functioned/operated to specification. d2 common device name revised on manufacturer device report in accordance with updated procedures. d4 model number was inadvertently entered incorrectly on the initial manufacturer device report number. d9 date device returned to manufacturer revised from the initial manufacturer device report number. f6/f8-should have been left blank on the initial manufacturer device report number. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number.
26955 | 1220648-2024-10919 | 19272476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19272476&pc=OZD | 04/30/2024 | 06/10/2024 |  | 05/08/2024 |  |  | 04/30/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1431282 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | The user facility reported the automated impella controller (aic) would not complete the purge solution prime at the start of a new case. a new console was used to complete the procedure. no patient harm was reported. | [5 empty] | The investigation for the reported console (aic) purge issue has been completed. after reviewing the logs, the reported issue with this console not being able to finish priming the purge could not be confirmed. in contrary to what the user reported, it appears that the console was able to complete priming after the start of a new case on the reported occurrence date. however, it was discovered that in one instance after step 3 of the case start wizard, the purge priming bar did not appear to user given priming was already completed at this step. after connecting the pump, the aic skipped straight to step 6 where purge fluid information is entered. this is likely what the user was reporting when they said that the console was not completing the purge solution priming at the start of a new case. the reason the purge priming bar did not appear on the aic is because the purge had already completed priming before the pump was connected. when reviewing the rt logs during this time frame it showed that there was about 800 mmhg of purge pressure before the pump was connected which could indicate one of the following scenarios: the yellow luers were already connected and pump was fully primed.¿ -or-there was increased purge pressure without the pump¿s purge lumen connected, which resulted in the skipping of the priming step when the user would have connected the yellow luers. console was tested after arriving in field service. after testing the console with a pump and purge cassette in the lab for a couple of hours, the reported issue could not be reproduced. console was able to prime without any issues and was functioning normally. the root cause of the reported issue was likely use related. added- d9 device return date d2-updated common device name h4-added device mfr date. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26956 | 1220648-2024-10920 | 19272730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19272730&pc=OZD | 05/01/2024 | 07/02/2024 |  | 05/08/2024 |  |  | 05/01/2024 | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467837 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2024 |  |  |  |  | The complainant reported a 68-year old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that they were having positioning issues. the pump dislodged into the aorta, rendering it unable to be pulled backed or advanced. the medical staff attempted to decrease potential energy in the cable and also cut sutures to decrease any tension. despite their efforts, the medical staff decided to explant and replace the impella 5.5 pump after troubleshooting. the patient remains on support. | [5 empty] | D9 updated as the pump and purge cassette were returned for investigation. the investigation into the positioning issues has been completed since the original report was filed. the pump was returned for investigation. visual inspection found no issues on the pump. as per clinical description, the pump was being pulled out as the pump was dislodged into the aorta due to patient movement. the root cause of positioning issue was use related. g1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer report number 1220648-2024-10920 was submitted. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
26957 | 1220648-2024-10921 | 19273395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19273395&pc=OZD | 02/21/2022 | 05/16/2025 |  | 05/08/2024 |  |  | 02/21/2022 | 05/08/2024 |  |  | 10/31/2023 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151872 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | Patient-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2022 |  |  |  |  | The user facility reported a 59-year-old male on impella cp for mechanical circulatory support. it was reported during implantation the impella cp device would not fit through the sheath. insertion was attempted, but unsuccessful. the impella cp was aborted due to the patient's vascular anatomy. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26958 | 1220648-2024-10924 | 19274813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19274813&pc=OZD | 03/02/2022 | 05/16/2025 |  | 05/08/2024 |  |  | 03/02/2022 | 05/08/2024 |  |  | 06/30/2023 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022095291 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2022 |  |  |  |  | The complainant reported a patient in st-elevation myocardial infarction (stemi) was implanted with an impella cp device for mechanical circulatory support. while on support, the patient¿s partial thromboplastin time (ptt) was elevated and bleeding from the nasogastric (ng) tube was noted. heparin was removed from the purge solution. the patient was reportedly hemodynamically stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26959 | 1220648-2024-10926 | 19274834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19274834&pc=OZD | 02/27/2022 | 05/08/2024 |  | 05/08/2024 |  |  | 02/27/2022 | 05/08/2024 |  |  | 11/30/2023 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168672 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Failure | Restricted Flow rate | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/27/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding. the nurse adjusted angle of entry and the oozing resolved. additionally, the impella was repositioned on two attempts due to being too deep. pressure was also increased to p9 from p8. the patient reportedly showed improvement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26960 | 1220648-2024-10922 | 19274884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19274884&pc=OZD | 02/23/2022 | 05/19/2025 |  | 05/08/2024 |  |  | 02/23/2022 | 05/08/2024 |  |  | 10/31/2023 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151994 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding, hemolysis, ventricular fibrillation (vf), atrial fibrillation (afib), and an issue with the positioning of the impella. the patient had vf arrest, that required cardio-pulmonary resuscitation for 30 minutes until the return of spontaneous circulation was achieved. the patient¿s urine was reported to be dark red since the pump was placed, however with the addition of lasix the urine was pink tinged. the patient also went into afib rapid ventricular response, amiodarone was administered to attempt to lower the heart rate. it was reported the patient made his way to the end of the bed and attempted to get up. the impella started alarming that the pump had been pulled back into the aorta. the patient was taken emergently down to the cardiac catheterization lab, to attempt to reposition the device across the aortic valve due to the pump not successfully advancing in the intensive care unit under echocardiogram. the pump cannula was found to be kinked; the kink was removed by advancing first, then pulling down while torquing the device. the impella 0.035 wire was removed and replaced with a new 14sheath and impella cp. the swan was repositioned again, pulmonary capillary wedge pressure was fourteen. it was reported there was considerable oozing at the insertion site, resolved by the nurse changed the dressing. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis: impella 5.5 with smart assist for use during cardiogenic shock & impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter as noted in the impella ifu ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle as noted in the impella ifu ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis as noted in the impella ifu ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ as noted in the impella ifu ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction as noted in the impella ifu ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26961 | 1220648-2024-10923 | 19274898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19274898&pc=OZD | 02/23/2022 | 05/08/2024 |  | 05/08/2024 |  |  | 02/23/2022 | 05/08/2024 |  |  | 08/31/2023 |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022130761 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/23/2022 |  |  |  |  | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced excessive bleeding from the chest tube. heparin was withheld. additionally, the patient developed a mottled foot that continued to get worse over a few days. the patient also required blood transfusions. the decision was made to remove the impella cp. . | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ limb ischemia : impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26962 | MW5154669 | 19278769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19278769&pc=OZD | 03/21/2024 | 05/06/2024 |  | 05/08/2024 |  |  |  | 05/08/2024 |  |  |  |  | 05/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 HEART PUMP |  |  | OZD | [5 empty] | JOHNSON & JOHNSON / ABIOMED, INC. | [8 empty] | No | [9 empty] | Voluntary report | Not Available | Patient Family Member or Friend | [6 empty] | 65 YR |  | Male | 77 KG | Non Hispanic | Asian, White | Injury | Yes | Yes | [7 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | ASPIRIN. | HEPARIN. | LEVOPHED. | MILRINONE. | VASOPRESSIN. | [18 empty] | Initial submission | [9 empty] | My husband had the impella 5.5 heart pump installed on (b)(6) 2024 at (b)(6) hospital and developed blood clots in his internal jugular vein and cephalic veins on (b)(6) 2024. we were not told about these clots until (b)(6) 2024.
26963 | 1220648-2024-10822 | 19257910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19257910&pc=OZD | 04/04/2024 | 05/07/2024 |  | 05/07/2024 |  |  | 04/04/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314860 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 04/04/2024 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that an automated impella controller (aic) failed to recognize an installed purge cassette during impella support. upon checking the purge disc area, the registered nurse stated that the purge disc was loose as a result of a cracked purge disc retainer. the aic was swapped as a result of the noted damage. there was no patient harm. | [5 empty] | The investigation has been completed. log files were reviewed and the reported purge disc not detected alarm issue was confirmed. there were multiple instances of purge disc not detected alarm found on the reported occurrence date. the console was tested. the reported purge disc not detected alarm issue was able to be reproduced. after visual inspection of the console, it was discovered that there was a hairline fracture on the left side of the purge disc retainer. the purge disc not detected alarms were determined to be caused by this fracture in the purge retainer. the purge disc not detected alarms were determined to be caused by this fracture in the purge retainer.
26964 | 1220648-2024-10823 | 19257958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19257958&pc=OZD | 04/05/2024 | 08/15/2024 |  | 05/07/2024 |  |  | 04/05/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1367614 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/05/2024 |  |  |  |  | The user facility reported a male patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the automated impella controller (aic) displayed a controller error alarm during patient support. the issue was temporarily resolved until two days later, the position waveform disappeared and the controller error alarm returned. support continued with the device until it could be returned following the completion of support. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the optical signal issue has been completed. data logs confirmed the reported failure mode given there was a controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as 16384 values due to the crc error. the console was returned for evaluation. the reported error was found to be caused by a failure mode which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, and has a low probability of reoccurrence. the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. f6/f8 should have been left blank in the initial report.
26965 | 1220648-2024-10824 | 19258008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19258008&pc=OZD | 04/27/2024 | 07/09/2024 |  | 05/07/2024 |  |  | 04/30/2024 | 05/07/2024 |  |  | 04/30/2025 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024339506 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an unknown age patient had percutaneous coronary intervention ongoing when the impella cp pump was primed for support. after priming, the pump stopped and was troubleshot by the replacement of the automated impella controller (aic). however, replacement of the aic did not remedy the issue. the team then replaced the impella pump, and all was fine, support proceeded. there was no report of harm to the patient. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported pump stop has been completed. impella failure alarm was triggered when the pump was connected to the aic. no product or data logs were returned. the root cause of the pump not detected issue was not determined because no product or data logs were returned.
26966 | 1220648-2024-10825 | 19258049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19258049&pc=OZD | 03/03/2022 | 05/07/2024 |  | 05/07/2024 |  |  | 03/03/2022 | 05/07/2024 |  |  | 11/30/2023 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169871 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/03/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced limb ischemia. the patient right lower extremity was mottled/purple. they were treated with warm blankets and an external femoral to femoral bypass was performed for distal perfusion. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26967 | 1220648-2024-10826 | 19258060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19258060&pc=OZD | 04/30/2024 | 07/02/2024 |  | 05/07/2024 |  |  | 04/30/2024 | 05/07/2024 |  |  | 05/31/2025 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024351908 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. upon starting impella 5.5 support, the patient became hypotensive and akinesis was observed via transesophageal echocardiogram. at this time, the pump also experienced a loss in motor current and waveforms. the patient required advanced cardiac life support measures and clot ingestion by the pump was suspected. the pump was subsequently replaced in response to the noted issues. | [5 empty] | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was submitted. the product and data logs were not returned for investigation. the root cause of low flow was unable to be determined as limited clinical details were provided and no product or logs were returned for review. g1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-10826 was submitted. h6 code 4627 was reported incorrectly on the initial manufacturer device report 1220648-2024-10826. new codes have been added to health effect - impact codes.
26968 | 1220648-2024-10827 | 19258127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19258127&pc=OZD | 03/11/2022 | 05/21/2025 |  | 05/07/2024 |  |  | 03/11/2022 | 05/07/2024 |  |  | 12/12/2023 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177778 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced significant bleeding from the nose and mouth, and the patient's heparin was held due to the bleeding. the patient¿s hemoglobin dropped, and the patient received two units of packed red blood cells. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26969 | 1220648-2024-10828 | 19258287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19258287&pc=OZD | 03/18/2022 | 03/26/2025 |  | 05/07/2024 |  |  | 03/18/2022 | 05/07/2024 |  |  | 11/30/2023 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165815 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding. the oozing was addressed with dressing the impella site to allow for angle matching, a forward tension suture, and pressure applied to the site. the patient remained on impella support and was successfully weaned over a 24 hours. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26970 | 1220648-2024-10829 | 19258344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19258344&pc=OZD | 04/30/2024 | 06/21/2024 |  | 05/07/2024 |  |  | 04/30/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025467757 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an attempted implant of an impella cp pump was unsuccessful on a patient in acute myocardial infarction/cardiogenic shock due to resistance from the patient's anatomy (unable to pass through the vasculature). the implant was subsequently aborted. there was no report of patient harm. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of delivery issue was most likely patient condition since circumferential, calcified lesion was noted at the origin. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
26971 | 1220648-2024-10830 | 19258598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19258598&pc=OZD | 04/01/2022 | 05/07/2024 |  | 05/07/2024 |  |  | 04/01/2022 | 05/07/2024 |  |  | 12/31/2023 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023175328 |  | 0550-0008 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/01/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient¿s pulmonary edema was worsening, and the decision was made to escalate to an impella 5.5. while on support, the patient developed heparin induced thrombocytopenia. heparin was discontinued and bivalirudin was infused for anticoagulation therapy. the patient remained on impella support and was successfully weaned. | [5 empty] | D6b: explantation date is unknown the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
26972 | 1220648-2024-10831 | 19258657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19258657&pc=OZD | 02/25/2022 | 05/07/2024 |  | 05/07/2024 |  |  | 02/25/2022 | 05/07/2024 |  |  | 08/31/2023 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022130319 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/25/2022 |  |  |  |  | The user facility reported a patient was placed onto impella cp support for high-risk percutaneous coronary intervention. while on support, the patient experienced limb ischemia that was resolved by removing the impella, though the patient was already sufficiently stable to remove it regardless of the presence of ischemia. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
26973 | 1220648-2024-10832 | 19258748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19258748&pc=OZD | 03/01/2022 | 03/26/2025 |  | 05/07/2024 |  |  | 03/01/2022 | 05/07/2024 |  |  | 11/30/2023 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162012 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/01/2022 |  |  |  |  | The complainant reported a patient in st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, site oozing was noted. pressure was held, forward sutured, angle matched, and femstop was placed. bleeding stopped when femstop was applied; however, site oozing was later noted again. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26974 | 1220648-2024-10833 | 19258775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19258775&pc=OZD | 04/08/2024 | 06/10/2024 |  | 05/07/2024 |  |  | 04/08/2024 | 05/07/2024 |  |  | 02/28/2025 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024336368 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Intraoperative Cardiac Valve Injury | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a female of unknown age with an acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, the patient's aortic valve was damaged. the impella 5.5 was explanted and a lvad was implanted instead. the impella was discarded. the patient was noted to be stable while this happened. the patient expired while no longer on impella support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ | The investigation for the reported heart valve damage has been completed. product was not returned for investigation. as per given clinical, there was damage to the aortic valve. the impella was explanted, and the patient remained stable with no harm. the cause of the heart valve damage issue was not determined since product was not returned and sufficient clinical information was not provided. the instructions for use referenced in the initial report is for the impella 5.5 with smart assist for use during cardiogenic shock in section: potential adverse events (unites states) . h1-corrected type of reportable event to death.
26975 | 1220648-2024-10834 | 19258835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19258835&pc=OZD | 04/08/2024 | 06/06/2024 |  | 05/07/2024 |  |  | 04/08/2024 | 05/07/2024 |  |  | 10/31/2025 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024413156 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2024 |  |  |  |  | The user facility reported a 37-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the user facility reported high purge pressure alarm while an impella 5.5 was in use with a patient. there was no patient harm reported. | [5 empty] | The investigation into the high purge pressure issue has been completed since the original report was submitted. no product was returned for evaluation. the data logs were reviewed and a gradual rise in purge pressure was observed over a period of 16 hours. no motor current spikes or speed deviations were observed. clinical details noted that the patient was receiving lots of blood products while on support due to a gastrointestinal bleed unrelated to the pump and no anticoagulation was used on the patient. the root cause of the high purge pressure is most likely due to gradual buildup of biomaterial as a result of no anticoagulation management while on support. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. h.6 medical device problem codes was revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
26976 | 1220648-2024-10835 | 19258880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19258880&pc=OZD | 04/08/2024 | 05/20/2024 |  | 05/07/2024 |  |  | 04/08/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025469165 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2024 |  |  |  |  | The user facility reported a 72-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella cp support was having intermittent pulses of the left lower extremity (lle) and there was concern for limb ischemia. the plan is to have ble [bilateral (or both) lower extremity] ultrasound to assess flow. patient passed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
26977 | 1220648-2024-10836 | 19258925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19258925&pc=OZD | 04/12/2024 | 06/03/2024 |  | 05/07/2024 |  |  | 04/12/2024 | 05/07/2024 |  |  | 07/31/2025 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024384057 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the patient developed an ischemic leg. a weaning trial was initiated and the impella was removed in the cath lab. a small dose of epinephrine was kept running. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia could not be determined as the device was not returned nor sufficient clinical details. the instructions for use referenced in the initial report is for the impella cp with smart assist system.
26978 | 1220648-2024-10837 | 19258977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19258977&pc=OZD | 03/02/2022 | 03/26/2025 |  | 05/07/2024 |  |  | 03/02/2022 | 05/07/2024 |  |  | 11/30/2023 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023160519 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced hemolysis. the urine was very dark. the position of the impella was verified to be stable. patient had hemolysis throughout the night with pfh >2000 platelets 145. the physician was advised to reposition the catheter despite echo imaging appearing ideal. the impella was pulled back 1cm and is now at 44cm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26979 | 1220648-2024-10838 | 19258993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19258993&pc=OZD | 04/18/2024 | 05/31/2024 |  | 05/07/2024 |  |  | 04/18/2024 | 05/07/2024 |  |  | 09/30/2025 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024405109 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella cp was brought to the operation room for 5.5 escalation. cp flows turned down and cp removed from left ventricle (lv). subsequently the right ventricle was distended and not filling the lv. decision made to start code and place peripheral va ecmo (extracorporeal membrane oxygenation). the patient stabilized and the cp was surgically removed and repaired. while removing, there was a noticeable amount of clot in the iliacs. this clot was removed, and the flow was restored. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ | The investigation into the thrombus issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary information was not provided, the root cause of the thrombus was not determined.
26980 | 1220648-2024-10839 | 19259051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19259051&pc=PYX | 10/29/2021 | 06/13/2024 |  | 05/07/2024 |  |  | 04/23/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The literary publication reported that there was a floating thrombus in the inferior vena cava (ivc) to the right atrium noted following impella rp removal for critically ill patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombus issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary information was not provided, the root cause of the thrombus was not determined.
26981 | 1220648-2024-10840 | 19259168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19259168&pc=OZD | 03/02/2022 | 03/26/2025 |  | 05/07/2024 |  |  | 03/02/2022 | 05/07/2024 |  |  | 06/30/2023 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022101203 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced some oozing from access site. all heparin was removed and sandbag on the groin. in addition, the heparin induced thrombocytopenia panel was sent and blood products were given. it was noted that the patient was stable, and oozing had subsided, however they were still holding heparin. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26982 | 1220648-2024-10841 | 19259231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19259231&pc=PYX | 12/05/2020 | 06/11/2024 |  | 05/07/2024 |  |  | 04/23/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0046-0011 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | Company representation |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaint from the literature review reported that the 49-year-old male patient in st elevated myocardial infarction was implanted with an impella rp device for mechanical circulatory support. after two days of impella rp support, there was bleeding from the axillary access site that required temporary suspension on anticoagulants and blood products transfusion. after ten days of support, the patient was successfully weaned of the right-ventricular support. | [5 empty] | D.4 serial number, lot number, expiration date and unique identifier (udi) # are unknown. g.6b explant date is unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
26983 | 1220648-2024-10842 | 19259236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19259236&pc=OZD | 04/22/2024 | 06/24/2024 |  | 05/07/2024 |  |  | 04/22/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2025481813 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella 5.5 implantation, air bubbles were detected sonographically. the patient subsequently showed slight neurological deficits. the patient was noted to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into the stroke/cva issue has been completed since the original report was filed. the pump was returned for investigation. as clinical information related to the stroke/cva was not provided, the root cause could not be determined. clinical details reported air bubbles were detected. the root cause of the embolism was not determined due to limited clinical information.
26984 | 1220648-2024-10843 | 19259386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19259386&pc=OZD | 03/02/2022 | 03/26/2025 |  | 05/07/2024 |  |  | 03/02/2022 | 05/07/2024 |  |  | 06/30/2023 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022093248 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had continuous oozing at the impella groin site. the heparin was stopped, and two (2) units of packed red blood cells were administered to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
26985 | 1220648-2024-10844 | 19259444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19259444&pc=PYX | 04/23/2024 | 07/02/2024 |  | 05/07/2024 |  |  | 04/23/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Positioning Failure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tricuspid Valve Insufficiency/ Regurgitation | [12 empty] | Life Threatening | Death | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant from the italian publication entitled "timing, positioning and ancillary therapies: the triad for successful impella rp" outlined three patient cases. case two had "tricuspid regurgitation emerged due to device interference with septal tricuspid leaflet" and case three had pump malpositioning and kinking observed. case two survived and case three expired. | This report will represent patient number two. patient number two had "tricuspid regurgitation emerged due to device interference with septal tricuspid leaflet". the patient survived. manufacturer report number 1220648-2024-12890 represents patient number three. |  |  |  |  | D.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella rp with smart assist system with automated impella controller section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. " in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | A.1, a.4 and a.5 are unknown. the investigation into the reported issue has been completed since the original report was submitted. the complaint is from a literature report with missing information. no case numbers, device serial numbers, or clinical information was provided. no product or data logs were available for review. the root cause of heart valve damage was not determined due to lack of clinical information and no returned product. literature citation: bertolin, s., maj, g., secco, g. g., pullara, a., gherli, r., cardinale, a., d¿ettore, n., cavozza, c., audo, a., & pappalardo, f. (2023). timing, positioning and ancillary therapies: the triad for successful impella rp. journal of cardiovascular medicine. https://doi.org/10.2459/jcm.0000000000001563 a.3 revised in accordance with updated procedures since the initial manufacturer device report 1220648-2024-10844 was submitted. b.1 product problem was selected on the initial manufacturer device report 1220648-2024-10844 and should not have been. b.2 death was selected on the initial manufacturer device report 1220648-2024-10844 and should not have been. b.5 revised to reflect patient number two only since the initial manufacturer device report 1220648-2024-10844 was submitted. e.1 facility name was revised as previously entered incorrectly on the initial manufacturer device report 1220648-2024-10844. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-10844 was submitted. h.6 codes 1802,1158 and 925 were reported incorrectly on the previously submitted report. new codes have been added to medical device problem code and health effect - impact code. h.10 a.1, a.4 and a.5 are unknown statement and literature citation were inadvertently omitted from the initial manufacturer device report 1220648-2024-10844. attachments: publication was inadvertently omitted from the initial manufacturer device report 1220648-2024-10844.
26986 | 1220648-2024-10845 | 19259454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19259454&pc=OZD | 04/22/2024 | 07/09/2024 |  | 05/07/2024 |  |  | 04/22/2024 | 05/07/2024 |  |  | 08/31/2024 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252794 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/22/2024 |  |  |  |  | The complainant from the study reported that during impella cp support, the patient experienced myocardial wall perforation that was classified as severe. the patient underwent an echocardiogram which revealed left ventricle rupture with aneurysm in the mid infero-lateral wall. the patient required prolonged hospitalization and required intervention. the patient underwent cardiac surgery with the placement of prolene 3-0 sutures with pledgets pm the middle posterior wall and simultaneous bypass for complete revascularization using venous graft. the event was determined to be unrelated to the device, but probably related to the impella procedure. the patient is noted to have recovered without residual effects. | [5 empty] | The device was not returned for evaluation. upon completion of the investigation, a supplemental mdr will be filed. potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation for the reported ventricle perforation has been completed. the patient experienced myocardial wall perforation that was classified as severe; wall perforation confirmed via echo. no product or data logs were returned. the root cause of the ventricle perforation was not determined due to lack of sufficient clinical information. added- h6 component code 925 d4-corrected model number f6/f8-should have been left blank in the initial report. g1 manufacturing site country. g2 report source added health professional.
26987 | 1220648-2024-10846 | 19259528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19259528&pc=OZD | 04/22/2024 | 06/24/2024 |  | 05/07/2024 |  |  | 04/23/2024 | 05/07/2024 |  |  | 10/31/2025 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024420159 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 51-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while having the impella implanted the patient developed ventricular fibrillation (vf). the patient required defibrillation (1x). the patient remained stable afterwards. the patient did not need inotropes or vasopressors anymore. the impella cp was successfully explanted a day later, and the patient was stable and awake | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation into the vf/vt/af/at (ventricular fibrillation) issue has been completed since the original report was submitted. the product was not returned for investigation. as the necessary information was not provided, the root cause of the vt/vf was not determined.
26988 | 1220648-2024-10847 | 19259605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19259605&pc=PYX | 04/23/2024 | 06/25/2024 |  | 05/07/2024 |  |  | 04/23/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | Disability | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaint from the literature review reported that four days later after bi-impella support, the patient developed acute pulmonary edema and a new onset of impella-mediated severe mitral regurgitation (mr) due to a3¿p3 flail. considering the profound thrombocytopenia, unclear neurological status, and acute liver failure, the surgical risk was prohibitive. therefore, the heart team decided on a mitraclip procedure. the patient expired from multi-organ failure after 22 days of impella support. | [5 empty] | D.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombocytopenia issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the issue was not determined due to insufficient clinical details.
26989 | 1220648-2024-10848 | 19259608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19259608&pc=OZD | 04/23/2024 | 07/11/2024 |  | 05/07/2024 |  |  | 04/23/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025478131 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the procedure for an impella cp placement for an of unknown age and gender patient, insertion was not possible into the aortic valve due to abnormal vasculature and the impella cp cannula kinked during implantation. therefore, a new impella was placed, and the procedure was completed successfully. the patient is noted to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation for the reported delivery issues has been completed. the pump was received for inspection and the canula kink was observed. the pump was unable to pass through the vasculature and pump kinked during implantation due to resistance at aortic valve during insertion. the root cause of the delivery issue is determined to be patient condition because of the resistance at the aortic valve during implantation of pump. the instructions for use referenced in the initial report is for the impella cp with smartassist system in the following sections: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ added- d9 device return date g1 corrected manufacturing site fax number.
26990 | 1220648-2024-10849 | 19259856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19259856&pc=OZD | 04/25/2024 | 07/02/2024 |  | 05/07/2024 |  |  | 04/25/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025443658 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The user facility reported a patient with unknown demographics in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient experienced bleeding from the insertion site. this was noted to be most likely due to low heparin clearance due to liver failure. the bleeding was able to be stopped after repositioning to 40 degrees and a forward suture. the patient was successfully weaned when hemodynamics had stabilized and the impella was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved by adding forward suture and repositioning to 40 degrees.
26991 | 1220648-2024-10850 | 19259928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19259928&pc=PYX | 04/25/2024 | 06/13/2024 |  | 05/07/2024 |  |  | 04/25/2024 | 05/07/2024 |  |  | 11/30/2025 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024434830 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The user facility reported a 45-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during patient preparation in the operation room for impella rp placement, it was noted resistance upon insertion of the ij swan and unable to advance it past the subclavian vein. it remained in the subclavian vein during the insertion of the rp flex. rp flex inserted without complication. flows at 4.5 liters. after insertion of the flex, the left ij swan noted to be tangled within the vein and was unable to be explanted. ultimately a new right femoral venous access was made to snare the swan. at this time there was a sharp marked increase in placement signal and pulmonary artery pressure, drop in motor current, and acute reduction in rp flows from 4.5 liters to 3 liters. md decision was to see how the patient tolerated 3 liters of flow and monitoring pump metrics. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the reported low pump flow has been completed. according to the clinical, on the first day of impella rp support the ij swan was removed and replaced. after a successful insertion, it could not be advanced past the subclavian vein. when attempting to remove the swan after inserting the pump it was discovered that it was tangled in the vein. as right femoral venous access was made to snare the swan, placement signal and pa pressure increased while motor current and flow decreased. the decision was then made to leave the pump in the patient and monitor support. no product was returned for a visual inspection. data log analysis confirmed that the placement signal, pa pressure, motor current, and flow shifts occurred. the motor current shifted downward by 100ma and no spikes were observed. the flow shifted downward because it is calculated by the inverse of placement signal. the cause of the low pump flow was not determined because no product was returned and not enough clinical information was given. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high risk cpi in using the automated impella controller with the impella catheter and using the automated impella controller with the impella rp with smartassist system catheter sections. d4-updated model number h5-changed labeled for single use to yes.
26992 | 1220648-2024-10851 | 19260014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19260014&pc=OZD | 04/26/2024 | 06/04/2024 |  | 05/07/2024 |  |  | 04/26/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025492165 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2024 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient lost distal pulses in their leg during impella cp support. due to the noted condition, plans were made to place a distal perfusion catheter and subsequently transfer the patient for ongoing support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use referenced in the initial report is for the cp with smartassist. in section: ¿potential adverse events (united states). d4-updated model number.
26993 | 1220648-2024-10854 | 19260196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19260196&pc=OZD | 04/29/2024 | 07/01/2024 |  | 05/07/2024 |  |  | 04/29/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443199 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella device for mechanical circulatory support. however, originally, the team did not start systemic heparin drip due to a hematoma at cutdown site. it was additionally noted, the patient has been receiving subcutaneous heprin doses, and there has been no impella device purge-related issue or ingestion of clot. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. additional information was received. d.6b explant date has been updated. d.4 revised model number from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures.
26994 | 1220648-2024-10855 | 19260202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19260202&pc=OZD | 04/28/2024 | 06/07/2024 |  | 05/07/2024 |  |  | 04/28/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453420 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient (age and primary indication unknown) was implanted with the impella 5.5 device for mechanical circulatory support. during support, there was high purge pressure and low purge flow. the purge cassette was replaced, and tissue plasminogen activator (tpa) was added to the purge. however, the issue did not resolve. additionally noted, there were no kinks observed on the purge line. there was no report of harm to the patient. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the high purge pressure was not determined since product and data logs were not returned.
26995 | 1220648-2024-10856 | 19260288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19260288&pc=OZD | 04/30/2024 | 09/11/2024 |  | 05/07/2024 |  |  | 04/30/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451014 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhagic Stroke | Ischemia Stroke | [10 empty] | Disability | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support, and while on support, the patient developed a brain bleed and subsequently expired. fifth day on support cardiac recovery was noted during ramp study, and a plan to wean the next day for possible removal of the impella pump was noted. however, overnight, the patient stopped responding to stimuli. on assessment, pupils were uneven; ct scan showed substantial brain bleed. care was withdrawn, and the patient expired. the physician attributed the event to impella device and noted the concern was potential for clot ingestion in the pump resulting in an embolic stroke that converted to hemorrhagic. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system-section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. | The following have been reported incorrectly or missing from manufacturer device report 1220648-2024-10856. g1 reporting contact fax number missing. h6 code 4418 and 4417 were reported incorrectly. new code was added to health effect - clinical code and component code. on manufacturer device report 1220648-2024-10856-1 h1 type of reportable event was incorrect. | The investigation into the stroke/cva issue has been completed since the original report was submitted. the device was not returned for investigation. as the clinical information related to the reported issue was not provided, the root cause of the stroke/cva could not be determined.
26996 | 1220648-2024-10857 | 19260337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19260337&pc=OZD | 04/29/2024 | 06/10/2024 |  | 05/07/2024 |  |  | 04/29/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467704 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2024 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient experienced access site bleeding. the patient had oozing at the access site and their hemoglobin and hematocrit was dropping. blood products were given to the patient and heparin was withheld. the patient remains on impella 5.5 support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and sufficient information is not provided in the clinical description.
26997 | 1220648-2024-10858 | 19260572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19260572&pc=OZD | 04/30/2024 | 07/03/2024 |  | 05/07/2024 |  |  | 04/30/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025469150 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | The user facility reported a 52-year-old male patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported that the impella cp had high purge pressure. the medical staff attempted to start tissue plasminogen activator, but the impella cp pump stopped. the medical staff then attempted to restart the pump several times unsuccessfully. the medical staff noted that when the impella cp was implanted, activated clotting time was 170. the impella cp was placed in the same site as the intra-aortic balloon pump. the impella cp was eventually replaced. the patient remains on support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the high purge pressure and the pump stop has been completed. the pump was returned for investigation. no sign of physical damage was noted. the pump electrical testing showed that one of the three motor current values was out of specification. the pump stopped with a high motor current alarm appearing during the test run. the purge system was blocked during testing. the catheter was cut open near the motor housing and purge fluid exited successfully. upon further investigation, biomaterial was revealed in the motor housing. the data logs show a 2-hour gradual rise in purge pressure followed by the several pump stop alarm with high motor current at the end of the case. the cause of the pump stop issue was biomaterial found inside motor housing. this also affected the purge pressure, causing it to be high. added- d9 device return date f6/f8 should have been left blank.
26998 | 1220648-2024-10859 | 19260802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19260802&pc=OZD | 04/30/2024 | 07/12/2024 |  | 05/07/2024 |  |  | 04/30/2024 | 05/07/2024 |  |  | 11/30/2025 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433106 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  |  | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Stroke/CVA | Ischemia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | The user facility reported a 38-year-old female patient of unknown ethnicity with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that there were no doplerable pulses in the left foot found. fem-fem bypass was inserted with adequate distal limb perfusion. it was further reported that the physician stated that the impella cp was difficult to maintain proper position. while escalating to central venous-arterial extra-corporeal membrane oxygenation, the impella cp inlet was observed to be above the aortic valve. the physician suspected the impella cp showered a clot from an unknown origin resulting in an embolic stroke. a clot was found on exterior of impella cp upon explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smart assist system for use during cardiogenic shock section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for limb ischemia, stroke, thrombus, and positioning issues has been completed. the product was not returned for investigation. data logs were returned and multiple position alarms were noticed. patient also had a severe mitral regurgitation. therefore, per clinical information and data log review, the root cause of the positioning issue was patient condition related to diseased valve. without additional clinical details, the root cause of the limb ischemia, thrombus, and stroke could not be determined. added concomitant medical products f6/f8-should have been left blank in the initial report. g1 reporting contact fax number missing.
26999 | 1220648-2024-10860 | 19260898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19260898&pc=OZD | 05/01/2024 | 06/11/2024 |  | 05/07/2024 |  |  | 05/01/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025460801 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. two sutures were placed to achieve hemostasis which were unsuccessful as the sutures would not hold to achieve closure. a 14 french sheath was placed, and hemostasis was obtained. the patient was taken to the operating room for removal of the impella cp pump. | [5 empty] | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue because the failure of proglides and sheath was placed to achieve the hemostasis. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27000 | 1220648-2024-10861 | 19260980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19260980&pc=OZD | 04/23/2024 | 06/10/2024 |  | 05/07/2024 |  |  | 04/23/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447818 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient had no pedal pulse during impella cp support. in response to the noted condition, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use referenced in the initial report is for the impella cp with smartassist system in the potential adverse events section. added- h5 labeled for single use should be yes d4-updated model number h1-updated type of reportable event to serious injury.
27001 | 1220648-2024-10862 | 19261103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19261103&pc=OZD | 05/01/2024 | 06/25/2024 |  | 05/07/2024 |  |  | 05/01/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025489012 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock had an attempted impella cp implant that was unsuccessful as the guidewire was unable to be pass through the pigtail/inlet. the decision for a new device was made. there was no report of patient harm. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation into the delivery issues has been completed since the original report was filed. the pump was returned for investigation. guidewire was not returned for review. the product had a red lumen that was pulled out about a couple of centimeters from the pump. as per clinical reports, the doctor was not able to wire the pump. the cause of the delivery issue was most likely use related.
27002 | 1220648-2024-10863 | 19261161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19261161&pc=OZD | 05/02/2024 | 03/06/2025 |  | 05/07/2024 |  |  | 05/02/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1814768 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a new automated impella controller (aic) was received and upon boot up the console alarmed with an advisory, "unexpected controller shutdown." the aic was restarted, and the alarm condition was no longer present. there was no patient involvement. | [5 empty] | The automated impella controller (aic) was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) power issues has been completed. console logs from the reported complaint date are consistent with what was reported in the complaint description. an unexpected controller shutdown [event set #603] alarm was observed in the imc logs on the first boot-up of the console by the customer on the complaint date. console was not returned for evaluation. according to the dhr for ic11549, this console was still at abiomed on (b)(6) 2024, which was the previous boot of the console prior to the alarm according to the imc logs. this console was not properly shut down in electronics manufacturing prior to shipment from abiomed, which resulted in an unexpected controller shutdown prompt appearing on the screen when the console was first powered on by the customer. quality inspection did a final inspection of the console prior to packaging of the console and is supposed to shutdown the console prior to disengaging the battery transport switch. it is likely that the battery transport switch was disengaged prior to the console completing the shutdown process. this was not an indication of equipment failure and the device operates within specification. the cause of the aic issue was a work instruction / technician error. corrections to the initial mfr report: d2 updated common device name d4-updated model number g1 mdr reporting contact email.
27003 | 1220648-2024-10864 | 19261295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19261295&pc=OZD | 11/13/2023 | 05/20/2024 |  | 05/07/2024 |  |  | 04/10/2024 | 05/07/2024 |  |  | 09/30/2025 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024402443 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured right common and external iliac artery dissection with a "possible" relationship to the impella device used: ¿pt had planned angiogram to assess the iliac arteries. found to have an active dissection involving the right common and external iliac arteries. pta and stenting of the rcia and eia performed.¿ | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation into the vascular damage - great vessel issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the vascular damage - peripheral vessels cannot be determined due to insufficient clinical information provided. d6a, d6b.
27004 | 1220648-2024-10865 | 19261352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19261352&pc=PYX | 11/16/2023 | 02/10/2025 |  | 05/07/2024 |  |  | 04/11/2024 | 05/07/2024 |  |  | 02/28/2025 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024340769 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Tachycardia | Heart Failure/Congestive Heart Failure | [10 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 04/11/2024 |  |  |  |  | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured worsening of congestive heart failure and experienced ventricular tachycardia those both possibly related the imeplla device. | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured worsening of renal function with a "possible" relationship to the impella device used: ¿acute renal failure superimposed on ckd. patient now had end-renal disease. permcath replaced on (b)(6) 2023 due to malfunction. baseline creatinine: 3.74-3.88 on (b)(6) 2023 peak creatinine: 4.72 on (b)(6) 2023.¿ |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. should additional information be received a supplemental mdr will be filed. instructions for use: impella rp system with automated impella controller & impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ b5 updated with additional information received h6 health effect - clinical code and health effect - impact code updated to reflect new information the following have been reported incorrectly or missing from manufacturer device report. f6/f8-should have been left blank. g1 reporting contact email revised in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The investigation for the reported renal failure has been completed. no product was returned for investigation. the data logs show that there are no motor current spikes or positioning alarms. there is a trend in the placement signal and the flow. the root cause of the renal failure is most likely due to the patients condition. d4-updated model number g1- updated mfr site name and address.
27005 | 1220648-2024-10866 | 19261456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19261456&pc=OZD | 01/01/2024 | 10/10/2024 |  | 05/07/2024 |  |  | 04/12/2024 | 05/07/2024 |  |  | 10/31/2025 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407776 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 32 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured worsening respiratory failure with a "possible / definite / probable" relationship to the impella device used: ¿patient was positive previously prior to the impella with rhinocirus/enterovirus 11/29/23 and was treated with antibiotics. patient was monitored for breathing throughout (also had hx of pe). patient intubated on (b)(6) 24 or for impella placement and was extubated 1/4/24 1600. patient was intubated again 1/20/24 until 1/22/24.¿ additionally, the patient endured progression of ventricular tachycardia with a "possible" relationship to the impella device used: ¿patient had vt run prior to admission on site from osh. after patient was started on amiodarone, patient did not have any vt since (ep following). after impella implant, patient was cardioverted for vt and cardioverted again 1/2/24 for second run of vt and was started on amiodarone drip and bolus. 1/8/24 for vt 200 hr. cardioverted and lidocaine started.¿ | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. added- h6 component code corrections to the initial mfr report: d4-model number. f6/f8 should have been left blank.
27006 | 1220648-2024-10867 | 19261582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19261582&pc=OZD | 04/10/2024 | 07/16/2024 |  | 05/07/2024 |  |  | 04/15/2024 | 05/07/2024 |  |  |  |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451922 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiogenic Shock | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating an allegation of cardiogenic shock: ¿cardiogenic shock 4/10-support: impella p7, epi 0.03, prn levo (at 0.14)/hemodynamics 2 pm (on epi 0.04, impella p7 with levo 0.1 for bp)/cvp 13/svo2 57 (was 35 before epi 0.04 this am) ci/co/svr 2.3/4.8/1125/ watching lactate, bmp, svo2 q4h for now/daily ldh with impella.4/12-cvp 12, svo2 62, co 7.9, ci 3.7, svr 505 lactate 4.34, down trending, epi weaned to 0.03 with transient increase in svo2, stable at 60. net + fluid balance since admit and worsening due to multiple drips. consolidating fluids as able, crrt with ufr uptitrated by nephro albeit limited by hypotension, vaso at 0.58 and climbing. doubt infection but given multiple lines, h/o endocarditis, declining svr, cultures negative.4/13- transition to inpatient hospice.¿ | It was further reported that care was withdrawn from the patient and the patient subsequently expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate the use of the device with the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | B.5 additional information was received, and abiomed's medical safety officer review was completed. the investigation for the reported cardiogenic shock has been completed. the device nor sufficient clinical information was returned for evaluation. as the necessary information was not provided, the root cause was not determined. b.7 revised as this information was inadvertently omitted from manufacturer device report 1220648-2024-10867. d.4 revised model number from manufacturer device report 1220648-2024-10867 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report 1220648-2024-10867 and should not have been. h.6 added codes 925, 4641, and 4644 as they were inadvertently left off manufacturer device report 1220648-2024-10867. h.6 code 4114 was reported incorrectly on manufacturer device report 1220648-2024-10867. a new code has been added to type of investigation codes. h.10 additional manufacturer narrative instructions for use (ifus) were reported on manufacturer device report 1220648-2024-10867 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. additional manufacturer narrative on manufacturer device report 1220648-2024-10867 stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded.
27007 | 1220648-2024-10868 | 19262525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19262525&pc=OZD | 02/26/2022 | 05/07/2024 |  | 05/07/2024 |  |  | 02/26/2022 | 05/07/2024 |  |  | 10/31/2023 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154826 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Positioning Failure | Increase in Pressure | [5 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/26/2022 |  |  |  |  | The complainant reported a patient in in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient arrived to the catheterization laboratory decompensating and required rounds of cardiopulmonary resuscitation (cpr) and defibrillation prior to impella placement. after the impella was inserted, the patient continued to require cpr before achieving a return of spontaneous circulation. after the cpr was performed, the impella had moved back into the aorta, and the medical team had to re-advance the impella into the left ventricle. it was reported that while the patient was on impella cp support, a "purge system blocked" alarm on console occurred. the purge system and fluid was changed with the physician with d10w, due to shortage of recommended purge solution of d5w. no kinks were found in the impella catheter. the purge system was within normal ranges afterwards. a recommendation to the physician was made to replace impella pump due to the possibility of pump failure, and the physician declined to remove and replace at that time. the medical team was advised to monitor motor current for signs of motor failure. no further alarms were noted, and the patient continued to be supported on the impella. an external femoral to femoral bypass was performed to the right leg, as the medical team was unable to find a pulse under doppler. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
27008 | 1220648-2024-10869 | 19265729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19265729&pc=OZD | 02/16/2022 | 05/07/2024 |  | 05/07/2024 |  |  | 02/16/2022 | 05/07/2024 |  |  | 11/30/2023 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023169086 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/16/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved by removing the impella, exchanging the sheath, and re-inserting the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27009 | 1220648-2024-10870 | 19265745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19265745&pc=OZD | 02/16/2022 | 05/20/2025 |  | 05/07/2024 |  |  | 02/16/2022 | 05/07/2024 |  |  | 11/30/2022 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2022143169 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support the physician had to reposition the impella and pull it back into the aorta. they were unable to regain access to the ventricle at bedside, so they took the patient to the cath lab where they were also unsuccessful in regaining ventricular access. the decision was made to replace with another impella. it was reported that the patient was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27010 | 1220648-2024-10871 | 19265786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19265786&pc=OZD | 02/25/2022 | 05/20/2025 |  | 05/07/2024 |  |  | 02/25/2022 | 05/07/2024 |  |  | 09/30/2023 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022137550 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 02/25/2022 |  |  |  |  | The user facility reported a 72-year-old patient was implanted with an impella cp support due to high-risk percutaneous coronary intervention (hrpci). upon explanting the impella cp, it was reported the physician experienced difficulty with the closing. the first perclose device did not work. the physician continued to attempt to use three other perclose devices. the last and fifth attempt appeared to have taken and allowed the impella cp site to be closed. one angio-seal was also used during the closure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27011 | 1220648-2024-10872 | 19265795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19265795&pc=OZD | 02/25/2022 | 05/07/2024 |  | 05/07/2024 |  |  | 02/25/2022 | 05/07/2024 |  |  | 09/30/2023 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142964 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 02/25/2022 |  |  |  |  | The user facility reported a 68-year-old male patient was implanted with an impella cp for mechanical circulatory support as a high-risk percutaneous coronary intervention. while on support, the patient experienced access site bleeding. to resolve the issue an extra perclose suture, two units of blood, and deployment of a covered stent was required. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27012 | 1220648-2024-10873 | 19265857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19265857&pc=OZD | 02/25/2022 | 05/07/2024 |  | 05/07/2024 |  |  | 02/25/2022 | 05/07/2024 |  |  | 10/31/2023 |  | 05/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156972 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 95 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 02/25/2022 |  |  |  |  | The user facility reported a 95-year-old patient was implanted with an impella cp for mechanical circulatory support. while on support the patient was reported to have a small ventricle space and had issues with ectopic beats. the impella was weaned from the patient and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27013 | 1220648-2024-10762 | 19246950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19246950&pc=OZD | 08/09/2022 | 03/07/2025 |  | 05/06/2024 |  |  | 08/09/2022 | 05/06/2024 |  |  | 02/29/2024 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186458 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2022 |  |  |  |  | The user facility reported a 73-year-old male on impella cp for mechanical circulatory support. it was reported that while on impella cp support, there was bleeding in the puncture point that could be stopped. however, intervention was unspecified. no other information provided. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27014 | 1220648-2024-10763 | 19246957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19246957&pc=OZD | 03/08/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 08/23/2022 | 05/06/2024 |  |  | 10/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2022144038 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/23/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. bleeding and hematoma were observed at the access site, and a blood transfusion was provided to the patient to remedy this issue. cardiac function was recovered, and the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27015 | 1220648-2024-10764 | 19246964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19246964&pc=OZD | 09/07/2022 | 05/30/2025 |  | 05/06/2024 |  |  | 09/07/2022 | 05/06/2024 |  |  | 06/30/2024 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235740 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported a 70-year-old male on impella cp for for hemodynamic support. while on support ooze from the site with no hematoma was noted. heparin in the purge was decreased to 25 iu/ml and eventually paused. delay in ventricular assist device due to heparin-induced thrombocytopenia which was also noted. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombocytopenia and access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
27016 | 1220648-2024-10765 | 19246971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19246971&pc=OZD | 09/10/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 09/10/2022 | 05/06/2024 |  |  | 05/31/2024 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230909 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/10/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was then found that the patient had an activated clotting time over 250 with oozing at the access site. systemic heparin was stopped, and the patient remained on impella cp support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27017 | 1220648-2024-10766 | 19246985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19246985&pc=OZD | 09/13/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 09/13/2022 | 05/06/2024 |  |  | 11/30/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023160173 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2022 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved with redressing, review of angle match / groin management best practice, and withholding of heparin. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartasssist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27018 | 1220648-2024-10767 | 19247011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19247011&pc=OZD | 10/31/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 10/31/2022 | 05/06/2024 |  |  | 02/29/2024 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191040 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | Yes | Pumping Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/31/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after implant, it was found that there had been "many suctions" after the patient returned to their room. the patient then also experienced hemolysis, which was tolerable and improved as the right heart function improved. the impella was then removed early due to the improving function of the heart, but it is unknown if the hemolysis was a factor in the early removal of the pump. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27019 | 1220648-2024-10768 | 19247051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19247051&pc=OZD | 03/09/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 03/09/2022 | 05/06/2024 |  |  | 10/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157899 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/09/2022 |  |  |  |  | The complainant reported a patient, of unknown race, in cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, the patient experienced bleeding from the impella access site. the physician made the decision to explant the impella and change the sheath for the 14fx25cm abiomed peel away sheath. the peel away sheath valve did not provide hemostasis and the physician chose to insert a 14fx30 cook sheath. the impella was flushed and re-wired and implanted with the repositioning sheath inserted into the cook. the impella was restarted in p2 and the groin was dressed and angle maintained. the groin site began bleeding again, and hemostasis was achieved with a purse string suture. the groin began bleeding again during support and systemic and purge heparin was discontinued, and new sutures were placed. hemostasis was achieved, and no further bleeding issues were noted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27020 | 1220648-2024-10769 | 19247139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19247139&pc=OZD | 03/11/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 03/11/2022 | 05/06/2024 |  |  | 11/30/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157147 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2022 |  |  |  |  | The complainant reported a patient, race and ethnicity unknown, noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, a heparin injection was administered prior to explanting the impella, despite being warned against injecting the heparin. bleeding occurred after patient movement and required closure with proglide and angio-seal. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27021 | 1220648-2024-10770 | 19247251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19247251&pc=OZD | 04/24/2022 | 05/14/2025 |  | 05/06/2024 |  |  | 04/24/2022 | 05/06/2024 |  |  | 12/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181480 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient then presented with a large hematoma at the groin access site. manual pressure was held, and epinephrine was given. the impella was removed due to the access site bleed, and the patient survived the procedure. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27022 | 1220648-2024-10771 | 19247261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19247261&pc=OZD | 03/11/2022 | 05/14/2025 |  | 05/06/2024 |  |  | 03/11/2022 | 05/06/2024 |  |  | 11/30/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161910 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site groin bleeding and hematoma at the end of the case. the bleeding was resolved with manual pressure and tightening the perclose sutures. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27023 | 1220648-2024-10774 | 19247486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19247486&pc=OZD | 03/11/2022 | 05/06/2024 |  | 05/06/2024 |  |  | 03/11/2022 | 05/06/2024 |  |  | 11/30/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023160163 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/11/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia and an external bypass was performed to resolve, however the foot remained cold. the vascular team decided to removed the impella cp because of the limb ischemia. the patient was weaned off the impella cp and the impella was successfully explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states). as noted in the impella ifu ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
27024 | 1220648-2024-10775 | 19247508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19247508&pc=OZD | 04/29/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 04/29/2022 | 05/06/2024 |  |  | 12/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176103 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was later found to have their right leg cooler than the left, and there were no doppler pulses on either. within an hour, there were doppler pulses to the left posterior tibial, and after two hours, doppler pedal pulses on the left. right lower extremity doppler pulses were still absent. the decision was made to explant the impella cp, and perclose was used to seal the access site. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
27025 | 1220648-2024-10776 | 19247726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19247726&pc=OZD | 03/11/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 03/11/2022 | 05/06/2024 |  |  | 08/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022133730 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient exhibited some ventricular tachycardia. the patient was treated with amiodarone and lidocaine. the patient was successfully weaned and the impella was explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | To make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant). ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records.
27026 | 1220648-2024-10777 | 19247797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19247797&pc=OZD | 03/03/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 03/03/2022 | 05/06/2024 |  |  | 11/30/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162647 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, a large hematoma and bleeding were noted when the peel away was removed and the repo sheath was inserted. additionally, there was no distal flow to the leg, and the patient was beginning to show signs of ischemia. due to the hematoma, an antegrade sheath could not be placed for external femorofemoral bypass. the impella pump was explanted once the patient was hemodynamically stable. manual pressure was applied, balloon tamponade from the contralateral side provided satisfactory hemostasis and restored flow to the limb. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
27027 | 1220648-2024-10778 | 19247975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19247975&pc=OZD | 03/18/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 03/18/2022 | 05/06/2024 |  |  | 12/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176441 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was then found that the patient was experiencing severe hemolysis and their acute kidney injury (aki) was worsening. the physician decided to explant due to the severity of the hemolysis. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27028 | 1220648-2024-10779 | 19248022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248022&pc=OZD | 03/21/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 03/21/2022 | 05/06/2024 |  |  | 01/31/2024 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174094 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | Appropriate Term / Code Not Available | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during placement, the pump was kinked in the right femoral artery, so the pump was switched to the left femoral artery. there was no reported injury to the patient. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27029 | 1220648-2024-10780 | 19248071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248071&pc=OZD | 04/06/2024 | 06/04/2024 |  | 05/06/2024 |  |  | 04/06/2024 | 05/06/2024 |  |  |  |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025464667 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | VENTILATOR | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2024 |  |  |  |  | The complainant reported a patient (race unknown) post cardiotomy cardiogenic shock / low cardiac output syndrome was urgently implanted with an impella 5.5 for mechanical circulatory support post coronary artery bypass surgery. the team was unable to stabilize the patient and decision was made to place the impella 5.5 for left ventricle decompression and to discontinue intra-aortic balloon pump. the impella 5.5 pump experienced high pressure purge during support. tissue plasminogen activator was run through the purge to manage the noted issue and support with the implanted pump was maintained. ultimately, care was withdrawn and the patient expired. the outcome is not related to the use of device. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the reported high purge pressure has been completed. the pump was returned for evaluation. upon visual inspection, biomaterial was found in the purge gap but not completely covering gap. the pump was connected to the automated impella controller (aic) and purged to attempt to reproduce high purge pressure alarms. high purge pressure was not reproduced, pump was turned on to p-9 and ran for approximately five minutes with no high purge pressure alarms occurring. no data logs were returned for evaluation. clinical details noted that patient was receiving multiple blood products and was not on any forms of anti-coagulation. additionally, tissue plasminogen activator was added to the purge after alarms in attempt to resolve high purge pressure. the root cause of the high purge pressure is most likely due to biomaterial. instructions for use for the related event are as follows: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
27030 | 1220648-2024-10782 | 19248126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248126&pc=OZD | 03/24/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 03/24/2022 | 05/06/2024 |  |  | 11/30/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167697 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | Appropriate Term / Code Not Available | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/24/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, an echocardiogram showed that the pump inlet was 1cm across the valve. the physician advanced the pump across the valve, and the pump became stuck in the papillary muscle. the patient continued to get frequent suction above p6 and remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27031 | 1220648-2024-10781 | 19248149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248149&pc=OZD | 04/06/2024 | 05/20/2024 |  | 05/06/2024 |  |  | 04/06/2024 | 05/06/2024 |  |  | 09/30/2025 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025444896 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. one day post impella pump placement, the patient experienced limb ischemia and was treated with a radial bypass to the femoral. two days later the patient status was noted as critical and was made a dnr by the family. patient during blood transfusion decompensated and went into a complete heart block where the patient did not recover. the patient expired two days afterwards after the impella pump had supported for just under six days. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system-section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. d6a.
27032 | 1220648-2024-10783 | 19248171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248171&pc=OZD | 05/09/2022 | 05/22/2025 |  | 05/06/2024 |  |  | 05/09/2022 | 05/06/2024 |  |  | 05/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022082313 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2022 |  |  |  |  | It was reported that a 51 year old male patient was placed on an impella cp for mechanical circulatory support. while on support, the patient developed bleeding at the impella site (left groin) due to a subcutaneous cut. a small retroperitoneal hematoma was seen and hemoglobin (hb) was decreased from 12 to 8. blood products were ordered, volume was given and the patient was stable. additionally, an abdominal ultrasound was performed and an acute retroperitoneal hemorrhage was found. the decision was made to explant the pump and close the groin. | [5 empty] | Impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27033 | 1220648-2024-10784 | 19248175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248175&pc=OZD | 04/12/2024 | 05/29/2024 |  | 05/06/2024 |  |  | 04/12/2024 | 05/06/2024 |  |  | 11/30/2025 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024434486 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that resistance was encountered during attempted delivery of an impella 5.5 pump. it was noted that the device was unable to advance through the artery due to the clavicle. after multiple unsuccessful attempts, the decision was made to abort the 5.5 implant. an impella cp pump was subsequently placed for patient support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the delivery issue was most likely patient condition since the pump would not pass through the artery due to the clavicle. section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ in accordance with updated procedures, section h10 has been revised from the initial submission.
27034 | 1220648-2024-10785 | 19248204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248204&pc=OZD | 04/22/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 04/22/2022 | 05/06/2024 |  |  | 11/30/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166098 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/22/2022 |  |  |  |  | The user facility reported an 88-year-old male on impella cp for mechanical circulatory support. bleeding was reported after peel away was removed. they attempted to tighten perclose, applied light manual pressure above insertion site, and provided a figure 8 suture. ultimately decision was made to explant and reassess right femoral for another impella cp placement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27035 | 1220648-2024-10786 | 19248239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248239&pc=OZD | 04/08/2024 | 06/19/2024 |  | 05/06/2024 |  |  | 04/15/2024 | 05/06/2024 |  |  |  |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450063 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2024 |  |  |  |  | Adverse event report, protect iv study. after hospital discharge: patient found dead in his bed. impella explant date (b)(6) 2024, date of death (b)(6) 2024; pi relatedness assessment impella procedure / death: possible; abiomed safety relatedness assessment: remote (unlikely). | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported death has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of death could not be determined as the device was not returned nor sufficient clinical details. the instructions for use referenced in the initial report is for the impella cp with smart assist system in section: potential adverse events (united states).
27036 | 1220648-2024-10788 | 19248338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248338&pc=OZD | 04/18/2024 | 07/05/2024 |  | 05/06/2024 |  |  | 04/18/2024 | 05/06/2024 |  |  | 10/31/2025 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024413431 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp support, the automated impella controller (aic) displayed optical sensor failure. the console was exchanged but the issue did not resolve. therefore, the pump was explanted and a new pump was placed to resolve the issue. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported optical signal issue has been completed. the product was returned and a light leak was found in the red handle on the patient cable side with no signs of excessive force. the console displayed optical sensor failure alarm after implant. data logs indicated that the placement signal spiked to 3000 during the case, and the optical sensor 5s parameters were characteristic of an out-of-box optical fiber break. the root cause of the optical signal issue was determined to be a damaged optical fiber line.
27037 | 1220648-2024-10789 | 19248343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248343&pc=OZD | 04/23/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 04/23/2022 | 05/06/2024 |  |  | 11/30/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162909 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2022 |  |  |  |  | The user facility reported a 67-year-old female on impella cp for mechanical circulatory support. it was reported that while on impella support, the patient experienced a loss of distal pulses. it was noted that a decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
27038 | 1220648-2024-10790 | 19248364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248364&pc=OZD | 03/08/2022 | 05/06/2024 |  | 05/06/2024 |  |  | 03/08/2022 | 05/06/2024 |  |  | 11/30/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169869 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/08/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced frequent suction alarms due to low pump flow. the clinician suggested a fluid bonus, however the physician chose to hydrate orally versus increasing the pump flow. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27039 | 1220648-2024-10748 | 19248429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248429&pc=OZD | 11/08/2023 | 05/21/2024 |  | 05/06/2024 |  |  | 04/05/2024 | 05/06/2024 |  |  | 08/31/2025 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397991 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 72 YR |  | Unknown |  |  |  | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/05/2024 |  |  |  |  | Information received via the loqi ( abiomed¿s long-term outcome and quality indicator) impella registry reported in (b)(6) of 2024, an adverse event attributed to the use of the impella cp. the patient of unknown (at this time) race and gender suffered from ventricular tachycardia. the patient survived the event. the relationship s listed as "possibly" and further clinical details are being pursued. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. the device information and the physician(reporter) name is not available at this time, additional report will be filed when informations is available. | Additional device information received. section d4 updated as per information. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. the instructions for use for impella cp with smartassist system in the in-hospital adverse events table 6.31 lists ventricular arrhythmia as a major complication, where only a fraction are attributed to the impella and events resolved without residual effects. added-a1-patient identifier, e1 name prefix, first name, last name. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27040 | 1220648-2024-10791 | 19248477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248477&pc=OZD | 03/14/2024 | 07/05/2024 |  | 05/06/2024 |  |  | 04/18/2024 | 05/06/2024 |  |  | 03/31/2024 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192044 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported a 87-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that 87-year-old male was on impella cp support for high risk percutaneous coronary intervention (hrpci). a few hours after a problem-free procedure and explant, the patient experienced bleeding in the groin. the bleeding was repaired with coiling and there was no further patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is for the impella cp with smart assist system section: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation.¿, section: general patient care considerations, and section: entering cardiac output. added- h6 component code 925. g1 manufacturing site zip code should be 5207421227 (received error message stating incorrect format for us postal zip code when trying to make the correction.) d4-corrected model number f6/f8-should have been left blank g1-corrected reporting contact email g1-corrected manufacturing site address, city, and fax number
27041 | 1220648-2024-10792 | 19248506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248506&pc=OZD | 03/08/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 03/08/2022 | 05/06/2024 |  |  | 12/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173618 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on impella support, the patient experienced wire induced arrhythmia and was successfully cardioverted. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | To make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records.
27042 | 1220648-2024-10793 | 19248535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248535&pc=OZD | 03/19/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 03/19/2022 | 05/06/2024 |  |  | 08/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022125333 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient was bleeding from the arterial impella access site. two units of packed red blood cells (prbc) and platelets were ordered and transfused. heparin in the purge fluid was decreased and systemic heparin was discontinued. it was noted that to resolve the issue a stitch was placed, and hemostasis was achieved via manual pressure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27043 | 1220648-2024-10794 | 19248545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248545&pc=OZD | 04/18/2024 | 06/25/2024 |  | 05/06/2024 |  |  | 04/18/2024 | 05/06/2024 |  |  |  |  | 05/06/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1647627 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Activation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during transport to explant, when the automated impella controller (aic) was disconnected from the ac power, a controller failure alarm appeared. after restart, the screen froze. the device was not removed due to the failure mode. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. the console logs from the reported day of event are consistent with the complaint and show that 17 seconds after the console presented with an advisory notification ¿ac power disconnected¿, all serial communication between epc and automated impella controller peripherals started failing, and corresponding processes were restarting in an attempt to recover communication. log data indicated that during this time, the pump was stopped for about seven seconds, and optical bench did not provide placement data for about 30 seconds (which is also when all communication issues resolved). after pump restart, there were still some unusual entries in the logs that could not be interpreted. after another 30 minutes, the pump and purge were disconnected from the console, and five minutes later battery failure alarm was set and resolved. this was presumably due to the staff disengaging the batteries in an attempt to shut down the console that may have been at this point unresponsive. a few seconds later, a user-requested shutdown was executed, and the console was promptly powered down. throughout the duration of the case, there were many interruptions of data streaming. each time recovering after the first and second handshake attempt. corresponding journals showed multiple unexpected reboots coinciding with the moments of "bad file descriptor" in the logs. there was no correlation noted between the reboots and wifi signal strength or access point connection. the console was tested but none of the issues could be reproduced. no anomalies were observed. the console components that potentially could have contributed to the issue are the 4-in-1 assembly as a communication hub and assembly as power source to all components. d.4 revised model number from the initial submission in accordance with updated procedures. g.1 revised manufacturing site fax number from the initial submission in accordance with updated procedures. h.6 codes 3270 and 403 were reported incorrectly on the previously submitted report. a new code has been added to medical device problem codes.
27044 | 1220648-2024-10795 | 19248561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248561&pc=OZD | 04/24/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 04/24/2022 | 05/06/2024 |  |  | 11/30/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165953 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2022 |  |  |  |  | The user facility reported a 75-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced significant bleeding at the groin insertion site. the physician adjusted the angle and redressed the site. the next day the patient was bleeding profusely from the groin/mouth/every stick. dextrose was administered and heparin was withheld until the bleeding subsided. two units of packed red blood cells were administered to the patient. the patient was reportedly hemodynamically stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27045 | 1220648-2024-10796 | 19248620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248620&pc=OZD | 03/08/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 03/08/2022 | 05/06/2024 |  |  | 10/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022150779 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support. during implantation, it was noted that the patient was found to have calcified iliacs, attempted to balloon and stent without success, therefore, the impella implantation aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27046 | 1220648-2024-10797 | 19248720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248720&pc=OZD | 04/29/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 04/29/2022 | 05/06/2024 |  |  | 01/31/2024 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188143 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2022 |  |  |  |  | The user facility reported a 78-year-old male was implanted with an impella cp for mechanical circulatory support. it was reported that the patient had access site bleeding around the impella sheath. the impella cp and sheath were removed and exchanged for 25cm sheath. the impella cp was re-inserted. hemodynamics were obtained. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27047 | 1220648-2024-10799 | 19248738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248738&pc=OZD | 03/08/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 03/08/2022 | 05/06/2024 |  |  | 10/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152040 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the patient was sent to the unit on impella support with plans to explant the next morning, however there were no distal pulses, so the decision was made to remove the impella. patient tolerated the procedure well and remained stable after explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
27048 | 1220648-2024-10798 | 19248739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248739&pc=OZD | 04/30/2024 | 02/25/2025 |  | 05/06/2024 |  |  | 04/30/2024 | 05/06/2024 |  |  |  |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED,INC |  |  | US |  | 1393502 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/30/2024 |  |  |  |  | The complainant in japan had a successful cp support for 3 days when the pump was explanted. the team then observed a controller failure and with restart the failure was reproduced. the team has removed the automated impella controller (aic) from service and per the instructions for use (ifu) a backup controller will be used. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) controller issues has been completed. data logs were reviewed which showed that the purge pressure sensor defective red controller error alarm was issued on date of occurrence. this log entry was preceded by a purge: sensor defective (voltage out of range) entry indicating that power to the purge pressure sensor through the purge unit driver was compromised. the console evaluated by field service which they replace the purge unit driver and performed preventative maintenance procedure as a troubleshooting step. the console was found to be functioning to specification after troubleshoot. the root cause for the purge pressure sensor defective alarm controller failure was a defective purge unit driver.
27049 | 1220648-2024-10800 | 19248770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248770&pc=OZD | 04/27/2024 | 07/02/2024 |  | 05/06/2024 |  |  | 04/27/2024 | 05/06/2024 |  |  |  |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471090 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Malposition of Device | Material Split, Cut or Torn | [5 empty] | Tachycardia | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2024 |  |  |  |  | The user facility reported a 73-year-old male of unknown race noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while on support, the patient had ventricular tachycardia (vt). the pump was found to be in the aorta, and it was repositioned. additionally, the patient had low platelets. the patient received one unit of blood products. furthermore, the sterile sleeve tore at the proximal end near the lock. the staff was repositioning the pump and there was not any slack given to the sterile sleeve and it ripped. moreover, the pump was found in the aorta so, the patient was brought to the operating room. the pump was explanted. the patient was reported as stable. | It was reported that the patient expired after care was withdrawn. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the product damage, thrombocytopenia, positioning issue, and the ventricular arrhythmia has been completed. no product was returned for investigation. during repositioning the clinical mentions that no slack was given causing the sterile sleeve to rip. the root cause of the product damage sterile sleeve damage was most likely use related. the root cause of the positioning issues was not determined since product was not returned and there is a lack of clinical details. the root cause of the thrombocytopenia was not determined due to insufficient clinical details. the root cause of the ventricular tachycardia could not be determined without additional clinical details. added-b5 mso review f6/f8 should have been left blank in the initial report.
27050 | 1220648-2024-10801 | 19248812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248812&pc=OZD | 05/01/2024 | 06/21/2024 |  | 05/06/2024 |  |  | 05/01/2024 | 05/06/2024 |  |  | 11/30/2025 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024436260 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2024 |  |  |  |  | The complainant reported a 74-year old female patient with acute myocardial infarction / cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the cp failed delivery and could not be implanted due to patient anatomy, so medical staff decided to abort the procedure. medical staff placed an intra-aortic balloon pump instead and consulted with cardiothoracic surgery. no patient harm has been reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of delivery issue was most likely patient condition since severe tortuosity was noted in the descending aorta. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. g.1 manufacturing site address line 1 revised as previously entered incorrectly. h.6 codes 4445, 463, 3038 and 4115 were reported incorrectly on the previously submitted report. new codes have been added. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
27051 | 1220648-2024-10802 | 19248839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248839&pc=OZD | 04/21/2024 | 07/11/2024 |  | 05/06/2024 |  |  | 04/21/2024 | 05/06/2024 |  |  |  |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446865 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | PHYSICIAN ASSISTANT | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/21/2024 |  |  |  |  | The user facility reported a male patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. md reported that the pump had stopped and could no longer be started; several attempts to start the pump failed, md reports that the patient could not be sufficiently anticoagulated since implant because of a hemoragic shock; a thrombus is possible; pp 550mmhg; even with a aic change to a new controller did not lead to the expected result. recommendation given to change the pump; in the meantime, the circuit was guaranteed with the ecmo (4l). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported pump stop has been completed. no product was returned for evaluation. no data logs were returned for evaluation. clinical details noted that pump was running for 22 days, and pump was able to be manually restarted after pump stops. patient was not on proper anticoagulation management techniques due to hemorrhagic shock. the root cause of the pump stop cannot be determined as no logs or product were returned for evaluation. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high risk cpi in the following sections: section: using the automated impella controller with the impella catheter section: using the automated impella controller with the impella rp with smartassist system catheter added h6 impact code 4629 d4-corrected model number f6/f8-should have been left blank in the initial report.
27052 | 1220648-2024-10803 | 19249097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19249097&pc=PYX | 03/04/2020 | 06/13/2024 |  | 05/06/2024 |  |  | 04/23/2024 | 05/06/2024 |  |  | 06/30/2020 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1353967 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 51 YR |  | Female |  |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombocytopenia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A publication was received indicating that a patient was implanted with an impella rp for support and that a thrombectomy that was performed. additionally, it was noted the patient had heparin induced thrombocytopenia (hit) while on support, and the purge solution was changed to angiomax. furthermore, the pump had suction when it was running above p-level p-7. the staff wanted to run at p-8, and it was suggested to give volume to the patient. | [5 empty] | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the low flow and the thrombocytopenia was unable to be determined as limited clinical details were provided and no product was returned for review. as the necessary information was not provided, the root cause of the thrombus was not determined.
27053 | 1220648-2024-10804 | 19249369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19249369&pc=OZD | 04/25/2024 | 06/13/2024 |  | 05/06/2024 |  |  | 04/25/2024 | 05/06/2024 |  |  |  |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440635 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2024 |  |  |  |  | The user facility reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. during insertion of the impella cp pump, the pigtail portion of the device perforated through the pericardium. the perforation was visual on the echocardiogram and fluoroscopy. the patient was taken to the operating room for open repair of the left ventricle and removal of the catheter. an intra aortic balloon pump was inserted for ongoing support and the pump was discarded. the patient is on bed rest. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system for use section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ | The investigation into the ventricle perforation issue has been completed since the original report was submitted. the device was not returned for investigation. per the clinical information and data log review, the root cause of the ventricle perforation issue was most likely improper positioning of the device.
27054 | 1220648-2024-10805 | 19249468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19249468&pc=OZD | 02/27/2022 | 05/20/2025 |  | 05/06/2024 |  |  | 02/27/2022 | 05/06/2024 |  |  | 08/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022125114 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | Yes | Increase in Pressure | Failure to Deliver | [5 empty] | Atrial Fibrillation | Tachycardia | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2022 |  |  |  |  | The user facility reported a 67-year-old patient needing mechanical circulatory support. it was reported that the patient was on impella cp for hemodynamic support. while on support, the initial impella femoral access attempt aborted due to severe peripheral vessel disease and unsuitable target for reperfusor sheath, higher purge flow rate and lower purge pressure, rapid afib (atrial fibrillation), episodes of vt (ventricular tachycardia)/v-fib (ventricular fibrillation) requiring defibrillation, and unsuccessful perclose attempts was noted. purge cassette was changed, no leaks observed, and flo seal applied to tract. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27055 | 1220648-2024-10807 | 19249545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19249545&pc=OZD | 04/27/2024 | 06/06/2024 |  | 05/06/2024 |  |  | 04/27/2024 | 05/06/2024 |  |  |  |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447009 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2024 |  |  |  |  | The user facility reported a patient undergoing coronary artery bypass was implanted with an impella cp device for mechanical circulatory support. during support, the patient on impella cp support was brought back to catheterization lab during the night due to leg pain and possible ischemia. external fem to fem bypass was placed for distal leg perfusion. it was further reported that the impella cp was weaned and explanted in the catheterization lab due to the distal leg perfusion being compromised. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
27056 | 1220648-2024-10806 | 19249546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19249546&pc=OZD | 02/28/2022 | 05/06/2024 |  | 05/06/2024 |  |  | 02/28/2022 | 05/06/2024 |  |  | 10/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152179 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/28/2022 |  |  |  |  | The user facility reported a 46-year-old patient needing mechanical circulatory support. it was reported that the patient was on impella cp support and while attempting to place the patient on support, the 14fx13cm abiomed sheath was inserted over a standard 0.035 table wire, which proved to be difficult, and the wire kinked. it was also reported that the patient experienced access site bleeding. the surgeon removed the 14fx13cm and replaced with the 14fx25cm abiomed sheath over the wire. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27057 | 1220648-2024-10808 | 19249666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19249666&pc=OZD | 03/08/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 03/08/2022 | 05/06/2024 |  |  | 09/30/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022146418 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, no flow was identified through re-access sheath. an external fem-fem bypass was done to perfuse down the right femoral artery. additionally, suction alarms were persistent on p3. the patient was successfully weaned off support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27058 | 1220648-2024-10809 | 19249766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19249766&pc=OZD | 04/08/2024 | 11/07/2025 |  | 05/06/2024 |  |  | 04/08/2024 | 05/06/2024 |  |  |  |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1132457 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 04/08/2024 |  |  |  |  | The complainant also reported that an aorta and left ventricle placement signal were lost during the alarms. | The user facility reported a 42-year-old male patient of unknown ethnicity with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the automated impella controller (aic) had a controller error alarm which occurred twice. both alarms resolved on their own. the patient remained hemodynamically stable. the aic was changed after the second alarm occurrence. no patient harm was reported. |  |  |  |  | The investigation into the automated impella controller error has been completed. data analysis: data logs confirm two instances of the controller error alarm. device analysis: the console was returned for evaluation. the controller error was due to an optical bench firmware communication protocol deficiency resulting in in misalignment of data communication packets received by embedded personal computer. the console software operated as designed. d8 added device return date d6a should have been blank-this device is not implanted f6 and f8 should have been blank. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: imc logs confirmed the reported failure mode given there were two instances of a controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. i.e. [07/04/24 22:35:08 imcgui: event set: #1045 from 9 (0)]. [08/04/24 04:44:35 imcgui: event set: #1045 from 9 (1)]. real time (rt) logs confirmed that signals received from optical bench (snr, contrast, gain) are represented as -16384 values due to the crc error. device analysis: console (ic2350) was sent back fs for further investigation. the console's front optical connector was inspected before cleaning and after cleaning. section 23 of the preventative maintenance was performed and the specifications met. the consoles fringe pattern revealed no observable abnormality. conclusion: the controller error and loss of placement signal is due to an optical bench fw communication protocol deficiency, resulting in misalignment of data communication packets received by epc. console sw operated as designed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal description. h6 updated. d2 common device name updated. g1 reporting contact fax number updated. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27059 | 1220648-2024-10810 | 19251211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19251211&pc=OZD | 04/28/2024 | 06/11/2024 |  | 05/06/2024 |  |  | 04/28/2024 | 05/06/2024 |  |  |  |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025467741 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/28/2024 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed a hematoma. the hematoma was managed by pressure dressing and blood was given. the bleeding stopped. the patient is stable. the patient remained on impella cp support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission
27060 | 1220648-2024-10811 | 19251422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19251422&pc=OZD | 03/03/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 03/03/2022 | 05/06/2024 |  |  | 12/31/2021 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023175392 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2022 |  |  |  |  | The complainant reported a patient noted as high-risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. during implant, a small hematoma and a tract ooze was encountered once the peel away sheath was placed, which was treated through manual pressure. during explant and peel away sheath removal, the perclose sutures were also removed, and oozing was noted which then required vascular surgery to close the access site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27061 | 1220648-2024-10812 | 19251798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19251798&pc=OZD | 04/02/2024 | 05/16/2024 |  | 05/06/2024 |  |  | 04/02/2024 | 05/06/2024 |  |  |  |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024435790 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported a 71-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that while on support, there was high purge pressure. the purge cassette was changed and tissue plasminogen activator was added to the purge. the patient remained on support until bridged left ventricular assist device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the high purge pressure was not determined as no product, no data logs, and insufficient clinical details were provided. in accordance with updated procedures, section d6 has been revised from the initial submission.
27062 | 1220648-2024-10813 | 19253321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19253321&pc=OZD | 04/27/2024 | 06/04/2024 |  | 05/06/2024 |  |  | 04/30/2024 | 05/06/2024 |  |  |  |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454441 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient experienced a stroke. the patient experienced left side deficits as a result of the stroke. support with the implanted pump was maintained following diagnosis. a couple of days later, low flow alarms appeared, and the decision was made to explant the pump. | [5 empty] | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into the low pump flow and the stroke/cva issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the low pump flow is most likely biomaterial based off the data log analysis and clinical information. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.
27063 | 1220648-2024-10814 | 19253518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19253518&pc=OZD | 03/02/2022 | 05/20/2025 |  | 05/06/2024 |  |  | 03/02/2022 | 05/06/2024 |  |  | 11/30/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161637 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 32 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support the purge pressure increase alarm occurred. when the position was adjusted, the pressure was improving. it was noted that pressure dropped and the following morning, it was operating about 400 mmhg. due to the heparin induced thrombocytopenia (hit) antibody testing positive, the patient was given argatroban. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27064 | 1220648-2024-10815 | 19255202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19255202&pc=OZD | 03/03/2022 | 05/19/2025 |  | 05/06/2024 |  |  | 03/03/2022 | 05/06/2024 |  |  | 09/30/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022137932 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the physician experienced positioning issues as the pump was pulled back into the descending aorta. an attempt was made to replace the pump across the aortic valve but was unsuccessful. the decision was made to remove the pump and replace it with the intra-aortic balloon pump (iabp). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27065 | 1220648-2024-10816 | 19255326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19255326&pc=OZD | 03/03/2022 | 04/07/2025 |  | 05/06/2024 |  |  | 03/03/2022 | 05/06/2024 |  |  | 08/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022138287 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced limb ischemia and the decision was made to explant the device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27066 | 1220648-2024-10817 | 19256199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19256199&pc=OZD | 12/06/2022 | 05/06/2024 |  | 05/06/2024 |  |  | 12/06/2022 | 05/06/2024 |  |  | 05/31/2024 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223102 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 12/06/2022 |  |  |  |  | The user facility reported a 64-year-old patient needing mechanical circulatory support. it was reported that the patient was placed on impella cp smart assist support for protection of the heart due to a chronic total occlusion at the right coronary artery. the patient developed bleeding/hematoma at the impella insertion site. the patient received a blood transfusion. additionally, the site was angle matched and suture was placed. the patient was successfully weaned from the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27067 | 1220648-2024-10818 | 19256226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19256226&pc=OZD | 01/31/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 01/31/2022 | 05/06/2024 |  |  | 10/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022153477 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had heparin induced thrombocytopenia (hit) and there was going to be no heparin in purge for 10 days. additionally, there was a pump stop and the patient was sent to the cath lab for a replacement pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: thrombocytopenia impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27068 | 1220648-2024-10819 | 19256307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19256307&pc=OZD | 02/25/2022 | 03/26/2025 |  | 05/06/2024 |  |  | 02/25/2022 | 05/06/2024 |  |  | 12/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172190 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced atrial fibrillation (af), bleeding and vascular damage. upon placement of the impella and extracorporeal membrane oxygenation (ecmo), the patient went into af. the treatment of both ecmo and impella was continued to stabilize the patient. the initial plan was to begin weaning the patient from ecmo. the patient had continued complication of bleeding requiring blood products and the cause was found to be due to the bilateral hematomas in the femoral arteries from the ecmo (decannulated) and impella. there was bleeding from nose and mouth which required 4 unit of red blood cells (rbc's). due to the ongoing bleeding, the pump was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27069 | 1220648-2024-10820 | 19256333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19256333&pc=OZD | 02/27/2022 | 05/06/2024 |  | 05/06/2024 |  |  | 02/27/2022 | 05/06/2024 |  |  | 10/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022155900 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 02/27/2022 |  |  |  |  | The user facility reported an 88-year-old patient needing mechanical circulatory support. while on support, it was reported that there was access site bleeding noted by the femoral bleeding. this was treated by giving the patient norepinephrine (levophed) and explanting the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27070 | 1220648-2024-10821 | 19256357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19256357&pc=OZD | 02/27/2022 | 05/06/2024 |  | 05/06/2024 |  |  | 02/27/2022 | 05/06/2024 |  |  | 12/31/2023 |  | 05/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171289 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 02/27/2022 |  |  |  |  | The user facility reported an 86-year-old patient was implanted with an impella cp for mechanical circulatory support. while on support the patient¿s groin site had a small amount of bleeding, which noted to be a skin tract ooze. purge solution was changed to include heparin, one unit of blood was infused, and reported that the area would continue to be monitored. it was later reported that the patient experienced episodes of atrial fibrillation (af) with drops of mean arterial pressure (map). pressure level was increased to 7, and a low dose of levophed was added for map support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27071 | 1220648-2024-10757 | 19245505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19245505&pc=OZD | 04/25/2024 | 07/10/2024 |  | 05/05/2024 |  |  | 04/27/2024 | 05/05/2024 |  |  |  |  | 05/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025462480 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2024 |  |  |  |  | The user facility in japan reported a 52-year-old female post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. it was reported that when impella cp was inserted the cp position was facing the rear wall and when the pump performance level was raised a suction alarm occurred. repositioning the impella was unsuccessful, and cp was replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low flow /suction and the positioning issues has been completed. neither product nor data logs were returned for review. based on clinical description, the root cause of positioning issue was most likely use related. ¿the cause of the low pump flow issue was not determined since product, logs were not returned and due to insufficient clinical information. f6/f8 should have been left blank in the initial report.
27072 | 1220648-2024-10759 | 19245512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19245512&pc=OZD | 04/28/2024 | 07/01/2024 |  | 05/05/2024 |  |  | 04/28/2024 | 05/05/2024 |  |  | 11/30/2025 |  | 05/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427534 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/28/2024 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 pump placed for support of a patient admitted with a history of cardiac comorbidities and mental status change. the patient was sent to the or on the impella and after 5 days the access site had a hematoma. the hematoma was treated by sandbag placement on the chest, 3 units of rbc (red blood cells), and removal of heparin. the pump remains on for support while neuro status checked post paralyzed. | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate the use of the device from the outcome. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is for the impella 5.5 with smart assist for use during cardiogenic shock in section: potential adverse events (united states). added-b5 mso review, added d6b, added h6 component code 925 d4-corrected model number f6/f8 should have been left blank h5-changed to yes g1 corrected reporting contract e-mail
27073 | 1220648-2024-10760 | 19245515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19245515&pc=OZD | 04/27/2024 | 06/21/2024 |  | 05/05/2024 |  |  | 04/29/2024 | 05/05/2024 |  |  |  |  | 05/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025467827 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2024 |  |  |  |  | The user facility reported a 42-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient developed a hematoma and heparin was temporarily paused and was resumed 2 days later. the patient remains supported by implanted 5.5 and is reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the pump was returned for investigation. no biomaterial was found. no catheter kinks were found. the sterile sleeve was detached. optical issue was discovered during investigation. the root cause of the optical signal issue was sensor silicone wear. the root cause of the access site bleeding cannot be determined. d6a, d6b.
27074 | 1220648-2024-10761 | 19245520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19245520&pc=OZD | 04/29/2024 | 06/24/2024 |  | 05/05/2024 |  |  | 04/29/2024 | 05/05/2024 |  |  |  |  | 05/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025467927 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Break | Material Separation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2024 |  |  |  |  | The user facility reported a 45-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella 5.5 sterile sleeve was broken. the site was cleaned with chlorhexidine and sterile sleeve reattached with a tegaderm. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (sterile sleeve damage) issue has been completed since the original report was submitted. the device was not returned for investigation. sterile sleeve was found damaged as per the clinical details. the cause of the sterile sleeve damage issue was not determined since product was not returned and sufficient clinical information was not provided. d6b.
27075 | 1220648-2024-10749 | 19244567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19244567&pc=OZD | 04/26/2024 | 06/03/2024 |  | 05/04/2024 |  |  | 04/26/2024 | 05/04/2024 |  |  |  |  | 05/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485800 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. impella cp insertion site post coronary artery bypass graft (cabg) surgery had bleeding issues. the physician attempted to control bleeding with mattress sutures around impella catheter without success. as the patient was stable post cabg surgery and able to maintain adequate hemodynamics without support, the decision was made to remove the impella pump. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is for the impella cp with smart assist system in sections general patient care considerations, entering cardiac output, and potential adverse events (united states). d4-updated model number.
27076 | 1220648-2024-10750 | 19244571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19244571&pc=PYX | 04/26/2024 | 06/13/2024 |  | 05/04/2024 |  |  | 04/26/2024 | 05/04/2024 |  |  | 03/31/2025 |  | 05/04/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024394992 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2024 |  |  |  |  | The user facility reported a 74-year-old male patient was implanted with an impella rp device for mechanical circulatory support via the right jugular left ventricular assist device placement. it was reported that during the echocardiogram it showed the heart underfilled and cardiac tamponade was noted. the bleed was in the superior vena cava (svc). the rp flex's 23 fr sheath was the presumed cause of the tamponade and bleed at the svc. the team repaired the jugular and replaced the rp with a venous-arterial extra-corporeal membrane oxygenation (va-ecmo). the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation. | The investigation into the vascular damage - peripheral vessel has been completed since the original report was submitted. the device was not returned for investigation. as per the clinical details, it was discovered that there was a bleed in the superior vena cava (svc). this bleed was believed to have resulted from the initial placement of a 23 french (fr) sheath. the medical team repaired it surgically. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27077 | 1220648-2024-10751 | 19244577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19244577&pc=PYX | 04/27/2024 | 09/26/2024 |  | 05/04/2024 |  |  | 04/27/2024 | 05/04/2024 |  |  | 01/31/2025 |  | 05/04/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024306799 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Death | Yes | No | Difficult to Insert | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Arrhythmia | Cardiac Arrest | Tachycardia | Ventricular Fibrillation | Cardiovascular Dysfunction/Insufficiency | [8 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/27/2024 |  |  |  |  | The complainant reported a 75 year old male patient was electively implanted with an impella rp for mechanical circulatory support, but the medical staff was unable to advance beyond the large venous cannula in the inferior vena cava. the patient had just undergone cardiac surgery. medical staff was intending on using a new rp after multiple unsuccessful attempts, however, while waiting for the second rp pump, the patient went into ventricular tachycardia and fibrillation arrest and passed away. it is unknown if the complications with the rp pump contributed to the patient's outcome. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: "potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ ¿do not advance or withdraw the impella rp flex with smart assist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-10751 was provided in sections b2, g3, and h1, and h6. | The investigation into the delivery issues and the vf/vt/af/at (vtach/vfib arrest) has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely use issue since an attempted was made by the staff to pass by another cannula which was placed along the route. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
27078 | 1220648-2024-10752 | 19244580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19244580&pc=OZD | 04/27/2024 | 06/03/2024 |  | 05/04/2024 |  |  | 04/27/2024 | 05/04/2024 |  |  |  |  | 05/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025458015 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and the patient had bleeding from the access site. the patient had bleeding and heparin was withheld. the impella access site was opened and packed to control bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist for use during cardiogenic shock in sections: section: potential adverse events (united states), which now includes a statement "cardiac or vascular injury (including ventricular perforation.¿ and general patient care considerations and entering cardiac output sections. d4-updated model number.
27079 | 1220648-2024-10753 | 19244586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19244586&pc=OZD | 08/14/2023 | 06/11/2024 |  | 05/04/2024 |  |  | 04/09/2024 | 05/04/2024 |  |  | 03/31/2025 |  | 05/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347896 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | We received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that hemolysis occurred with a definite relationship to the impella cp used on this patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ "patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis has been completed. product was not returned for investigation. data logs were returned and reviewed. there were "impella position unknown" and "impella in aorta" alarms observed just a few hours before hemolysis was reported but no other supporting information about imaging, position, etc has been provided. therefore, the root cause of the hemolysis issue was not determined since product was not returned and insufficient clinical information provided. the instructions for use referenced in the initial report is for the impella cp with smartassist system in the potential adverse events (united states), use of echocardiography for positioning of the impella catheter, and hemolysis sections. d4-updated model number g1-corrected mfr site name and address.
27080 | 1220648-2024-10754 | 19244791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19244791&pc=OZD | 04/27/2024 | 06/03/2024 |  | 05/04/2024 |  |  | 04/27/2024 | 05/04/2024 |  |  |  |  | 05/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466326 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility a patient in acute myocardial infarction/cardiogenic shock on impella cp support was transferred to another facility with impella p-level of p-9 and urine appearing dark red. labs came back hemolyzed. the patient had high afterload and was dependent on high flow from the impella pump. it was further reported that the patient also had signs of right leg ischemia despite having a 6fr. antegrade connected to 14fr. peelaway sideport. the right leg was cooler, appeared pale, and no pulses could be found. however, the patient had linear blood pressure. following successful placement of an impella 5.5 device, the impella cp was weaned and explanted by vascular surgery with primary closure of artery. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported hemolysis has been completed. the device was not returned for evaluation. the data logs were reviewed. there were no motor current spikes, suction events, position alarms or a trend in the placement signal. due to insufficient information and lack of product returned, the root cause of the hemolysis cannot be determined. the root cause of the limb ischemia could not be determined. d4-updated model number
27081 | 1220648-2024-10755 | 19244808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19244808&pc=PYX | 04/27/2024 | 06/21/2024 |  | 05/04/2024 |  |  | 04/27/2024 | 05/04/2024 |  |  |  |  | 05/04/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025478757 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | Restricted Flow rate | Increase in Pressure |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient in postcardiotomy cardiogenic shock had an impella rp flex placed for bipella support along with the impella cp. there was suspected clot ingesting in the impella rp pump, and there was a spike in motor current followed observed low flows. simultaneously, purge pressure high and purge block alarms presented. stat chest xray showed the impella rp pump was in suboptimal position and had possibly pulled back where the outflow was at the level of the pulmonic valve or possibly in the right ventricle. the impella rp flex was successfully repositioned; however, this did not resolve the alarms or low flows. the decision was made to explant the impella rp pump. the impella rp pump was explanted without issues and hemostasis obtained at right internal jugular by purse string suture. impella cp therapy continued without incident. | [5 empty] | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation for the reported low pump flow, high purge pressure, and access site bleed has been completed. the pump was returned for evaluation. no cannula or catheter kinks were observed. no biomaterial was found around the impeller or purge gap. the low pump flow and high purge pressure were not reproduced when running with the returned purge cassette at p-9. clinical details stated that repositioning the pump did not resolve the low flow and high purge pressure issues. clinical mentions that act was not maintained. data logs show 25 hours into support there is a motor current (mc) spike, upwards ps shift, speed deviation and drop in flows. the purge pressure immediately begins to rise with drop in purge flow over a span of about 2 minutes triggering hpp and purge system blocked alarms. the low pump flows and high purge pressure do not resolve. the root cause of the low pump flow issue was most likely biomaterial. the root cause of the high purge pressure issue was most likely biomaterial. the root cause of the access site bleeding - minor issue was not determined since there is a lack of clinical details. the instructions for use referenced in the initial report is for the impella rp with smartassist system in section: using the automated impella controller with the impella rp with smartassist system catheter.
27082 | 1220648-2024-10756 | 19244826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19244826&pc=OZD | 04/26/2024 | 06/07/2024 |  | 05/04/2024 |  |  | 04/27/2024 | 05/04/2024 |  |  | 10/31/2025 |  | 05/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024415328 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the impella cp was being prepped to insert for support of an unknown age patient admitted in acute myocardial infarction/cardiogenic shock. in the time of prep, the user had an error as the user mistakenly pulled out the red guidewire lumen. the decision was made to replace the impella pump. there was no report of patient harm. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely use issue since the user mistakenly pulled out the red guidewire lumen.
27083 | 1220648-2024-10695 | 19235838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19235838&pc=OZD | 09/22/2022 | 03/26/2025 |  | 05/03/2024 |  |  | 09/22/2022 | 05/03/2024 |  |  | 04/30/2024 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221537 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2022 |  |  |  |  | The complainant reported a 80 year old male patient was implanted with an impella cp device for mechanical circulatory support. while on support, access site bleeding was noted. as a result, the impella cp was explanted. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27084 | 1220648-2024-10697 | 19237095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19237095&pc=OZD | 03/21/2022 | 03/26/2025 |  | 05/03/2024 |  |  | 03/21/2022 | 05/03/2024 |  |  | 12/31/2023 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180081 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. after removal of the 14 fr peel away sheath, a steady flow of bleeding was noted at the access site. manual pressure was held for 10 minutes without any improvement in the bleeding. decision was made by the physician to remove impella, and exchange with the other 14 fr sheath and reinsert the impella catheter for support. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27085 | 1220648-2024-10698 | 19237298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19237298&pc=OZD | 03/11/2022 | 05/03/2024 |  | 05/03/2024 |  |  | 03/11/2022 | 05/03/2024 |  |  | 09/30/2023 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022143168 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/11/2022 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, a significant groin bleed at the impella access site was noted after the patient was coughing. the nurse resolved the bleeding with placement of a sandbag. the patient received three units of packed red blood cells and heparin was discontinued. additionally, the patient¿s urine turned from yellow prior to the impella to red and central venous pressure was low. additionally, the patient's laboratory results were hemolyzed. the physician gave fluid orders to the staff but expressed he did not want to remove the impella immediately. the patient was successfully weaned and was reported to have strong palpable pulses in the foot and hemostasis achieved after the impella was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿.
27086 | 1220648-2024-10699 | 19237315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19237315&pc=OZD | 04/03/2024 | 05/28/2024 |  | 05/03/2024 |  |  | 04/03/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442955 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2024 |  |  |  |  | The complainant reported a patient (race unknown) with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient developed a hematoma. heparin was withheld and then resumed when the hematoma was reduced. the patient remains on impella 5.5 support on the date of this report. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely anticoagulation management as heparin was stopped and hematoma was reduced. the instructions for use referenced in the initial is from ifu for impella 5.5 with smart assist for use during cardiogenic shock, sections: general patient care considerations, entering cardiac output, and potential adverse events (united states).
27087 | 1220648-2024-10700 | 19237412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19237412&pc=OZD | 04/04/2024 | 05/17/2024 |  | 05/03/2024 |  |  | 04/04/2024 | 05/03/2024 |  |  | 10/31/2025 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414456 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed hemolysis on the first day. the patient's urine was red and the lab results came back hemolyzed. as the positioning of the pump was verified to be correct, the staff was advised of other means to treat the hemolysis. cardiology declined further troubleshooting/treatment for the condition and support with hemolysis lasted two more days until the device was explanted. care was withdrawn and the patient expired. abiomed does not have enough information to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation for the reported hemolysis has been completed. pump was not returned for investigation. data logs were reviewed and there were no motor current spikes or positioning issues observed. the root cause of the hemolysis issue was not determined since product was not returned and insufficient clinical information provided and data logs were inconclusive. the instructions for use referenced in the initial report are found in the ifu for impella cp with smartassist system in section: suction , section: hemolysis, section: use of echocardiography for positioning of the impella catheter, and section: potential adverse events (united states) which now includes cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, section h10 has been revised from the initial submission.
27088 | 1220648-2024-10701 | 19237444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19237444&pc=OZD | 04/04/2024 | 05/20/2024 |  | 05/03/2024 |  |  | 04/04/2024 | 05/03/2024 |  |  | 11/30/2025 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024426036 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was supported by multiple pumps (impella cp and escalated impella 5.5) for acute myocardial infarction/cardiogenic shock. the pump support of the impella 5.5 pump was for over six (6) days when the pump was explanted at the time of care withdrawal after the cerebral vascular accident (cva) was diagnosed. the cva was embolic. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (unites states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ | The investigation into the stroke/cva has been completed since the original report was submitted. the device was not returned for investigation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. d6a, d6b.
27089 | 1220648-2024-10702 | 19237470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19237470&pc=OZD | 04/05/2024 | 06/04/2024 |  | 05/03/2024 |  |  | 04/05/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2025464982 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [5 empty] | BI/PAP C/PAP. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2024 |  |  |  |  | The complainant reported a patient (age and race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella cp support, the patient experienced under perfusion of the left leg. the leg color was different from the non-access site leg. the device was not removed due to the failure mode. the team made the decision to withdraw care and the patient expired. abiomed does not have enough information to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the pump was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ g.1 revised manufacturing site address line 1, manufacturing site address line 2, manufacturing site us city or (foreign state/province/territory) and manufacturing site us zip code (55555) or (55555-4444) or foreign postal code (10 digits) as previously entered incorrectly. h.6 added code 3038 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
27090 | 1220648-2024-10703 | 19237560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19237560&pc=OZD | 03/04/2022 | 05/03/2024 |  | 05/03/2024 |  |  | 03/04/2022 | 05/03/2024 |  |  | 12/31/2023 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170168 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/11/2022 |  |  |  |  | The user facility a 54-year-old patient, of unknown ethnicity, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. it was reported while on support the patient experienced red urine and a pump stop was noted. no information was provided on how the hematuria was treated. upon pump stop, the impella cp was removed and replaced with a new impella cp. no patient harm was reported due to the pump stopping. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
27091 | 1220648-2024-10704 | 19237573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19237573&pc=OZD | 04/24/2024 | 06/13/2024 |  | 05/03/2024 |  |  | 04/24/2024 | 05/03/2024 |  |  | 11/30/2025 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024423525 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient in need of support for high-risk surgery. the pump was supporting but, the site was oozy and so the team gave multiple units of blood products. patient support is ongoing and the pump was not explanted despite the ooze. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella 5.5 with smart assist states in sections: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27092 | 1220648-2024-10705 | 19237591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19237591&pc=OZD | 03/12/2022 | 05/14/2025 |  | 05/03/2024 |  |  | 03/12/2022 | 05/03/2024 |  |  | 10/31/2023 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022153068 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Positioning Failure | [6 empty] | Ischemia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there were no distal pulses in foot, but the staff continued to wrap foot in warm blankets. also, the pigtail was found to be behind pap chord, (trabeculae), but inlet mostly free-floating. the impella was pulled back into aorta. echo showed good position. the patient also experienced ventricular tachycardia with failed cardioversion three times. the impella was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27093 | 19237633 | 19237633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19237633&pc=OZD | 04/11/2024 | 04/16/2024 | 04/16/2024 | 05/03/2024 | 04/16/2024 | 05/03/2024 |  | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP SET WITH SMARTASSIST |  |  | OZD |  |  | 0048-0003 |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0048-0003 |  | No |  |  | Not Available |  |  |  | Health Professional |  |  | User Facility report | No | Risk Manager | [6 empty] | 2 MO |  | Female |  |  |  | Malfunction | No | Yes | [7 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Elderly female with history of atrial fibrillation. procedure: pci (percutaneous coronary intervention) with impella and rotablator. during the procedure, bleeding was witnessed. as the sheath was slowly removed, inching up and peeling, it was visualized that the sheath was ripped when it was in patient and visualized as it was exposed. sheath removed. no known harm to patient at this point. manufacturer response for impella 14f peel away sheath, impella cp set with smartassist (per site reporter). will obtain.
27094 | 1220648-2024-10710 | 19237844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19237844&pc=OZD | 03/22/2022 | 03/26/2025 |  | 05/03/2024 |  |  | 03/22/2022 | 05/03/2024 |  |  | 12/31/2023 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176460 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. after the implant, it was noted that there was no distal flow; the patient experienced limb ischemia. as a result, the impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
27095 | 1220648-2024-10711 | 19237898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19237898&pc=OZD | 03/23/2022 | 03/26/2025 |  | 05/03/2024 |  |  | 03/23/2022 | 05/03/2024 |  |  | 12/31/2023 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180728 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/23/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the left leg was found to be mottled and distal pulses could not be identified; the patient experienced limb ischemia. the pump was then explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
27096 | 1220648-2024-10712 | 19237941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19237941&pc=OZD | 04/24/2024 | 07/15/2025 |  | 05/03/2024 |  |  | 04/24/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467921 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Mitral Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | Death |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/24/2024 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as escalation of care from an impella cp device. the impella 5.5 pump was placed, but with extracorporeal membrane oxygenation (ecmo) decannulation, the pump came out of appropriate position. the malpositioned pump was explanted and replaced. the team additionally noted the patient was suffering from new mitral regurgitation that resolved with the pump removal. the team also drew plasma free hemoglobin and the patient had signs of hemolysis. they believe the hemolysis was due to a combination of the pump-ecmo. the patient remains on impella 5.5 support. | Additional information received from the clinical site that the patient expired on (b)(6) 2024. the relationship being "unknown" to the use of the impella. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: axillary insertion of impella 5.5 catheter. ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter. ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation for the reported hemolysis, positioning, and heart valve damage issues has been completed. product was not returned for investigation. data logs were returned and investigated. multiple suction alarms were seen during the time of alleged hemolysis, but the exact time for hemolysis is unknown. patient was also on ecmo which can be a potential contributor to hemolysis. therefore, the root cause of the hemolysis issue was not determined since product was not returned and insufficient clinical information provided and data log investigation is inconclusive. the pump was pulled back while decannulating ecmo and repositioning the pump. no other information is provided. the pump was explanted due to failed positioning attempt. therefore the root cause of the positioning issue was most likely use related. the patient updates on 4/30/23 mention that the impella was displaced into aorta and repositioning was difficult. data logs show multiple "impella position unknown" and "impella in aorta" alarms around that time along with spread placement signal. the mr was resolved once the pump was explanted and implanted again. therefore, the root cause of the heart valve damage issue was not determined since product was not returned and insufficient clinical information provided. | B2 death and date of death added. b5 additional information received added. h1 type of report updated to death. h6 revised to reflect death d6b explant date was erroneously entered on the initial manufacturer device report number 1220648-2024-10712. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-10712. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures 1220648-2024-10712. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-10712.
27097 | 1220648-2024-10713 | 19237962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19237962&pc=OZD | 09/14/2022 | 05/03/2024 |  | 05/03/2024 |  |  | 09/14/2022 | 05/03/2024 |  |  | 02/29/2024 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196811 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/14/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support the physician tried to reposition the device due to shallow placement and the pigtail in the papillary. multiple attempts were made without success or significant rhythm/flow changes. additionally, a small track ooze on the impella side was noted. heparin was withheld. subsequently the device successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27098 | 1220648-2024-10714 | 19238147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19238147&pc=OZD | 04/14/2024 | 06/04/2024 |  | 05/03/2024 |  |  | 04/14/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025449287 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Migration or Expulsion of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2024 |  |  |  |  | The user facility reported a 40-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella cp support was having persistent ectopy. echo showing pump measuring 6cm. the cp was pulled back at bedside, but impella jumped back into the aorta and was unable to recross. the impella cp was removed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. data logs confirmed pump was in aortic area. pump then was stopped manually & disconnected from the console. based on clinical description & data analysis, the root cause of positioning issue was most likely use related.
27099 | 1220648-2024-10715 | 19238203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19238203&pc=OZD | 04/25/2024 | 06/13/2024 |  | 05/03/2024 |  |  | 04/25/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205389 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The user facility reported a 67-year-old male in pre-operation placement was implanted with an impella 5.5 for mechanical circulatory support. the automated impella controller (aic) was showing a controller failure alarm that would not shut off. no patient harm has been reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - controller failure (red alarm) has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the reported failure mode was unable to be reproduced. however, dextrose deposit was found on the purge system, and the purge door was not able to open smoothly, most likely due to the dextrose deposit on the purge door hinge, which is unrelated to the reported failure mode. the root cause of improper shutdown was not determined due to the inability to reproduce.
27100 | 1220648-2024-10716 | 19238307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19238307&pc=OZD | 04/15/2024 | 05/01/2025 |  | 05/03/2024 |  |  | 04/15/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025456888 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/15/2024 |  |  |  |  | The complainant had a new impella cp pump opened because the flag of the previous impella cp was no longer in place and the doctor had to perform the ventricular passage for a second time. the implantation was performed with the patient resuscitated. when the placement wire was in the ventricle, the physician threaded the pump onto the wire and wanted to advance it, but before implantation physician saw that the placement wire was no longer in the ventricle. the ventricle had to be passed again. as the flag for the pump guide was already out and the physician wanted to give the patient maximum implantation safety; therefore, physician decided to have a new pump opened. there was no reported patient harm. | [5 empty] | The device was not returned for evaluation. when the investigation is complete a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. product was not returned for review. clinical details were sufficient to determine the root cause. the root cause of delivery/positioning issue was most likely due to use related. the instructions for use for the impella cp with smartassist system states: section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | Failure mode changed to delivery issue as the guidewire (gw) was not in optimal position to deliver the pump. gw was not returned for analysis. pump was returned for analysis. visual inspection found no issues on the pump. the root cause of delivery issue was most likely use related due to improper use during support. d9 date device returned to manufacturer added. d1 brand name was erroneously entered on the initial manufacturer device report number 1220648-2024-10716. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2024-10716. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-10716. g1 reporting contact email revised on manufacturer device report 1220648-2024-10716 in accordance with updated procedures. g1 manufacturing site country was inadvertently omitted on the initial manufacturer device report number 1220648-2024-10716. g2 report source; health professional was inadvertently omitted on the initial manufacturer device report number 1220648-2024-10716. h6 component code 3038 was erroneously entered on the initial manufacturer device report number 1220648-2024-10716.
27101 | 1220648-2024-10717 | 19238373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19238373&pc=OZD | 03/26/2022 | 05/03/2024 |  | 05/03/2024 |  |  | 03/26/2022 | 05/03/2024 |  |  | 11/30/2023 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163758 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/26/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced three episodes of ventricular tachycardia, each episode was successfully treated with defibrillation. additionally, a small hematoma and oozing formed at the access site when the peel away sheath was removed, the physician placed a reposition sheath. manual pressure was applied for approximately ten minutes and the physician cinched down two proglide sutures around the sheath and hemostasis was achieved. the patient remained on impella support for 24 hours and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. . impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer as noted in the impella ifu: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27102 | 1220648-2024-10718 | 19238397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19238397&pc=OZD | 04/17/2024 | 06/25/2024 |  | 05/03/2024 |  |  | 04/17/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2025449798 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Unknown |  |  |  | Injury | No | Yes | Material Rupture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 04/17/2024 |  |  |  |  | The complainant reported that during impella cp implant procedure for 69-year-old patient, the impella ruptured in two pieces. there were no delivery issues prior to the damage but the vessels were noted to be tortuous. the device was removed as a result of the issue. there was no reported patient injury. | [5 empty] | The impella device was received from the customer on 01-may-2024 and therefore, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (detached inflow/outflow - great vessel) issue has been completed since the original report was filed. the pump was returned for investigation. visual inspection found that pump had inflow cage detached from cannula with detached tip at tip of introducer sheath and nitinol coil stayed connected to the rest of the pump and was stretched through the introducer. clinical details noted that patient had known cad and tortuosity. the root cause of the inflow cage and pigtail detachment was due to use with excessive force. the root cause of the introducer damage was due to patient's tortuosity.
27103 | 1220648-2024-10719 | 19238467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19238467&pc=OZD | 03/18/2022 | 03/26/2025 |  | 05/03/2024 |  |  | 03/18/2022 | 05/03/2024 |  |  | 10/31/2023 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151721 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Infusion or Flow Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved by discontinuing systemic heparin and monitoring the activated clotting time results. additionally, the patient experienced saturated flow that was resolved by decreasing purge flow. the pump also had suction alarms (low pump flow) that were resolved by removal and replacing with a new impella pump. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smartassist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Access site bleeding: the cause of the access site bleeding could not be determined as minimal clinical details are known. saturated flow: the cause of the saturated flow was not determined as no logs or product was returned. low pump flow: the cause of the low pump flow was not determined as no product or logs were returned. this report is being filed as part of a retrospective review of historical records.
27104 | 1220648-2024-10720 | 19238492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19238492&pc=OZD | 04/17/2024 | 06/24/2024 |  | 05/03/2024 |  |  | 04/17/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025451644 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/17/2024 |  |  |  |  | The user facility reported a 66-year-old male patient placed on 5.5 support as care was escalated from intra-aortic balloon pump for the coronary artery bypass graft (cabg) procedure. it was reported that on day 5 of support the patient was prepped for cabg and an air embolus was noted on the echocardiogram. the patient was removed from the bypass pump and the air resolved. the patient remained on support for 17 days until weaned and explanted. | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured hemolysis with a "possible" relationship to the impella device used: ¿per md note in chart: hemoglobin stable, likely has mild hemolysis in the setting of being on impella¿. the patient recovered with no sequelae. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | Section b5, g2 and h6 were updated as per additional information received. | The investigation into the embolism and the hemolysis issues has been completed since the original report was submitted. the product was not returned for investigation. the root cause of the embolism and the hemolysis issues was not determined since product was not returned and insufficient clinical information related to the incident provided.
27105 | 1220648-2024-10721 | 19238527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19238527&pc=OZD | 04/25/2024 | 06/13/2024 |  | 05/03/2024 |  |  | 04/25/2024 | 05/03/2024 |  |  | 11/30/2025 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025436914 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The user facility reported a 68-year-old male of unknown race and ethnicity in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed hemolysis and experienced bleeding from their access site during impella cp support. femostop was placed at the site to manage the oozing and the patient was given four units of fresh frozen plasma and four units of packed red blood cells to treat the conditions. support was maintained with the implanted pump following the transfusion. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major and the hemolysis issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the access site bleeding and the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided.
27106 | 1220648-2024-10722 | 19238633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19238633&pc=OZD | 04/05/2022 | 05/03/2024 |  | 05/03/2024 |  |  | 04/05/2022 | 05/03/2024 |  |  | 11/30/2023 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161933 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/05/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the physician reported possible limb ischemia. the decision was made to remove the extracorporeal membrane oxygenation and remove the impella cp then replace the pump with an impella 5.5. the impella cp was explanted and replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) as noted in the impella ifu: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
27107 | 1220648-2024-10723 | 19238723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19238723&pc=OZD | 04/25/2024 | 06/27/2024 |  | 05/03/2024 |  |  | 04/25/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466172 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The user facility reported a 65-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that impella cp was prepped after impella 5.5 was removed but the impella cp would not pass the innominate artery. the decision was made to implant the impella cp via femoral access. it was further reported that the patient lost pulses at the limb and reperfusion sheath was placed. support was maintained with the implanted pump and the patient was reported as stable. | Additional information. it was further reported that care was withheld and the patient expired. per abiomed medical safety office review: abiomed does not have enough information to associate use of device with outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ | The investigation into the delivery issues and the limb ischemia - reversible issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of delivery issue was most likely patient condition since circumferential calcified lesion was noted at origin of innominate artery. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. d6b added. f6 and f8 removed.
27108 | 1220648-2024-10724 | 19238834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19238834&pc=OZD | 03/15/2022 | 05/03/2024 |  | 05/03/2024 |  |  | 03/15/2022 | 05/03/2024 |  |  | 12/31/2023 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173803 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/15/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, access site bleeding was noted by the oozing at the insertion site. the oozing was treated by redressing the site with proper angle matching along with a pant leg cut of the tegaderm which had been pushing tight on the sheath. as oozing persisted, the site was redressed, one unit of blood was given to the patient. the pump was explanted, and a small hematoma appeared along with some persistent oozing, which was resolved through manual pressure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27109 | 1220648-2024-10725 | 19238922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19238922&pc=OZD | 04/25/2024 | 07/03/2024 |  | 05/03/2024 |  |  | 04/26/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2025470104 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2024 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had extracorporeal membrane oxygenation/impella cp support ongoing when the patient had thrombosis confirmed. the impella cp was explanted as a result and the clot was observed at catheter shaft. the patient was returned to the intensive care unit without impella cp support. the team alleges that the anticoagulation pause may have caused the thrombosis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined.
27110 | 1220648-2024-10726 | 19239133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19239133&pc=OZD | 04/26/2024 | 06/13/2024 |  | 05/03/2024 |  |  | 04/26/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443146 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an impella 5.5 device implanted as elective placement for mechanical circulatory support. however, five minutes following start up, the implanted impella pump experienced a spike in motor current and subsequently stopped functioning. attempts to restart the pump were unsuccessful, and the decision was made to replace the pump and console for ongoing patient support. there was no report of patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported pump stop has been completed. data logs confirm a pump stop alarm 13 which was unable to be restarted. no biomaterial was found on the returned pump and it ran correctly in the lab. no other pump abnormalities were found. the root cause of the pump stop issue was not determined since the pump stop was not reproduced in the lab.
27111 | 1220648-2024-10727 | 19239209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19239209&pc=OZD | 04/23/2022 | 03/26/2025 |  | 05/03/2024 |  |  | 04/23/2022 | 05/03/2024 |  |  | 08/31/2023 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022133006 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2022 |  |  |  |  | The complainant reported a patient was being implanted with an impella cp device for mechanical circulatory support due to aortic stenosis. the impella cp implantation was aborted due to the inability to cross the aortic valve. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27112 | 1220648-2024-10728 | 19239339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19239339&pc=OZD | 04/26/2024 | 01/30/2025 |  | 05/03/2024 |  |  | 04/26/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1619563 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an impella 5.5 device implanted as elective placement for mechanical circulatory support. however, five minutes following start up, the implanted impella pump experienced a spike in motor current and subsequently stopped functioning. attempts to restart the pump were unsuccessful, and the decision was made to replace the pump and console for ongoing patient support. there was no report of patient harm. | [5 empty] | The automated impella controller (aic) was not received from the customer; and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of automated imepella controller (aic) - pump stop has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d2 common device name was revised on manufacturer device report 1220648-2024-10728 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-10728 in accordance with updated procedures. g1 reporting contact fax number missing on manufacturer device report 1220648-2024-10728. h6 component code was reported incorrectly on manufacturer device report 1220648-2024-10728. new component code was added.
27113 | 1220648-2024-10729 | 19239562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19239562&pc=OZD | 04/28/2022 | 03/26/2025 |  | 05/03/2024 |  |  | 04/28/2022 | 05/03/2024 |  |  | 02/29/2024 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193202 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced thrombus that was resolved with embolectomy. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
27114 | 1220648-2024-10731 | 19239574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19239574&pc=OZD | 04/19/2024 | 10/10/2024 |  | 05/03/2024 |  |  | 04/19/2024 | 05/03/2024 |  |  | 05/31/2025 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024358334 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp support, the patient experienced oozing from the access site. the hemoglobin dropped from 7 to 5.7. oozing was attempted to be resolved with sutures, but the issue did not resolve completely. therefore, a pressure bandage was also applied. patient received 2 frozen plasmas and blood products. the impella was not removed as a result of the failure. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Manufacturer device report 1220648-2024-10731 was created in error. all investigation findings can be found in manufacturer device report 1220648-2024-10624.
27115 | 1220648-2024-10730 | 19239575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19239575&pc=OZD | 04/30/2022 | 03/26/2025 |  | 05/03/2024 |  |  | 04/30/2022 | 05/03/2024 |  |  | 07/31/2023 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022113430 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. while on support oozing/bleeding around the site was noted. no additional information was provided related to the resolution of access site bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for evaluation. this report is being filed as part of a retrospective review of historical records.
27116 | 1220648-2024-10732 | 19239683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19239683&pc=OZD | 04/26/2024 | 06/04/2024 |  | 05/03/2024 |  |  | 04/26/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447798 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2024 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant had an impella cp placed for support of a patient admitted with acute myocardial infarction and was transferred to the tertiary center for coronary artery bypass grafting). days after the pump was implanted, the team alleged the pump was placed within the left ventricle and there was now mural clot present. additionally, the patient had signs of hemolysis. the urine color had changed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smart assist system. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding ¿ minor and the hemolysis could not be determined due to insufficient clinical information and lack of product returned. as the necessary information was not provided, the root cause of the thrombus was not determined.
27117 | 1220648-2024-10733 | 19239744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19239744&pc=OZD | 04/21/2024 | 06/11/2024 |  | 05/03/2024 |  |  | 04/21/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025465574 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support and there were positioning issues. placement was technically challenging of the impella pump into the left ventricle secondary to patient's heart anatomy. despite use of best practices, efforts to reposition the pump into ideal left ventricle position were unsuccessful as the physician only managed to find a rather shallow spot for the inlet. once in "position," the pump was maximized for patient support. after pci, it was noted the physician made a final attempt to reposition the pump for more optimal flow, but it was unsuccessful. the pump's p-level was decreased, the pump was pulled across the aortic valve and later removed. there was no report of adverse effects to the patient. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. according to the clinical, while inserting the impella cp device it became difficult to insert the pump into the lv due to patient anatomy. after placing the pump in a shallow position support was maximized. later the same day another attempt was made to reposition the pump with the same negative results. the decision was then made to wean the patient off of support and remove the device. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the positioning issues was patient condition related. d4-updated model number
27118 | 1220648-2024-10734 | 19239785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19239785&pc=OZD | 03/25/2022 | 05/03/2024 |  | 05/03/2024 |  |  | 03/25/2022 | 05/03/2024 |  |  | 09/30/2023 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2022146485 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/25/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella was very positional and continued to dive in too deep. the pump was ran at p4 until the position could be corrected. it is unknown if the staff was able to correct the position. the patient successfully weaned from the impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27119 | 1220648-2024-10735 | 19239832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19239832&pc=OZD | 04/24/2024 | 05/29/2024 |  | 05/03/2024 |  |  | 04/24/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025468758 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Disability | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient in acute myocardial infarction/cardiogenic shock implanted with an impella cp device for support suffered a stroke. head scan showed right posterior cerebral artery stroke, and the patient was deemed not a candidate for intervention. the patient was made a do not resuscitate (dnr), and heparin was held due to high risk of hemorrhagic conversion of the stroke. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system-section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. d4-updated model number
27120 | 1220648-2024-10736 | 19239985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19239985&pc=OZD | 04/24/2024 | 06/13/2024 |  | 05/03/2024 |  |  | 04/24/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025464841 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Dislodged or Dislocated | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support when there were positioning issues. there was a pump in aorta alarm and the staff attempted to reposition the pump multiple times. the pump obtained proper position and after the patient was ambulating, the impella became displaced. the staff were unable to reposition again, the pump was removed. reportable due to positioning issues. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the positioning issue was not determined because no product was returned, and not enough clinical information was given.
27121 | 1220648-2024-10737 | 19239989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19239989&pc=OZD | 04/24/2024 | 06/13/2024 |  | 05/03/2024 |  |  | 04/24/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025449976 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Separation | Material Perforation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | The user facility reported a 46-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support when the sterile sleeve ripped out of the impella hub. this occurred during a repositioning of the pump. the physician immediately was aware and pushed the sleeve back into the hub. the sleeve was cleaned and reinforced with tegaderm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (sterile sleeve damage) has been completed since the original report was submitted. the device was not returned for investigation. per the clinical information provided, the sterile sleeve was ripped during repositioning of the pump by applying excessive force. therefore, the root cause of the product damage (sterile sleeve damage) issue was most likely use related to excessive force.
27122 | 1220648-2024-10738 | 19240377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19240377&pc=OZD | 04/26/2024 | 06/12/2024 |  | 05/03/2024 |  |  | 04/26/2024 | 05/03/2024 |  |  | 09/30/2025 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400463 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2024 |  |  |  |  | The complainant reported a 67-year-old male patient with post cardiotomy cardiogenic shock was to be implanted with an impella cp for mechanical circulatory support, but the medical staff experienced delivery issues. the patient underwent runs of ventricular tachycardia when the diagnostic catheter was being placed. the doctor believed the patient could not tolerate the impella resting in their left ventricle without causing arrhythmias. the impella introducer was used on the patient. however, the pump was not used on the patient. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related, the doctor felt that the patient could not tolerate the impella resting in their left ventricle without causing arrhythmias as per the clinical description. as the necessary information was not provided, the root cause of the vt/vf was not determined. in accordance with updated procedures, sections d6 have been revised from the initial submission.
27123 | 1220648-2024-10739 | 19240993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19240993&pc=OZD | 04/26/2024 | 06/25/2024 |  | 05/03/2024 |  |  | 04/26/2024 | 05/03/2024 |  |  | 11/30/2025 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024431326 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2024 |  |  |  |  | The user facility reported a 62-year-old female with a high risk of percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support, but the medical staff experienced delivery issues. the patient underwent access site bleeding. the perclose failed to take, and the medical staff was unable to advance the sheath into the arteriotomy due to a hostile groin, a large pannus, and likely a hematoma pushing the sheath away. the medical staff consulted vascular surgery and administered two units of red blood cells, two perclose sutures, and an angio-seal to obtain hemostasis. the patient survived the event. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of access site bleeding was determined to be patient condition because they were unable to advance the 11 french sheath into arteriotomy of the patient. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 code 4625 was reported incorrectly on the previously submitted report.
27124 | 1220648-2024-10742 | 19241846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19241846&pc=OZD | 04/25/2024 | 06/07/2024 |  | 05/03/2024 |  |  | 04/25/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED,INC |  |  | US |  | 2025463883 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The user facility reported a 70-year-old female patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. it was noted that the device was unable to pass through the right axillary artery and graft due to existing calcification in the artery. after multiple failed attempts, the decision was made to abort the implant. the patient remained support with an impella cp device. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist for use during cardiogenic shock in sections: warnings & cautions: warnings, section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves.
27125 | 1220648-2024-10743 | 19241889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19241889&pc=OZD | 04/26/2024 | 06/03/2024 |  | 05/03/2024 |  |  | 04/26/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461230 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2024 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient was on impella 5.5 support and there was a thrombus noted upon explant. the physician found a large thrombus on the outflow upon explanting the impella. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined.
27126 | 1220648-2024-10740 | 19242414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19242414&pc=OZD | 04/04/2024 | 06/14/2024 |  | 05/03/2024 |  |  | 04/04/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454404 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient went into ventricular tachycardia (vt). the pigtail catheter was advanced over the 0.035" j wire into the left ventricle and the patient went into vt. the patient was shocked four times. the procedure proceeded and the patient went into vt again. the patient was shocked. the implant was successful, and support continued. throughout support, the patient had further arrhythmia issues. the patient did not have a history of vt prior to the impella insertion. | There is not enough information to associate use of device with outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the vf/vt/af/at (ventricular tachycardia) issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
27127 | 1220648-2024-10744 | 19242642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19242642&pc=OZD | 04/02/2024 | 06/12/2024 |  | 05/03/2024 |  |  | 04/04/2024 | 05/03/2024 |  |  | 08/31/2025 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024383014 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella 5.5 support had ultrasound of the cut down site showing hematoma. two days later the patient was taken to the operation room for a wash-out of the axillary. it was further reported that the patient showed decline in hand and finger numbness tingling post evacuation of blood clot in axillary. systemic anticoagulation was held due to evacuation of hematoma. the patient is stable. the patient remains on support on the date of this report. | There is no association between impella use and outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding - major issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. impella rp with smart assist system with automated impella controller section: warnings and cautions : ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ in accordance with updated procedures, sections b5, d6, and h10 have been revised from the initial submission.
27128 | 1220648-2024-10741 | 19242805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19242805&pc=OZD | 04/25/2024 | 07/01/2024 |  | 05/03/2024 |  |  | 04/25/2024 | 05/03/2024 |  |  | 02/28/2025 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024324044 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Migration or Expulsion of Device | Positioning Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The user facility reported a 52-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp migrated to the aorta during patient transfer. the physician decided to remove the impella in the intensive care unit. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. according to the clinical information, on the second day of impella cp support during patient transfer, the pump migrated into the aorta. the decision was then made to remove the pump. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the positioning issues was patient movement. b.7 revised as this information was inadvertently omitted from the previously submitted report. h.6 medical device problem codes were revised in accordance with updated procedures. a new code has been added to medical device problem codes.
27129 | 1220648-2024-10745 | 19242815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19242815&pc=OZD | 04/25/2024 | 06/13/2024 |  | 05/03/2024 |  |  | 04/25/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485906 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was also on extracorporeal membrane oxygenation. during support, the team observed signs of limb ischemia on the impella cp accessed leg and so a perfusion catheter was placed. the patient remains on impella cp support on the date of this report. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the limb ischemia and the low pump flow has been completed. according to the clinical, on the first day of impella cp and ecmo support the pump was unable to go above p3 without suction. blood volume was given and an antegrade sheath was inserted. limb ischemia symptoms were also observed so a perfusion catheter was placed. no product was returned for a visual inspection. data log analysis confirmed the presence of multiple suction alarms but no other abnormalities. the most likely cause of the low pump flow was patient condition related. the root cause of the limb ischemia could not be determined without additional information. the instructions for use referenced in the initial report is for the impella cp with smart assist system in section: potential adverse events (united states). in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27130 | 1220648-2024-10746 | 19242971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19242971&pc=OZD | 04/26/2024 | 06/04/2024 |  | 05/03/2024 |  |  | 04/26/2024 | 05/03/2024 |  |  | 04/30/2025 |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352781 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [5 empty] | VENTILATOR | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. following explant of the impella cp pump, the access site was noted to be infected with pus coming out. the site was surgically closed and a wound vac was placed over the site. an impella 5.5 pump was placed via the opposite axillary site for ongoing support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined.
27131 | 1220648-2024-10747 | 19243112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19243112&pc=OZD | 04/26/2024 | 06/13/2024 |  | 05/03/2024 |  |  | 04/26/2024 | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025468761 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2024 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp pump was explanted as a result of a large hematoma and bleeding around the access site during patient support. a peripheral balloon was placed to occlude the flow and manual pressure was held to achieve hemostasis. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for impella cp with smart assist system states. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) :¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d4-corrected model number
27132 | MW5154540 | 19248861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248861&pc=OZD | 04/13/2024 | 04/30/2024 |  | 05/03/2024 |  |  |  | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0048-3092 |  |  | ABIOMED, INC. | [8 empty] | No | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | NA |  | Unknown |  |  |  | Death | No | No | [7 empty] | Device Embedded In Tissue or Plaque | [12 empty] | Death | [28 empty] | Initial submission | [9 empty] | Patient pulled on external portion of impella line causing it to break clean through into two pieces making it non operable and non repairable. impella remained inside of patient. patient expired within minutes.
27133 | MW5154541 | 19248876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19248876&pc=OZD | 04/12/2024 | 04/30/2024 |  | 05/03/2024 |  |  |  | 05/03/2024 |  |  |  |  | 05/03/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  |  |  | 2025466827 |  |  |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 81 YR |  | Male | 101 KG | Non Hispanic | White | Death | Yes | No | [7 empty] | Myocardial Infarction | [12 empty] | Death | [28 empty] | Initial submission | [9 empty] | Impella device that was placed during an emergent procedure for a myocardial infarction, needed to be removed at a later time due to low degree of functioning. when removed - questionable piece of placement sheath lodged in motor.
27134 | 1220648-2024-10658 | 19228328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19228328&pc=OZD | 03/18/2022 | 03/26/2025 |  | 05/02/2024 |  |  | 03/18/2022 | 05/02/2024 |  |  | 05/31/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022083009 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on impella support, a fluoroscopy revealed no flow distal to impella insertion site. the peel away sheath was removed and the staff discussed an external femorofemoral bypass, as access gained on contralateral left side showed flow distal to insertion site. upon arrival to the unit, no dopplerable pulses were present and a vascular ultrasound revealed no flow distal to impella insertion site. an external femorofemoral bypass was placed to restore flow to the leg. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
27135 | 1220648-2024-10659 | 19228420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19228420&pc=OZD | 03/21/2022 | 05/02/2024 |  | 05/02/2024 |  |  | 03/21/2022 | 05/02/2024 |  |  | 11/30/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169158 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/21/2022 |  |  |  |  | The complainant reported a patient, race unknown, was taken to the cardiac catheterization laboratory (ccl) for impella placement. due to a positive stress test, and shortness of breath for the last two months the physician determined coronary artery bypass graft was needed and the ventricle was unloaded prior to the surgery and the patient was implanted with an impella cp. it was reported while on support, the patient experienced access site bleeding, bleeding from the nose and mouth, and bloody urine. the physician did not remove the impella due to the patient being unstable and gave orders for no heparin until the bloody urine and bleeding stopped. the patient was taken back to the operating room for post operative bleeding, additional surgery was performed, and the bleeding was successfully managed. the patient received six units of packed red blood cells. the patient was transferred to another facility and venoarterial extracorporeal membrane oxygenation weaning medication was added and work up being done for transplant or left ventricular assist device. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output . "use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis . ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. section: suction . ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. this report is being filed as part of a retrospective review of historical records.
27136 | 1220648-2024-10660 | 19228664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19228664&pc=OZD | 10/30/2023 | 05/22/2024 |  | 05/02/2024 |  |  | 04/04/2024 | 05/02/2024 |  |  | 08/31/2025 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024398230 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | A notification was received long-term outcome and quality indicator impella registry that the patient developed hemolysis during impella cp support. a transthoracic echocardiogram showed that the impella cp pigtail was facing towards the papillary muscle. should the patient¿s lactate dehydrogenase levels continue to rise, repositioning or further intervention would become necessary. it is unknown how the hemolysis resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis. | The investigation into the hemolysis issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the hemolysis cannot be determined as there is insufficient clinical information provided, and a lack of product or data logs returned.
27137 | 1220648-2024-10663 | 19228822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19228822&pc=OZD | 03/17/2022 | 03/26/2025 |  | 05/02/2024 |  |  | 03/17/2022 | 05/02/2024 |  |  | 08/31/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127482 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump position was in the aorta with plans to remove the device. device was repositioned the following morning with no issues. device was successfully weaned and explanted. the patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27138 | 1220648-2024-10664 | 19229001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19229001&pc=OZD | 03/08/2022 | 03/26/2025 |  | 05/02/2024 |  |  | 03/08/2022 | 05/02/2024 |  |  | 10/31/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152152 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. upon implant, the impella cp was placed into the common femoral but was unable to pass into the aorta and the decision was made to abort the impella cp placement and place a impella 2.5 instead. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci ¿section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27139 | 1220648-2024-10665 | 19229261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19229261&pc=OZD | 03/12/2022 | 03/26/2025 |  | 05/02/2024 |  |  | 03/12/2022 | 05/02/2024 |  |  | 11/30/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165051 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the nurse broke the sidearm when performing purge fluid bag change. a decision was made to swap out the pump for continuation of therapy. the impella cp was removed and replaced without issue and support initiated. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk cpi. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
27140 | 1220648-2024-10666 | 19229310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19229310&pc=OZD | 03/15/2022 | 05/02/2024 |  | 05/02/2024 |  |  | 03/15/2022 | 05/02/2024 |  |  | 12/31/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165981 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 03/15/2022 |  |  |  |  | The user facility reported a 82-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced several short episodes of ventricular tachycardia during the night. cardio-pulmonary resuscitations were performed on the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27141 | 1220648-2024-10667 | 19229330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19229330&pc=OZD | 03/22/2022 | 09/22/2025 |  | 05/02/2024 |  |  | 03/22/2022 | 05/02/2024 |  |  | 12/31/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172825 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2022 |  |  |  |  | The complainant reported a patient was taken to the cardiac catheterization laboratory (ccl) for impella placement. the impella was successfully implanted in the right groin. while on support, the patient experienced access site bleeding, atrial tachycardia and sepsis. the impella was removed and repair to the right iliac for bleeding done by the vascular surgeon. it was reported the groin sutures were not holding the pump in place and a bit of the catheter was sticking out, this was fixed by adding forward tension sutures and redressing, no further oozing was noted. reportedly the physician suspected the patient had a urinary tract infection which could be the reason for the septic issue. the impella site was noted to be clean and dry, the patient¿s urine was reported to be clear yellow, and the coronary artery bypass graft was planned for the next day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-10667 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-10667 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-10667. h6 component code revised from the initial submission in accordance with updated procedures.
27142 | 1220648-2024-10668 | 19229343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19229343&pc=OZD | 03/17/2022 | 05/02/2024 |  | 05/02/2024 |  |  | 03/17/2022 | 05/02/2024 |  |  | 10/31/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022153213 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/17/2022 |  |  |  |  | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. the patient experienced bleeding from the chest tube and at the impella site. fem-stop was placed and bleeding at the site was controlled. heparin was withheld. no additional information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27143 | 1220648-2024-10669 | 19229373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19229373&pc=OZD | 03/17/2022 | 05/02/2024 |  | 05/02/2024 |  |  | 03/17/2022 | 05/02/2024 |  |  | 11/30/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166874 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/17/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma. the physician decided to wean and explant the impella due to the hematoma. the patient was hemodynamically stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27144 | 1220648-2024-10670 | 19229410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19229410&pc=OZD | 03/17/2022 | 05/02/2024 |  | 05/02/2024 |  |  | 03/17/2022 | 05/02/2024 |  |  | 12/31/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163697 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/17/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced heavy bleeding at the impella access site. the physician decided to explant the device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27145 | 1220648-2024-10672 | 19229456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19229456&pc=OZD | 03/22/2022 | 10/29/2025 |  | 05/02/2024 |  |  | 03/22/2022 | 05/02/2024 |  |  | 01/31/2024 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-EU |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC |  |  | US |  | 2023178449 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2022 |  |  |  |  | The complainant reported that while on impella cp support the purge pressure was high. purge-cassette was changed by the staff under guidance. 50 units of heparin in purge was noted, and due to critical coagulation, heparin in purge reduced to 25units/ml. purge values were then reported as resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | After evaluation, the complaint does not meet the criteria for mdr reportability as defined by 21 cfr 803.50. the issue will be tracked and trended internally.
27146 | 1220648-2024-10673 | 19229581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19229581&pc=OZD | 03/24/2022 | 05/02/2024 |  | 05/02/2024 |  |  | 03/24/2022 | 05/02/2024 |  |  | 12/31/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176052 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/24/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced marginal hematomas and no distal perfusion to either leg. the surgeon was deciding on the impella 5.5 or extra-corporeal membrane oxygenation. the patient remained sedated and was escalated to veno-arterial extracorporeal membrane oxygenation (va-ecmo). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. impella cp with smart assist system instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output . use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
27147 | 1220648-2024-10674 | 19229608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19229608&pc=OZD | 10/18/2022 | 05/02/2024 |  | 05/02/2024 |  |  | 10/18/2022 | 05/02/2024 |  |  | 07/31/2024 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253309 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 10/18/2022 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced bleeding at the impella site. the patient was administered two units packed red blood cells. additionally, some hematuria was noted so the impella was repositioned (advanced) to reduce risk of hemolysis. subsequently it was reported that the patient was hemodynamically stable and the hemolysis resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
27148 | 1220648-2024-10675 | 19229609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19229609&pc=OZD | 03/24/2022 | 03/26/2025 |  | 05/02/2024 |  |  | 03/24/2022 | 05/02/2024 |  |  | 12/31/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023172977 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/24/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient experienced runs of ventricular tachycardia/ventricular fibrillation accompanied by suction alarms. no resolution was specified in either case, but the patient reverted to stable. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27149 | 1220648-2024-10676 | 19229763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19229763&pc=OZD | 03/13/2022 | 05/02/2024 |  | 05/02/2024 |  |  | 03/13/2022 | 05/02/2024 |  |  | 09/30/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022145829 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/13/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved with manual pressure, withholding of heparin, and one unit of blood given. additionally, the patient experienced a thrombus with no resolution specified and hemolysis that resolved after explanting the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient condition including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
27150 | 1220648-2024-10678 | 19230154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19230154&pc=OZD | 03/13/2022 | 05/02/2024 |  | 05/02/2024 |  |  | 03/13/2022 | 05/02/2024 |  |  | 09/30/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140100 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/13/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support after presenting with shortness of breath and a diagnosis of cardiomyopathy. during support, the patient was having vascular issues from possible right femoral vein site (for possible dialysis) as well as right femoral artery impella site. manual pressure being held over both right femoral sites. the decision was made to remove both sites due to increased hematoma and bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
27151 | 1220648-2024-10679 | 19230388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19230388&pc=OZD | 11/21/2022 | 03/26/2025 |  | 05/02/2024 |  |  | 11/21/2022 | 05/02/2024 |  |  | 07/31/2024 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257739 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. when the physician attempted to reposition the device, the patient continuously went into ventricular tachycardia. patient also experienced an oozy groin from the scalpel incision that the physician made due to difficulty advancing the sheath on insertion. heparin-induced thrombocytopenia was suspected. heparin was switched to argatroban. the patient was subsequently escalated to the impella 5.5 pump. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event (access site bleeding) are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system instructions for use for the related event (thrombocytopenia) are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system instructions for use for the related event (difficulty inserting/removing introducer) are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27152 | 1220648-2024-10680 | 19230544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19230544&pc=OZD | 04/03/2024 | 05/10/2024 |  | 05/02/2024 |  |  | 04/03/2024 | 05/02/2024 |  |  |  |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025463849 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2024 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 pump placed for support of a patient admitted in cgs/ami (cardiogenic shock / acute myocardial infarction). the pump was placed along with ecmo (extracorporeal membrane oxygenation) and the iabp (intra-aortic balloon pump) was removed. the pump support was ongoing for 1 day when the pump had a bleed observed. the bleeding site had an exploration and cautery done. the support proceeded til the care was withdrawn. the patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. d4 updated model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27153 | 1220648-2024-10681 | 19230665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19230665&pc=OZD | 03/15/2024 | 05/20/2024 |  | 05/02/2024 |  |  | 04/09/2024 | 05/02/2024 |  |  | 11/30/2025 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024423404 |  | 0048-0045 |  | No |  |  | No |  | No | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Sepsis | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | We received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that insertion site infection occurred with a possible relationship to the impella cp used on this patient. | [5 empty] | The investigation into the infection/sepsis has been completed since the original report was submitted. the device was not returned for investigation. because clinical evidence has not been provided, the root cause of the infection cannot be determined. d6a, d6b | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿
27154 | 1220648-2024-10682 | 19230731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19230731&pc=OZD | 04/02/2024 | 05/15/2024 |  | 05/02/2024 |  |  | 04/02/2024 | 05/02/2024 |  |  | 11/30/2025 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434092 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during cardiac surgery. the pump remained on for support along with extracorporeal membrane oxygenation. the patient experienced persistent ooze from the access site. the team had to infuse blood products and monitor the neurological status of the patient due to underlying anoxic brain injury and being made do not resuscitate by the family. the patient was listed as an organ donor and expired while on support. there is no association between impella use and outcome. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
27155 | 1220648-2024-10683 | 19231005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19231005&pc=OZD | 03/31/2022 | 05/02/2024 |  | 05/02/2024 |  |  | 03/31/2022 | 05/02/2024 |  |  | 11/30/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170067 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 03/31/2022 |  |  |  |  | The user facility reported a 74-year-old patient on impella cp had access site had oozing and hemolysis. the impella was inserted and the patient had access site had oozing and hematoma. the nurse was holding manual pressure, heparin removed from purge. the patient¿s urine was pink, hemolysis remains unconfirmed due to patient improvement. the impella was weaned down and heparin was removed due to persistent suspected hemolysis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
27156 | 1220648-2024-10684 | 19231170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19231170&pc=OZD | 04/12/2024 | 06/04/2024 |  | 05/02/2024 |  |  | 04/12/2024 | 05/02/2024 |  |  |  |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447867 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 76-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that following impella cp implant, the patient developed a spontaneous cardiac tamponade and began to decompensate. cardiopulmonary resuscitation was performed, and venoarterial extracorporeal membrane oxygenation was initiated to counteract the condition. upon arrival to the operating room, the patient¿s chest was opened and the impella pigtail was found to have perforated the left ventricle posterior wall. despite the subsequent surgical intervention, the patient passed away. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torqueing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation into the ventricle perforation issue has been completed since the original report was filed. the device was returned for investigation. after reviewing the clinical information, data logs, and returned pump, it was determined that the cause of the ventricle perforation was improper positioning of the device.
27157 | 1220648-2024-10685 | 19231294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19231294&pc=OZD | 04/03/2022 | 03/26/2025 |  | 05/02/2024 |  |  | 04/03/2022 | 05/02/2024 |  |  | 04/30/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022070663 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/03/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient then developed oozing at the impella cp access site and on the impella dressing. the patient then had a drop of platelets from 141 to 76. there was a slight difficulty in maintaining therapeutic activated clotting time range, therefore, systemic heparin was increased from 2500 to 2900. there was a discussion with the care team of the patient being heparin-induced thrombocytopenia (hit) positive or having hemolysis, as there was a lack of increase in hemoglobin despite the patient receiving 7 units of blood. however, the patient remained on impella support. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). this report is being filed as part of a retrospective review of historical records.
27158 | 1220648-2024-10686 | 19231556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19231556&pc=OZD | 04/25/2024 | 06/13/2024 |  | 05/02/2024 |  |  | 04/25/2024 | 05/02/2024 |  |  |  |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453259 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The user facility reported a 54-year-old female with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that they were unable to advance the impella 5.5 into the patient's axillary artery beyond the graft anastomosis due to the patient¿s anatomy. the medical staff attempted the implant multiple times, using multiple wires. the implant was eventually aborted, and the medical staff implanted an intra-aortic balloon pump instead. no patient harm has been reported. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition.
27159 | 1220648-2024-10687 | 19231806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19231806&pc=OZD | 04/01/2024 | 06/05/2024 |  | 05/02/2024 |  |  | 04/02/2024 | 05/02/2024 |  |  | 03/31/2025 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024329695 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  |  |  | Death | Yes | Yes | Failure to Advance | [6 empty] | Perforation of Vessels | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2024 |  |  |  |  | The complainant reported a patient (race unknown) in acute myocardial infarction/cardiogenic shock was to be emergently implanted with an impella cp device for mechanical circulatory support. the impella cp was unable to be advanced due to a vessel dissection at the femoral artery. the procedure was discontinued as the patient's condition did not improve due to the coronary situation. the patient required resuscitation prior to procedure. the patient expired in the procedure area. abiomed does not have enough information to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the vascular damage - peripheral vessel was determined to be patient condition related since the patient was a bailout and had to be resuscitated for some time beforehand. the root cause of the delivery issues was patient condition related because due to the dissection, the pump could not longer be advanced. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high risk pci in the warnings & cautions section. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
27160 | 1220648-2024-10688 | 19232321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19232321&pc=OZD | 04/26/2024 | 07/10/2024 |  | 05/02/2024 |  |  | 04/26/2024 | 05/02/2024 |  |  |  |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025476197 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 30 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/26/2024 |  |  |  |  | The user facility reported a 30-year-old patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was inserted before priming was completely completed, so it had to be reinserted. the 0.018 inch guidewire became fully bent and it was difficult to reinsert. a new 0.018 inch guidewire was used. there was no patient harm. | [5 empty] | The investigation for the reported product damage has been completed. no product was returned for evaluation. clinical details states that the impella was inserted before priming was complete. 0.018 inch guidewire became completely bent. the root cause of the delivery issue was use-related due to pump reinsertion following improper technique. added- d6b, h6 impact code 4629 f6/f8 should have been left blank in the initial report. g1 reporting contact fax number missing. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27161 | 1220648-2024-10689 | 19233211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19233211&pc=OZD | 06/09/2023 | 07/16/2024 |  | 05/02/2024 |  |  | 04/09/2024 | 05/02/2024 |  |  | 04/30/2025 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024338142 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | We received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that severe hypotension occurred with a definite relationship to the impella 5.5 used on this patient. | [5 empty] | The investigation for the reported hypotension has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hypotension could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿
27162 | 1220648-2024-10690 | 19233457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19233457&pc=PYX | 04/23/2024 | 06/13/2024 |  | 05/02/2024 |  |  | 04/23/2024 | 05/02/2024 |  |  |  |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Literature |  |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Thrombosis/Thrombus | Tricuspid Valve Insufficiency/ Regurgitation | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | A publication was shared with abiomed ¿ clinical outcomes in acute right ventricular failure with percutaneous right ventricular assist devices: impella rp and protek duo. the publication was from the american college of cardiology in 2021. the publication contained a retrospective review of impella rp and protek duo. the noted complications in impella rp patients included: cannula repositioning, cannula thrombosis, excessive hemolysis, and severe tricuspid regurgitation. | [5 empty] | D.4, d.6a, d.6b and h.4 catalog number, serial number, lot number, expiration date, unique identifier (udi) #, implant date, explant date and device manufacture date are unknown. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. rp with smart assist system with automated impella controller. section: contraindications (united states). ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ section: warnings & cautions: cautions ¿patients with tricuspid or pulmonary valve stenosis or insufficiency, and patients with prosthetic tricuspid or pulmonary valves, may be compromised by the use of the impella rp flex with smartassist system catheter.¿ rp with smart assist system with automated impella controller. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. " in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the reported hemolysis, heart valve damage, thrombus, and the positioning issue has been completed. pump was not returned for investigation. the cause of the heart valve damage issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the thrombus could not be determine without sufficient clinical information.
27163 | 1220648-2024-10691 | 19233620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19233620&pc=OZD | 04/23/2024 | 06/13/2024 |  | 05/02/2024 |  |  | 04/23/2024 | 05/02/2024 |  |  |  |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440316 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | The user facility reported a 53-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported the patient was on impella cp support due to having cardiogenic shock. while on support, the patient had access site bleeding. a quick clot was applied, manual pressure was performed and two units of packed red blood cells were provided. additionally, there were suction events. two units of blood cells, three albumin and precedex were started to help with the low pump flow. furthermore, the patient had renal failure on support. one unit of blood products was given to the patient. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the major access site bleeding, the renal failure, and the low pump flow was not determined because no product was returned, and not enough clinical information was given.
27164 | 1220648-2024-10692 | 19235708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19235708&pc=OZD | 03/31/2022 | 05/02/2024 |  | 05/02/2024 |  |  | 03/31/2022 | 05/02/2024 |  |  | 12/31/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178274 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 03/31/2022 |  |  |  |  | The user facility reported a 79-year-old male patient was implanted with an impella cp for mechanical circulatory support. while on support, a large hematoma was noted at the impella site. manual pressure was held, and the hematoma was expressed. the access site continued to bleed without pressure. systemic heparin was withheld. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. this report is being filed as part of a retrospective review of historical records.
27165 | 1220648-2024-10693 | 19235718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19235718&pc=OZD | 03/31/2022 | 05/02/2024 |  | 05/02/2024 |  |  | 03/31/2022 | 05/02/2024 |  |  | 12/31/2023 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172994 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 03/31/2022 |  |  |  |  | The user facility reported a 68-year-old male patient was implanted with an impella cp for mechanical circulatory support. while on support, perfusion of the left leg was very slow. it was noted that the patient was slowly weaned off impella. a decision was made to remove the device due to concern of potential leg ischemia. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
27166 | 1220648-2024-10694 | 19235732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19235732&pc=OZD | 09/21/2022 | 05/02/2024 |  | 05/02/2024 |  |  | 09/21/2022 | 05/02/2024 |  |  | 07/31/2024 |  | 05/02/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231536 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/21/2022 |  |  |  |  | The user facility reported a 68 year old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced continues runs of vt (ventricular tachycardia) and bleeding in the chest. systematic heparin was stopped due to the bleeding. the physician also experienced positioning issues. the patient was combative overnight, and the impella was pulled back across the aortic valve. the physician was unable to view the impella on the ultrasound and repositioned blindly. he tried to re-advance across valve unsuccessfully and decision was made to exchange the impella cp for an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27167 | 1220648-2024-10616 | 19219788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19219788&pc=OZD | 05/10/2022 | 03/28/2025 |  | 05/01/2024 |  |  | 05/10/2022 | 05/01/2024 |  |  | 01/31/2024 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184103 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2022 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced minimal urinary output and hematuria. the patient received bumex and was placed on dialysis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
27168 | 1220648-2024-10617 | 19220014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220014&pc=OZD | 03/18/2024 | 06/07/2024 |  | 05/01/2024 |  |  | 04/03/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443202 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 20 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | Thrombocytopenia | [10 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 04/03/2024 |  |  |  |  | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured recurring thrombocytopenia with a "possible" relationship to the impella device used: ¿the patient developed reportable thrombocytopenia on (b)(6) 2024. admission platelets 272 k/ul, platelets (b)(6) 2024 49 k/ul, lowest platelet count 39 k/ul ((b)(6) 2024), and recovered above reportable threshold (b)(6) 2024 (57 k/ul). the patient experienced recurring thrombocytopenia (b)(6) 2023 with platelets as low as 12 k/ul. | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured recurring hemolysis with a "definite" relationship to the impella device used: ¿the patient developed recurring hemolysis. admission total hemoglobin (b)(4), admission bilirubin 2.4 mg/dl (ldh not done). lowest total hemoglobin (b)(4) ((b)(6) 2024) peak ldh 38298 u/l ((b)(6) 2024), peak bilirubin (b)(4) mg/dl ((b)(6) 2024). plasma free hemoglobin peak while on impella support 45.0 mg/dl ((b)(6) 2024).¿ the patient outcome is "unknown" at this time. | We received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that ventricular tachycardia occurred with a definite relationship to the impella 5.5 used on this patient. |  |  |  | B5 and h6 has been updated as per additional information received. | The investigation for the reported ventricular tachycardia (vt), hemolysis, and thrombocytopenia has been completed. the device was not returned for evaluation. clinical details were insufficient to determine the root cause of the thrombocytopenia and the vt. however, based upon clinical details, the root cause of the hemolysis was patient condition related and unrelated to impella since the hemolysis began to develop five days after impella removal. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist for use in the potential adverse events (united states) section. d4-updated model number | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿
27169 | 1220648-2024-10622 | 19220213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220213&pc=OZD | 04/13/2024 | 06/03/2024 |  | 05/01/2024 |  |  | 04/13/2024 | 05/01/2024 |  |  | 10/31/2025 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445628 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient exhibited access site bleeding. the bleeding was treated with temporary stop of heparin and a purse string suture. during explant, the patient experienced vt which resolved on its own. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major and the vf/vt/af/at (ventricular tachycardia) issues has been completed since the original report was submitted. the root cause of the access site bleeding issue was most likely use related, the bleeding was treated with temporary stop of heparin and a purse string suture as per the clinical description. as the necessary information was not provided, the root cause of the vt/vf was not determined.
27170 | 1220648-2024-10623 | 19220216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220216&pc=OZD | 04/04/2022 | 05/01/2024 |  | 05/01/2024 |  |  | 04/04/2022 | 05/01/2024 |  |  | 02/28/2023 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022056003 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/04/2022 |  |  |  |  | The complainant reported a patient was taken to the cardiac catheterization laboratory for impella placement and that the physician could not get the impella to calm the left ventricle and the patient was shocked five times. the physician attempted to resolve this by placing the impella deeper and shallower, but the patient would still go into ventricular fibrillation. the decision was made to remove the impella due to the size and irritated left ventricle. the physician reported upon follow-up that the patient was chest pain free. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27171 | 1220648-2024-10624 | 19220286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220286&pc=OZD | 04/18/2024 | 07/15/2024 |  | 05/01/2024 |  |  | 04/18/2024 | 05/01/2024 |  |  | 05/31/2025 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024358334 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp support there were problems with the purge system. there was dripping at the yellow luer lock. the purge system was changed but the issue did not resolve. the purge pressure was 0 and the purge flow was >30. a purge filter bypass was successfully performed and the purge values went back to normal. the patient was stable and the pump parameters were normal. | Additional information was provided that stated that the patient experienced oozing at the access site with a hemoglobin drop from 7g/dl to 5.7g/dl. sutures were placed but the bleeding continued and the patient received blood products the next day. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the reported access site bleed and purge leak has been completed. pump was returned for evaluation. upon visual inspection, a purge sidearm bypass was present on the pump. the original sidearm was not returned for inspection. clinical details noted that purge pressure was 0 and it was not clear where the leak was coming from on the sidearm. sidearm bypass was performed and it resolved the issue. upon visual inspection of the repositioning sheath, no damage was found. clinical details noted that impella groin site was found to be oozing and was attempted to resolve with sutures without success. pressure bandages were applied and resolved oozing. the root cause of the access site bleeding cannot be determined as no damage was found on the repositioning sheath and bleeding was able to be managed with pressure bandages. the root cause of the purge leak was most likely due to a damaged sidearm as the low purge pressure was resolved with sidearm bypass. the exact location of the leak was unknown since purge sidearm was not returned. the instructions for use for the impella cp with smart assist system in the following sections: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
27172 | 1220648-2024-10625 | 19220404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220404&pc=OZD | 03/23/2022 | 03/28/2025 |  | 05/01/2024 |  |  | 03/23/2022 | 05/01/2024 |  |  | 12/31/2023 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171624 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/23/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. upon impella explant, perclose device failed three times. the event was managed through balloon tamponade and manual pressure. patient tolerated the procedure well. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartaassist system. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27173 | 1220648-2024-10626 | 19220534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220534&pc=OZD | 03/08/2022 | 05/01/2024 |  | 05/01/2024 |  |  | 03/08/2022 | 05/01/2024 |  |  | 12/31/2023 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171677 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/08/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient developed a hematoma at the right groin, active bleeding noted. it was reported that the vascular surgeon then successfully removed pump, manta 14f was placed for vessel closure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27174 | 1220648-2024-10627 | 19220618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220618&pc=OZD | 04/21/2024 | 06/07/2024 |  | 05/01/2024 |  |  | 04/21/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025447944 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/21/2024 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced some bleeding from their access site during impella cp support. heparin was put on hold and a suture was added to the site to manage the condition. support was maintained throughout and following the intervention. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of access site bleeding is determined to be use issue because the issue was resolved by suturing site and achieved hemostasis. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
27175 | 1220648-2024-10628 | 19220673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220673&pc=OZD | 04/24/2024 | 05/28/2024 |  | 05/01/2024 |  |  | 04/24/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025445433 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella 5.5 support coughed when working with pt/ot and impella suctioned, and maps (transesophageal echocardiography) dropped. the patient was helped back to bed and began to slur speech. the patient became unresponsive and intubated. ct (computed tomography) head and chest with full neuro exam showed no acute neurological concerns. eeg (electrocardiogram) shows mild diffuse encephalopathy, and no cva (cerebral vascular accident). the patient was on systemic heparin at the time of the event. it is unknown if impella contributed to this event. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. d4-updated model number | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
27176 | 1220648-2024-10629 | 19220676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220676&pc=OZD | 03/16/2022 | 05/01/2024 |  | 05/01/2024 |  |  | 03/16/2022 | 05/01/2024 |  |  | 11/30/2023 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  |  |  | 2023164610 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/16/2022 |  |  |  |  | The user facility reported a 36-year-old patient needing mechanical circulatory support. it was reported that while patient was on impella cp, the patient experienced access site bleeding that was resolved with discontinuation of heparin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27177 | 1220648-2024-10630 | 19220720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220720&pc=OZD | 04/24/2024 | 06/04/2024 |  | 05/01/2024 |  |  | 04/24/2024 | 05/01/2024 |  |  | 12/31/2024 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024298256 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient developed bleeding in the right femoral artery post implant. bleeding was successfully managed with a pump exchange. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is for the impella cp with smart assist system.
27178 | 1220648-2024-10631 | 19220767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220767&pc=OZD | 03/21/2022 | 03/28/2025 |  | 05/01/2024 |  |  | 03/21/2022 | 05/01/2024 |  |  | 12/31/2023 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171673 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed fever which was resolved by changing antibiotics. additionally, atrial fibrillation (af) was noted and was able to be resolved by administering magnesium, however af returned temporarily and subsided when electrolytes were balanced. patient's condition remained stable throughout the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿¿ this report is being filed as part of a retrospective review of historical records. | Infection/sepsis. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: -patient symptoms and medical history obtained by the treating clinician. -physical examination findings. -results of special investigations: laboratory test results & imaging studies. -microbiological culture and sensitivity results. -response to previous treatments and therapies. -any relevant epidemiological information or exposure history. -concomitant devices in use. -preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. -care of access site. because this evidence has not been provided, the root cause of the infection cannot be determined. cardiac arrythmia to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant). ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalitie. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records.
27179 | 1220648-2024-10632 | 19220803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220803&pc=OZD | 04/06/2022 | 05/01/2024 |  | 05/01/2024 |  |  | 04/06/2022 | 05/01/2024 |  |  | 01/31/2024 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184807 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/06/2022 |  |  |  |  | The complainant reported a patient, noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support after pulsatility dropped, it was noted that the impella cp was in too deep and needed to be pulled back. the physician made several attempts to find a good position. however, the patient¿s anatomy was too small, and the outlet was only passed halfway through the aortic arc, the patient¿s ventricle was reported to be small. the patient remained stable throughout the procedure and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27180 | 1220648-2024-10633 | 19220846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220846&pc=OZD | 04/24/2024 | 06/30/2024 |  | 05/01/2024 |  |  | 04/24/2024 | 05/01/2024 |  |  | 10/31/2025 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424901 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Inaccurate Delivery | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | The complainant reported a patient undergoing cardiac surgery was to be implanted with an impella cp device for mechanical circulatory support. during attempted delivery, positioning was lost in the left ventricle and the impella had to be removed. the second pump was subsequently placed using transesophageal echocardiography and fluoroscopy. there was no patient harm. | [5 empty] | The investigation for the delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was determined to be use issue because the guide wire was placed into the left ventricle and during the insertion of the impella cp pump, they lost the position and removed the pump and they were able to get the positioning later when they tried the insertion of an impella 5.5. hence, the root cause was determined to be use issue. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves, ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ a.5 revised ethnicity as previously entered incorrectly. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 medical device problem code and component code were revised in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
27181 | 1220648-2024-09871 | 19220879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220879&pc=OZD | 04/01/2024 | 06/05/2024 |  | 05/01/2024 |  |  | 04/10/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025463839 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | The user facility reported a 69-year-old male patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the patient had bleeding from a previous impella device¿s access site. anticoagulation was withheld to control the bleeding, and the patient developed deep vein thrombosis in their right arm. this was not treated at the time due to the bleeding concerns. the patient remains on support. | [5 empty] | The investigation has been completed. the data logs confirmed the failure mode and clinical narrative. the logs showed after the pump started, motor current spiked followed low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. product was not returned for review. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 reporting contact facility name was inadvertently left off the previously submitted report and was added. h.6 added code 925 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
27182 | 1220648-2024-09870 | 19220880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220880&pc=OZD | 04/09/2024 | 06/06/2024 |  | 05/01/2024 |  |  | 04/09/2024 | 05/01/2024 |  |  | 04/30/2025 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330400 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that an implanted impella 5.5 pump began to experience high purge pressure during patient support. at first, a tissue plasminogen activator protocol was run through the purge lines to counteract the increase in pressure. upon reversing the upwards trend in pressure, the tpa was switched to heparin in the purge and the issue was considered to be resolved. there was no patient harm. | [5 empty] | The investigation into the high purge pressure issue has been completed since the original report was filed. the pump was returned for investigation. there was no biomaterial found around the impellers or the outflow cage. the pump was ran and the failure could not be reproduced. the data logs were returned and investigated. the data logs showed a gradual rise in the purge pressure from 400 mmhg to over 1050 mmhg in four hours leading to "high purge pressure" and "purge system blocked" alarms. tissue plasminogen activator (heparin) was added at this point and the purge pressure started to decrease and returned to around 500 mmhg. as per the data log review and product evaluation, the root cause of the high purge pressure issue was biomaterial. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 medical device problem code revised from the initial submission in accordance with updated procedures. codes 4114 (device not returned) and (3233) results pending completion of investigation were inadvertently omitted from the previously submitted report and have been added to this report. the product has been returned since then and the investigation has been completed. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27183 | 1220648-2024-10329 | 19220881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220881&pc=OZD | 05/24/2022 | 03/28/2025 |  | 05/01/2024 |  |  | 05/24/2022 | 05/01/2024 |  |  | 12/31/2023 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180657 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient¿s urine was noted to be bright red, and lab samples were hemolyzed. an initial echocardiogram showed the device to be too deep at 4.5-5 cm. the physician noted that the device was not locked nor attached to the reposition sheath. a few attempts were made to reposition, but the physician didn¿t want to dislodge the device completely. the pump was later found to be entirely outside the left ventricle (lv), and the physician was unable to re-advance it into the lv. the device was then explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27184 | 1220648-2024-09967 | 19220882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220882&pc=OZD | 03/27/2024 | 05/08/2024 |  | 05/01/2024 |  |  | 03/27/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025461193 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient developed bleeding from their access site during impella 5.5 support. anticoagulation was put on hold in response to the condition. the bleeding was noted to have resolved and support with the implanted pump was maintained. | [5 empty] | The investigation into the access site bleeding ¿ major issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was determined to be anticoagulation management because the anticoagulation not ordered and kept on hold due to bleeding concerns at the insertion site the previous night to achieve hemostasis. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27185 | 1220648-2024-10197 | 19220883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220883&pc=OZD | 05/19/2022 | 09/21/2024 |  | 05/01/2024 |  |  | 05/19/2022 | 05/01/2024 |  |  | 01/31/2024 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191222 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  |  | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | Ventricular Fibrillation | No Clinical Signs, Symptoms or Conditions | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/19/2022 |  |  |  |  | The user facility reported a 63-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported the patient was presented with chest pains (cp) and shortness of breath (sob). the patient agreed to ventricular septal defect (vsd) repair and coronary artery bypass graft (cabg) placement with ecpella if necessary, after bypass procedure. the medical professional pulled the impella into the descending aorta to avoid cross clamping and the bypass surgery was performed as planned. after the bypass procedure, the medical professional was able to float the impella back across the valve for optimal positioning. the patient¿s oximetry statistics began to drop and suction greater than pump speed level p7 were observed. ecpella was successful and the cp was set to p2 for unloading. overnight the patient went into ventricular fibrillation (vfib) and suction alarms were reported by the nurse on call. the impella cp was repositioned due to the patient experiencing hemolysis and volume replacement was given. the patient was weaned from the impella, and subsequently expired on support. | The patient went into ventricular fibrillation (vfib) and suction alarms were reported by the nurse on call. the impella cp was repositioned due to the patient experiencing hemolysis and volume replacement was given. the patient was weaned from the impella, and subsequently expired on support. use of the device cannot conclusively be associated with the outcome. patient's peri-procedural condition was critical. the patient's peri-procedural condition was critical. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-10197 was provided.
27186 | 1220648-2024-10195 | 19220884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220884&pc=OZD | 05/24/2022 | 03/28/2025 |  | 05/01/2024 |  |  | 05/24/2022 | 05/01/2024 |  |  | 07/31/2023 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022109804 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/24/2022 |  |  |  |  | The user facility reported a patient on impella cp support due to acute myocardial infarction and cardiogenic shock. during implantation, a .018 wire was inserted into the pigtail, but as the pigtail was being backed out, the wire slipped back. the physician advanced the wire, but it was suspicious of being behind the papillary. the pigtail was fully removed and the impella cp backloaded onto the .018 wire and advanced into place. however, immediately it could be seen that the device was in the papillary. the .018 wire was removed and the impella cp was gently pulled back, but popped into the aorta. the physician attempted to advance the device without a wire but it would not cross. it was decided to remove the impella cp. the staff also noted that the patient did not have doppler pulses in the left foot and a lack of doppler pulses indicates limb ischemia. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27187 | 1220648-2024-10196 | 19220885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220885&pc=OZD | 04/06/2024 | 06/12/2024 |  | 05/01/2024 |  |  | 04/06/2024 | 05/01/2024 |  |  | 10/31/2025 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024415164 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 71-year-old female patient on impella cp support had access site bleed leading to the pump being explanted. the patient was on anticoagulants, and the patient's partial thromboplastin time at the time of the bleed was 120. hemostasis was able to be achieved by correcting sheath and suturing. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was determined to be use issue because hemostasis was able to be achieved by correcting sheath and suturing. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27188 | 1220648-2024-10620 | 19220886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220886&pc=OZD | 02/20/2024 | 05/01/2024 |  | 05/01/2024 |  |  | 04/05/2024 | 05/01/2024 |  |  | 10/31/2025 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422738 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/05/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to transplant. during impella 5.5 support, a high purge pressure alarm was encountered. the purge cassette was replaced, but the issue remained. a tissue plasminogen activator protocol was subsequently initiated and the purge pressure alarms subsided. there was no patient harm. the patient was successfully weaned and the impella 5.5 was explanted. | [5 empty] | The investigation into "high purge pressure" has been completed. the device was not returned for evaluation. clinical review: purge pressure high alarms reported on 04/05/24. no purge line kinks observed. tpa was added to purge but unknown if this action resolved the purge blockage. data log review: data logs were reviewed, and several high purge pressure and purge system blocked alarms were noted on 04/05/24. high purge pressure rises gradually within 20 to 30 minutes and later starts to resolve after 5 hours. the purge pressure gradually reduces in 4 hours to normal range resolving purge issues. no motor current spike observed due to likely gradual biomaterial ingestion. the cause of the high purge pressure issue was most likely due to biomaterial per log and clinical review. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
27189 | 1220648-2024-10621 | 19220887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220887&pc=OZD | 03/22/2022 | 03/28/2025 |  | 05/01/2024 |  |  | 03/22/2022 | 05/01/2024 |  |  | 12/31/2023 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023175045 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, concerns of possible clot in the impella cannula was noted. impella cp was explanted and replaced with a new one. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical record. | When making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions. ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts. ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
27190 | 1220648-2024-10619 | 19220888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220888&pc=OZD | 10/30/2024 | 05/20/2024 |  | 05/01/2024 |  |  | 04/04/2024 | 05/01/2024 |  |  | 07/31/2025 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024388875 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | We received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted ventricular tachycardia occurred with a possible relationship to the impella cp used on this patient. the patient was administered mg and vt subsided. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation into the vf/vt/af/at (ventricular tachycardia) issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary information was not provided, the root cause of the vt/vf was not determined. impella rp system with automated impella controller & impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
27191 | 1220648-2024-10618 | 19220889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220889&pc=OZD | 11/02/2023 | 07/26/2024 |  | 05/01/2024 |  |  | 04/04/2024 | 05/01/2024 |  |  | 06/30/2025 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024376361 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Renal Failure | Tachycardia | [10 empty] | Hospitalization | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/04/2024 |  |  |  |  | Additional information received from our medical safety team via loqi (abiomed's long-term outcome and quality indicator impella registry), indicating an allegation of recurrent episodes of svt (supra ventricular tachycardia) with the impella cp device used on this patient. patient was hospitalized to address the issue and is stable. | We received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), indicating an allegation of worsening respiratory failure, recurrent atrial fibrillation with rvr, renal failure and worsening cardiogenic shock with the impella cp device used on this patient. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | Additional information received from the clinical study coordinator indicated that the event is no longer related to the impella device. no further reports will be forthcoming.
27192 | 1220648-2024-10634 | 19220906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220906&pc=OZD | 03/23/2022 | 03/26/2025 |  | 05/01/2024 |  |  | 03/23/2022 | 05/01/2024 |  |  | 12/31/2023 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171078 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/23/2022 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, an external femoral to femoral bypass was placed for leg perfusion. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition any concomitant medication. vascular size or variations thereof detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
27193 | 1220648-2024-10635 | 19220919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220919&pc=OZD | 03/22/2022 | 05/01/2024 |  | 05/01/2024 |  |  | 03/22/2022 | 05/01/2024 |  |  | 12/31/2023 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180195 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/22/2022 |  |  |  |  | The complainant reported a patient, race unknown, was implanted with an impella cp for continued support of left ventricular unloading. while on support, the patient developed limb ischemia. it was reported no pulsers were noted to the right lower extremity and a doppler revealed sluggish flow. the physician performed an external femoral bypass using 5 french sheaths, which reportedly improved the lower extremity¿s color and temperature. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿. this report is being filed as part of a retrospective review of historical records.
27194 | 1220648-2024-10636 | 19220925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220925&pc=OZD | 04/24/2024 | 06/03/2024 |  | 05/01/2024 |  |  | 04/25/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447770 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient's left lower extremity became mottled with no appreciable pulse during impella cp support. a distal perfusion catheter was inserted to manage the condition and the patient's symptoms improved. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. d4-updated model number
27195 | 1220648-2024-10637 | 19220988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220988&pc=OZD | 04/25/2024 | 06/11/2024 |  | 05/01/2024 |  |  | 04/25/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025485740 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The user facility reported a 65-year-old male with cardiomyopathy was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that complications were encountered during attempted delivery of an impella 5.5 pump. it was documented that despite multiple wire strategies, the pump was unable to pass the origin of the innominate artery due to a circumferential calcified lesion. the implant was aborted, and plans were made to implant an impella cp pump for patient support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: caution: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist for use during cardiogenic shock. in accordance with updated procedures, sections d6 have been revised from the initial submission.
27196 | 1220648-2024-10638 | 19220994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19220994&pc=OZD | 04/25/2022 | 05/01/2024 |  | 05/01/2024 |  |  | 04/25/2022 | 05/01/2024 |  |  | 12/31/2023 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180197 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/25/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had a left femoral hematoma at the impella cp insertion site. the staff held pressure until a femstop was obtained. there was no heparin in the purge due to the hematoma, systemic heparin was continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27197 | 1220648-2024-10639 | 19221107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19221107&pc=OZD | 04/25/2024 | 06/13/2024 |  | 05/01/2024 |  |  | 04/25/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025485017 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The user facility reported a 75-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that after the procedure, the impella was removed and the first perclose tore and was unable to deploy, the second perclose was also unsuccessful. a balloon tamponade was performed, and the patient was taken to vascular surgery to close the access site. additionally, there was a clot noted. a thrombectomy was performed. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major and the thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was determined to be use issue as the first perclose tore and was unable to be deployed, second perclose was also unable to be deployed. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
27198 | 1220648-2024-10640 | 19221161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19221161&pc=OZD | 04/30/2022 | 03/26/2025 |  | 05/01/2024 |  |  | 04/30/2022 | 05/01/2024 |  |  | 08/31/2023 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127144 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the right lower extremity was found to be dusky below the knee. the physician felt that both extremities were finished and cool due to being clamped down from the bicarbonate drip. the device began to be explanted through a 14fr sheath, but the sheath could not be withdrawn even after being flushed. an 8fr dilator was placed just inside the 14fr hemostatic valve to attempt to aspirate a clot out of the proximal end of the sheath, but this was unsuccessful. the physician decided to leave the device in place and allow the vascular surgeon to explant it and clear any remaining clots. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia: when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. thrombus: when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions. ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts. ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
27199 | 1220648-2024-10641 | 19221213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19221213&pc=OZD | 04/22/2024 | 06/11/2024 |  | 05/01/2024 |  |  | 04/22/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442871 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Difficult to Remove | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. however, the impella pump was going to be exchanged due to "purge system blocked" alarm that did not resolve with tissue plasminogen activator. upon attempting to explant the pump, it could not be removed. the axillary artery was ¿cinched¿ beyond the graft anastomosis causing an impingement of the vessel and prohibiting the impella pump explant. the impella pump remained in the descending aorta on p-level p-2 and placement monitoring disabled until the following day that the patient was taken for intervention. the impella pump was eventually removed from ascending aorta through the graft with a decision to not place a new impella device. the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the reported high purge pressure (hpp) issue and the removal issues has been completed. pump was not returned for investigation. data logs were returned and investigated. the hpp failure was reported there were no high purge pressure alarms seen in the data logs to correspond to the failure mentioned. as per the data log review, there was no high purge pressure issue found during the case and insufficient data was returned. the root cause of the removal issue was most likely use related to improper implant technique. the root cause of the saturated flow issue was not determined since product was not returned and data log review is inconclusive.
27200 | 1220648-2024-10642 | 19221292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19221292&pc=OZD | 06/28/2022 | 03/26/2025 |  | 05/01/2024 |  |  | 06/28/2022 | 05/01/2024 |  |  | 11/30/2023 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151467 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 93 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a (b)(6) female was implanted with an impella cp for mechanical circulatory support. it was reported that access site bleeding and a hematoma. the patient was given a blood transfusion as intervention. following explant of the device, it was reported that infection of the access site occurred. no additional information provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist: section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27201 | 1220648-2024-10643 | 19221497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19221497&pc=PYX | 07/01/2018 | 07/05/2024 |  | 05/01/2024 |  |  | 04/23/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Positioning Failure | Detachment of Device or Device Component | [5 empty] | Foreign Body In Patient | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | This report will represent the patient that the impella rp device fractured on removal and all components of the device were removed percutaneously without incident. manufacture device report 1220648-2024-12947 represents the patient that had to return to the catheterization lab for repositioning. | A literature review was performed, and a publication was found inclusive of a impella rp review of cases from 2007 to 2018 published/authored by the team at new york presbyterian columbia. in the review, two complaints were noted. there was a patient who with impella rp removal, the impella rp fractured, and the team was successful in removing all parts of the pump from native patient vasculature. another patient had to be returned to the catheterization lab for impella rp malpositioning issues. |  |  |  |  | The investigation for the pigtail detachment has been completed. the impella device was not returned for evaluation. the root cause of the cannula/ pigtail detachment issue was not determined since the product was not returned and insufficient clinical information and case information was provided. literature citation: coromilas, e. j., takeda, k., ando, m., cevasco, m., green, p., karmpaliotis, d., kirtane, a., topkara, v. k., yuzefpolskaya, m., takayama, h., naka, y., burkhoff, d., colombo, p. c., & garan, a. r. (2019). comparison of percutaneous and surgical right ventricular assist device support after durable left ventricular assist device insertion. journal of cardiac failure, 25(2), 105¿113. https://doi.org/10.1016/j.cardfail.2018.12.005 b.1 product problem was selected on the initial manufacturer device report 1220648-2024-10643 and should not have been b.5 revised to reflect only one patient since the initial manufacturer device report 1220648-2024-10643 was submitted. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-10643 was submitted. h.6 codes 1158 and 925 were reported incorrectly on manufacturer device report 1220648-2024-10643. h.10 a.1, a.4 and a.5 are unknown statement and literature citation were inadvertently omitted from the initial manufacturer device report 1220648-2024-10643. attachments: publication was inadvertently omitted from the initial manufacturer device report 1220648-2024-10643. | D.14 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. h.4 device manufacture date is unknown. d.6a and d.6b implant and explant dates are unknown. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella rp smart assist system with automated impella controller. section: warnings & cautions: warnings. ¿handle with care. the impella rp with smartassist system catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
27202 | 1220648-2024-10644 | 19221565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19221565&pc=OZD | 09/16/2022 | 03/26/2025 |  | 05/01/2024 |  |  | 09/16/2022 | 05/01/2024 |  |  | 05/31/2024 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228494 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/16/2022 |  |  |  |  | The user facility reported a 67-year-old female on impella cp support due to vfib (ventricular fibrillation) and vtach (ventricular tachycardia). while on support minimal bleeding was noted. the impella cp access site was maintained for va ecmo (venoarterial extracorporeal membrane oxygenation). the purge solution was switched to bicarbonate. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27203 | 1220648-2024-10645 | 19221585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19221585&pc=OZD | 03/22/2022 | 03/26/2025 |  | 05/01/2024 |  |  | 03/22/2022 | 05/01/2024 |  |  | 12/31/2023 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177764 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Atrial Flutter | Hemolysis | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/24/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced a couple of episodes of atrial fibrillation with flutter. the patient was shocked twice, and amiodarone bolus and drip were given. it was reported that the patient's laboratory results were reported as hemolyzed, and two units of packed red blood cells were administered. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27204 | 1220648-2024-10646 | 19221610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19221610&pc=OZD | 03/29/2022 | 05/01/2024 |  | 05/01/2024 |  |  | 03/29/2022 | 05/01/2024 |  |  | 09/30/2023 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022140295 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/29/2022 |  |  |  |  | The user facility reported that a 48 year old male patient was taken to the cardiac catheterization laboratory (ccl) for impella implant prior to a coronary artery bypass graft (cabg) being placed. while on impella support, it was reported that a hematoma was observed at the access site. the impella cp was explanted three (3) hours after the patient¿s admission to the medical facility. the complication was managed utilizing percutaneous transluminal angioplasty and a covered stent to the right axillary artery. it was reported that the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27205 | 1220648-2024-10647 | 19221664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19221664&pc=OZD | 09/16/2022 | 03/26/2025 |  | 05/01/2024 |  |  | 09/16/2022 | 05/01/2024 |  |  | 06/30/2024 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235844 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/16/2022 |  |  |  |  | The complainant reported a patient undergoing electrophysiology/ablation was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced bleeding. there was oozing at the groin site and the site was redressed, which solved the issue with some success. heparin was dropped. it was later reported that the patient experienced a substantial groin bleed which required 2.5 hours of manual pressure. a stitch was applied with a pressure dressing to resolve the issue, and the patient remained on impella cp support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27206 | 1220648-2024-10648 | 19226191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19226191&pc=OZD | 04/24/2024 | 06/13/2024 |  | 05/01/2024 |  |  | 04/24/2024 | 05/01/2024 |  |  | 06/30/2025 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024427768 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | The user facility reported a 58-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, there was a sudden spike in motor current followed by a pump stop. the staff was unable to restart the pump. the pump was removed and replaced with an impella 5.5. after removal, there was a significant clot burden found inside the inlet of the pump. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smart assist system. section: warnings & cautions: warnings: section: pre-support evaluation; section: impella cp with smart assist catheter insertion; section: axillary insertion of the impella cp with smart assist catheter; section: use of impella with transcatheter aortic valves; ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smart assist for use during cardiogenic shock and high risk percutaneous coronary intervention (pci). section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smart assist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smart assist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk pci. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into the reported issues has been completed since the original report was filed. the pump was returned for investigation. there was heavy biomaterial found throughout the cannula from inlet to outlet and above the impeller. multiple attempts were seen in the logs to restart the pump but no success. per the product investigation and data logs, the root cause of the pump stop issue was biomaterial found lodged in the cannula and around the impeller head.
27207 | 1220648-2024-10651 | 19226192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19226192&pc=OZD | 04/24/2024 | 06/07/2024 |  | 05/01/2024 |  |  | 04/24/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1811664 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | The user facility's biomedical department released the automated impella controller and observed an alarm of "unexpected controller shutdown" on startup screen. there was no harm or injury noted as no patient use at the time. | [5 empty] | The investigation into the aic - a/c power issues has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. it was discovered that the console had not been properly shut down prior to shipment from abiomed, which resulted in ¿unexpected shutdown¿ alarm when the console was powered on next time at the customer¿s location. this was not an indication of equipment failure, and the device operates within specification. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27208 | 1220648-2024-10650 | 19226193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19226193&pc=PYX | 04/24/2024 | 06/12/2024 |  | 05/01/2024 |  |  | 04/24/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025479861 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella rp flex device for mechanical circulatory support. it was reported that the patient had an attempted impella rp flex implant and there were issues with removing the guidewire. there was significant resistance met with the guidewire not being able to be removed from the impella. the impella rp and the wire had to be removed from the patient. when the impella was outside the body, the wire was removed from the impella with considerable effort and resistance. the staff confirmed the correct wire was used. a new impella was opened and implanted without issue. the guidewire was found to have a partial fracture. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (detached inflow/outflow - great vessel) has been completed since the original report was submitted. the device was not returned for investigation. per the clinical information provided, the patient was obese and serial dilation was done pre pump implant. therefore, the root cause of the product damage (guidewire damage) issue was not determined since product was not returned for evaluation and insufficient clinical information was provided.
27209 | 1220648-2024-10649 | 19226195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19226195&pc=OZD | 04/24/2024 | 06/30/2024 |  | 05/01/2024 |  |  | 04/24/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025458141 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | The user facility reported a 77-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp insertion had a complication with the patient¿s anatomy. the impella sheaths were exchanged due to vessel calcification. once the impella was placed, the flows would not reach above 1.5 liters despite position check and volume given. the impella cp was removed. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk. percutaneous coronary intervention. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation has been completed since the original report was filed. the pump was returned for investigation. visual inspection found a kink on the cannula. no other issues were found on the rest of the pump. the pump ran without issue under bench test. data logs found the low flow when the pump was started and remained thru the rest of the case that triggered auto suction response and suction alarms. the root cause of low flow was kinked cannula. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
27210 | 1220648-2024-10652 | 19226196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19226196&pc=OZD | 04/24/2024 | 06/13/2024 |  | 05/01/2024 |  |  | 04/24/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025444355 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Unspecified Infection | Sepsis | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | The user facility reported a 42-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. impella 5.5 support was ongoing for a patient admitted with cardiogenic shock/acute myocardial infarction and cardiomyopathy. the labs were elevated, and the team alleged there was hemolysis despite urine color not tinging. no hematuria was noted. the team made the decision to explant and replaced the impella 5.5 due to the hemolysis and signs of infection/sepsis. a new pump was placed and remains on while awaiting transplant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella 5.5 with smartassist for use during cardiogenic shock. section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the hemolysis and infection has been completed. the pump was returned for investigation. a small free clot was found in the cannula near the inlet cage. no other physical abnormalities were observed. the pump ran as expected during the test. according to the clinical report, the device was explanted due to rapidly rising ldh levels after it was repositioned. the device was moved 1 cm and hemolysis was suspected with clear yellow urine. despite the surgical team's advice, the doctor refused to pull the device back and left it deep. data logs show several impella position unknown alarms and suction alarms during the hemolysis event. the cause of the hemolysis issue was most likely the improper positioning of the device. the root cause of the infection could not be determined without additional clinical details. added- d9 device return date. impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27211 | 1220648-2024-10654 | 19226197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19226197&pc=OZD | 04/24/2024 | 07/17/2024 |  | 05/01/2024 |  |  | 04/25/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC, |  |  | US |  | 1391515 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/25/2024 |  |  |  |  | The user facility's biomedical department released that the automated impella controller (aic) has a broken port cover. the aic was removed from service. | [5 empty] | The investigation for the console (aic) damage has been completed. data logs revealed no issues. the console was returned for evaluation. upon visual inspection, the pump receptacle cover had broken off completely. the cause of the issue was a defective component. f6/f8 should have been left blank in the initial report. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
27212 | 1220648-2024-10655 | 19226198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19226198&pc=OZD | 04/25/2024 | 06/20/2024 |  | 05/01/2024 |  |  | 04/25/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1813537 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported during set up of the new automated impella controller (aic), the battery activation buttons are not visible from underneath the console and were unable to activate the battery. there was no report of patient involvement. | [5 empty] | The investigation for the reported console (aic) red alarm has been completed. the console was returned for evaluation. during functional testing, the battery switch not being visible was confirmed. after opening up the console it was discovered that the battery switch was still inside the console but was not seated in place. the data logs were reviewed and it was evident that the consoles battery had been engaged via the battery switch at some point during its use. there were instances of the console remaining on while disconnected from ac power without shutting down which indicates that the battery was engaged. the reported issue was determined to be use related. it is suspected that the battery switch on this console was most likely tampered with. since there is evidence in the imc logs of the console remaining on while disconnected from ac power, this infers that the console's battery switch was engaged at some point during use. this indicates that the console's battery switch was accessible at some point. added- d9 device return date. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27213 | 1220648-2024-10656 | 19226199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19226199&pc=OZD | 02/28/2024 | 05/21/2024 |  | 05/01/2024 |  |  | 04/04/2024 | 05/01/2024 |  |  | 10/31/2025 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024413168 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | We received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that hemolysis occurred with a definite relationship to the impella 5.5 used on this patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported access site bleed and hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and hemolysis could not be determined.
27214 | 1220648-2024-10653 | 19226202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19226202&pc=OZD | 04/25/2024 | 06/24/2024 |  | 05/01/2024 |  |  | 04/25/2024 | 05/01/2024 |  |  |  |  | 05/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1503435 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 24 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. during patient support, the automated impella controller (aic) screen suddenly went black. the patient's blood pressure dropped at the same time. the aic may have stopped driving. however, there were no problems with hemodynamics as extracorporeal membrane oxygenation machine was used in combination with the aic. the aic restarted on its own and worked without further issues. there was no report of patient harm. | [5 empty] | The investigation into the aic - power on issues/blank screen has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. as described in the complaint, the pump was inserted into the patient. the pump was run continuously until there was a string of "scumain:reopen ncc_signal" log entries that culminated in a heartbeat check showing the calculator process had crashed and a partial reset by the software watchdog. the motor was terminated and the user asked for the pump to be resumed so therapy was lost for 20 seconds. the pump was restarted but removed from the console 40 minutes later. shortly before the calculator crash, the memory system had filled up and needed to delete old files to make space for new files to be created. the failure mode was not reproduced through running a pump and putting the console into the burn-in room. the root cause of the unexpected reboot was due to a calculator process crash. the cause of the process crash was unable to be determined due to the inability to reproduce. d.6a date was reported in error on the previously submitted report. there is no implant date for an automated impella controller. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. | The automated impella controller (aic) has been received from the customer. however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
27215 | 1220648-2024-10657 | 19227117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19227117&pc=OZD | 03/30/2022 | 05/01/2024 |  | 05/01/2024 |  |  | 03/30/2022 | 05/01/2024 |  |  | 01/31/2024 |  | 05/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183808 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 03/30/2022 |  |  |  |  | The user facility reported an 84-year-old female patient was implanted with an impella cp for mechanical circulatory support as high risk percutaneous coronary intervention. while on support, the patient experienced limb ischemia, which was resolved by explanting the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
27216 | 1220648-2024-10535 | 19211986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19211986&pc=OZD | 12/14/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 12/16/2022 | 04/30/2024 |  |  | 08/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022124724 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support during a ventricular tachycardia ablation. while on impella cp support, the patient experienced significant oozing at the access site with the repositioning sheath in place. the medical team changed the dressing and partially reversed heparin. patient's condition was stable after the procedure and the impella cp was successfully explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27217 | 1220648-2024-10536 | 19212377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212377&pc=OZD | 01/06/2024 | 06/05/2024 |  | 04/30/2024 |  |  | 04/01/2024 | 04/30/2024 |  |  | 11/30/2025 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024425887 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. we received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that left lower extremity ischemia occurred with a possible relationship to the impella cp used on this patient. | The patient's outcome was reviewed by abiomed's medical safety officer. it was determined that there was not enough information to associate use of device with outcome. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use referenced in the initial is for the impella cp with smart assist system in sections: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ added- b5 mso review d4-updated model number
27218 | 1220648-2024-10537 | 19212380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212380&pc=OZD | 04/11/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 04/11/2022 | 04/30/2024 |  |  | 12/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176807 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. once the impella was inserted, a retrograde dissection was noted, and contralateral access was required for percutaneous coronary intervention (pci). access for pci obtained through the left femoral artery with a 7x45 destination sheath, which was placed into the mid-right coronary artery, mid-left coronary artery, and proximal left anterior descending artery. thrombolysis in myocardial infarction was performed and flow was achieved in all vessels. physicians then placed an 8x60 everflex stent to seal the dissection. the dissection was successfully managed, and hemostasis was achieved with perclose x2. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27219 | 1220648-2024-10538 | 19212452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212452&pc=OZD | 01/12/2024 | 09/23/2024 |  | 04/30/2024 |  |  | 04/01/2024 | 04/30/2024 |  |  | 10/31/2025 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424204 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stenosis | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | We received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that left common femoral artery showed luminal stenosis occurred with a possible relationship to the impella cp used on this patient. | [5 empty] | The investigation for the reported luminal stenosis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. added- h6 component code 925 d4-corrected model number. f6/f8 should have been left blank in the initial report. g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿
27220 | 1220648-2024-10539 | 19212470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212470&pc=OZD | 03/24/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/24/2022 | 04/30/2024 |  |  | 10/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157108 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/24/2022 |  |  |  |  | The complainant reported a patient was being implanted with an impella cp device for mechanical circulatory support for a high-risk percutaneous cardiac insertion case. access was obtained in the right femoral artery without issue. however, the physician was unable to advance the pump through the abdominal aorta due to a second sheath from the left femoral artery which he had intended to use for the percutaneous coronary intervention. the impella cp kept getting caught on the sheath and tore a piece of the sheath off while attempting. ultimately, attempts were stopped, and the case was aborted. the patient will be brought back at a later date for impella protected pci. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. this report is being filed as part of a retrospective review of historical records.
27221 | 1220648-2024-10540 | 19212498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212498&pc=OZD | 04/09/2024 | 06/28/2024 |  | 04/30/2024 |  |  | 04/09/2024 | 04/30/2024 |  |  | 10/31/2025 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024421117 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | Difficult to Insert | Patient-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp implant procedure for 57-year-old male patient had to be aborted as the device was unable to be delivered due to patient anatomy. the vessel conditions were noted to be tortuous and there was no serial pre-dilation done. there was no reported patient harm. the patient is stable. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition since the pump could not be implanted due to the patient anatomy. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ e.1 address line 1 revised from the initial submission in accordance with updated procedures. g.1 revised manufacturing site fax number from the initial submission in accordance with updated procedures. h.6 medical device problem code was revised in accordance with updated procedures. a new code has been added to medical device problem codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
27222 | 1220648-2024-10541 | 19212512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212512&pc=OZD | 04/10/2024 | 04/30/2024 |  | 04/30/2024 |  |  | 04/10/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1200547 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) does not start up and displays 'system check failed'. the console was therefore replaced. there was no reported patient harm. | [5 empty] | (h3) the investigation into the reported low pump flow has been completed since the original report was filed. product was returned for investigation. the failure mode was reproduced in the lab. firmware check command performed on a test pc showed that the sau_powermod_1 firmware lacked information. the internal pbm circuitry was visually inspected and evidence of corroded copper traces and via of the pbm was observed. the corroded copper vias were in the vicinity of the super capacitors c69 and c70. the data analysis, imc logs confirm the issuance of the system self check failure alarm, following the console was rebooted with numerous powermod_1 communication errors between the pbm and epc. the root cause of the system self check failure was communication errors between the pbm and the epc caused by corrosion of the pbm . the corrosion of the pbm was due to leaking of electrolyte from the super capacitors c69 and c70.
27223 | 1220648-2024-10542 | 19212548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212548&pc=OZD | 04/12/2024 | 05/21/2024 |  | 04/30/2024 |  |  | 04/12/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC |  |  | ABIOMED, INC. |  |  | US |  | 1204668 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during customer training from local team, the automated impella controller (aic) did not detect the purge disk. the purge system was changed but the issue did not resolve. therefore, a new aic was used and the problem resolved. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned. the logs were reviewed for the reported issue that the purge disc could not be detected was confirmed. the purge disc flag was found to be broken. the root cause of the issue was a broken purge disc flag. a.3 revised as from the initial submission in accordance with updated procedures. e.1 name prefix/title was inadvertently omitted from the previously submitted report and address line 1 was revised as it was previously entered incorrectly. h.6 code 2199 was reported incorrectly on the previously submitted report.
27224 | 1220648-2024-10545 | 19212706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212706&pc=OZD | 04/14/2024 | 06/03/2024 |  | 04/30/2024 |  |  | 04/14/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025449632 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2024 |  |  |  |  | The user facility reported a 29-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complaint had placed ecmo (extracorporeal membrane oxygenation) and then the cp for a 1-day postpartum patient with cgs (cardiogenic shock) and ef (ejection fraction) of 10%. during the support there was bleeding from access site treated by pressure and blood product delivery. the blood was also noted for the low cvp (central venous pressure). additionally, heparin was adjusted. the team also had placed lines for limb perfusion and concerns of limb ischemia. after 11 days the pump was explanted with plans for heart transplant. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ . ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿. | The investigation into the access site bleeding ¿ major and the limb ischemia ¿ reversible issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved by adding additional sutures. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
27225 | 1220648-2024-10544 | 19212709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212709&pc=OZD | 04/13/2024 | 07/03/2024 |  | 04/30/2024 |  |  | 04/13/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2025466165 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support for patient, there were purge pressure high alarms. the purge line was noted to be free of kinks. the purge cassette was not replaced. tpa was added to the purge. the motor current was monitored, and pump exchange was recommended. there was no reported patient harm. | Additional information: it was reported that patient care was withheld and the patient expired. per abiomed medical safety office, abiomed does not have enough information to associate use of device with outcome. |  |  |  |  | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the high purge pressure issue has been completed since the original report was filed. the pump was returned for investigation. the product was inspected, and biomaterial was found around impeller which upon clearing, did not reproduce high purge pressure. the root cause of high purge pressure was most likely biomaterial as high purge pressure did not reproduce. instructions for use for the related event: high pressure purge. impella 5.5 with smart assist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ b5 additional information added. b7 additional information added d9 updated. g1 manufacturing reporting contact and manufacturing site address corrected. d6b added. h6 medical device problem code corrected. h10 instructions for use corrected.
27226 | 1220648-2024-10543 | 19212845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212845&pc=OZD | 04/12/2024 | 06/13/2024 |  | 04/30/2024 |  |  | 04/12/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025437799 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Positioning Failure | Difficult to Insert | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was unable to be delivered due to calcifications and kinks. the device was unable to pass through the vasculature. kinks and calcification were noted when the pump was pushed forward. the procedure was aborted, however, an attempt will be made to implant a different impella using the opposite femoral artery. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related, the impella cp was unable to be delivered due to calcifications and kinks as per the clinical description.
27227 | 1220648-2024-10548 | 19212868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212868&pc=OZD | 03/31/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 03/31/2022 | 04/30/2024 |  |  | 10/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022153872 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 48 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/31/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, loss of bilateral pulses in the legs, clots and a tear on the distal end of the anastomoses were noted. the patient was taken to the operating room where clots were surgically removed to resolve the limb ischemia. the tear that occurred on the distal end of the anastomoses was likely due to the 45 degree bevel of the graft and was sutured to achieve hemostasis and the patient was transfused with two units of packed red blood cells and one unit of platelets. patient condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿. caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27228 | 1220648-2024-10546 | 19212882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212882&pc=OZD | 04/14/2024 | 07/15/2024 |  | 04/30/2024 |  |  | 04/14/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1202894 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Computer Operating System Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/14/2024 |  |  |  |  | The user facility reported a 55-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a brief controller error alarm during impella 5.5 support. the error coincided with a loss in placement and lv (left ventricle) signals on the aic (automated impella controller). the issue resolved on its own after a couple minutes, but a backup console was placed outside of the patient's room in case the issue recurred. there was no patient harm. | [5 empty] | The impella device was received from the customer on (b)(6)2024 and therefore, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the automated impella controller (aic) error and placement alarms has been completed. data logs were reviewed and confirm the alarm. the device was received and the analysis revealed the error and loss of placement signal was due to an optical bench firmware communication protocol deficiency resulting in misalignment of data communication packets received by the embedded personal computer. the console software operated as designed. f6 and f8 should have been blank.
27229 | 1220648-2024-10549 | 19212908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212908&pc=OZD | 04/12/2024 | 07/03/2024 |  | 04/30/2024 |  |  | 04/15/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025450512 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for 41-year-old female patient, the pump was repositioned due to 'position in ventricle' alarm. the pump stopped after a few minutes. due to patient condition worsening, the decision was made to explant the pump. following explant, it was noted that the corda tendinea was sucked into the pump. | [5 empty] | The investigation for the pump stop and the positioning issue has been completed. pump was returned for evaluation. upon visual inspection, no mechanical or purge related abnormalities were noted. a catheter kink was observed near the 35cm mark of catheter but did not affect pump functionality. pump was run in a bucket of water at p-9 for 10 minutes while purging and a pump stop was not reproduced. no data logs were returned for evaluation. clinical details noted that after repositioning, a "pump in ventricle" alarm was triggered and a few minutes later a pump stop occurred. pump was explanted and it was noted that a corda tendinae had been sucked into the pump. the root cause of the pump stop was biomaterial due to improper positioning. the instructions for use referenced in the initial report is for the impella cp with smart assist for use during cardiogenic shock and high risk cpi in the following sections: section: using the automated impella controller with the impella catheter. section: using the automated impella controller with the impella rp with smartassist system catheter. section: impella stopped. added- d9 device return date, added- d6b. corrected g1 manufacturing site address line 1,2, manufacturing site city, zip code, fax number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
27230 | 1220648-2024-10547 | 19212920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212920&pc=OZD | 03/25/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/25/2022 | 04/30/2024 |  |  | 01/31/2024 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182829 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/25/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, there was access site bleeding, noted by the very oozy groin. this was treated by applying pressure and giving the patient protamine. the decision was then made to explant the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records.
27231 | 1220648-2024-10550 | 19212968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212968&pc=OZD | 04/08/2024 | 05/14/2024 |  | 04/30/2024 |  |  | 04/16/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1754606 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that battery smoke appeared and the automated impella controller (aic) displayed failure alarm. there was no reported patient harm. | [5 empty] | The investigation for the reported console (aic) battery failure has been completed. the console was returned for evaluation. upon visual inspection, there were clear burn marks around c199 on the pbm. the console was able to turn on and boot up without the failure mode recurring. a new pbm was inserted and the failure mode once again did not happen. data logs were reviewed finding the console immediately had event sets 360 and 23 for battery failure and battery level low respectively confirming the "failure alarm" from the complaint. these alarms resolved after three seconds. the ac power registered as disconnected seven seconds later so there was no ac power at the time of failure. the cause of the aic issue was a defective pbm board but the cause of the burnt pbm is unknown. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
27232 | 1220648-2024-10551 | 19212982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19212982&pc=OZD | 03/04/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 03/04/2022 | 04/30/2024 |  |  | 11/30/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169215 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella support, there were no pulses found in the patient's right foot and hemolysis was suspected due to the patient¿s urine changing from clear to brown within one hour of the impella cp implant. the patient was escalated to impella 5.5. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27233 | 1220648-2024-10552 | 19213124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213124&pc=OZD | 04/16/2024 | 06/20/2024 |  | 04/30/2024 |  |  | 04/16/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2025459092 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported impella 5.5 support had high purge pressure 950mmhg and low pump flow 1,5ml. the purge line was checked, and the cannula was checked via echo, and there was no biomaterial visible. the motor current was monitored but the pump was ultimately explanted. there was no reported patient harm and the patient is noted to be stable. | [5 empty] | The investigation for the reported high purge pressure has been completed. no product was returned for evaluation. clinical details noted that no kinks were observed on purge line, heparin was used in purge, and no biomaterial was visible on echo. data logs were reviewed and left data logs from case show a gradual rise in purge pressure over a period of 7 hours before high purge pressure alarms were triggered. no motor current spikes were observed before purge pressure increase. the root cause of the high purge pressure is most likely due to a gradual buildup of biomaterial. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist in section: using the automated impella controller with the impella catheter. g1-corrected mfr site name and address. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
27234 | 1220648-2024-10554 | 19213148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213148&pc=OZD | 03/25/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/25/2022 | 04/30/2024 |  |  | 01/31/2024 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183973 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/25/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the physician attempted to reposition the pump. an echocardiogram was performed to view the placement and the physician decided to explant and replace the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27235 | 1220648-2024-10555 | 19213310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213310&pc=OZD | 03/20/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 03/20/2022 | 04/30/2024 |  |  | 12/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181092 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/20/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, no distal pulses were felt in the left lower extremity. the patient underwent a fem fem bypass. it is unknown if the intervention resolved the ischemia. additionally, it was reported that there was vascular repair done to groin due to wrong size sheath being used. no further information was provided. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27236 | 1220648-2024-10556 | 19213361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213361&pc=OZD | 03/26/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/26/2022 | 04/30/2024 |  |  | 12/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170422 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/26/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there was access site bleeding, noted by the hematoma at the site. this was treated through manual pressure, increasing the repositioning sheath angle, and discontinuing heparin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27237 | 1220648-2024-10553 | 19213435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213435&pc=OZD | 03/15/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 03/15/2022 | 04/30/2024 |  |  | 10/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154119 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was noted a positioning issue as the pigtail appeared to be stretched out and connected to a papillary muscle. the decision was made to not reposition the pump as there were good flows and a risk of the pump coming out across the aortic valve. additionally, the pump's p-level was dropped down to p6. patient was successfully weaned from impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27238 | 1220648-2024-10557 | 19213450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213450&pc=OZD | 03/23/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/23/2022 | 04/30/2024 |  |  | 06/30/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022107837 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/23/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia. a femorofemoral bypass was performed to restore the flow down the patient¿s leg. additionally, the patient experienced hemolysis. the position of the impella was checked to determine the cause of hemolysis. it is unknown of the hemolysis was resolved. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
27239 | 1220648-2024-10558 | 19213628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213628&pc=OZD | 04/16/2024 | 06/11/2024 |  | 04/30/2024 |  |  | 04/16/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451356 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/16/2024 |  |  |  |  | The complainant reported that during the procedure for an impella cp placement for an 80-year-old male patient, insertion was not possible. there was too much kinking due to patient¿s anatomy and calcification and the impella cp cannula kinked. the staff will do a ct angio and retry with a vascular surgeon. the patient is noted to be stable. | [5 empty] | The impella device has not been received. at the conclusion of the investigation, a supplemental mdr will be filed. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. the instructions for use referenced in the initial report is for the impella cp with smartassist system in the following sections: section: warnings & cautions: warnings section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves d4-updated model number.
27240 | 1220648-2024-10559 | 19213630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213630&pc=OZD | 03/26/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/26/2022 | 04/30/2024 |  |  | 12/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173067 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/26/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported that there was a high-pressure alarm, the lines were checked for kinks, and none were found. abiomed support advised the team to change the purge solution from d5 with heparin to tissue plasminogen activator protocol (tpa), and the abiomed support representative went on site to assist the medical team with the tpa protocol. later in the day the purge flow drifted lower again, the purge solution was changed to tpa protocol due to flow rates dropping a third time. the patient remained on impella support for three additional days and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27241 | 1220648-2024-10560 | 19213637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213637&pc=PYX | 04/18/2024 | 06/26/2024 |  | 04/30/2024 |  |  | 04/18/2024 | 04/30/2024 |  |  | 11/30/2025 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024433044 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate use of the device with the outcome. | The user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella rp had low flow after insertion, and it was replaced. a new rp was implanted successfully. there was no patient harm. |  |  |  |  | The investigation for the low pump flow has been completed. the impella device was not returned for investigation. data logs were reviewed and did not find any placement signal trends. during support, the flow dropped to below 3.5 l/min at p-9. no indication for biomaterial ingestion observed. the cause of the low pump flow cannot be determined since the complaint device not returned for investigation and no sufficient clinical data provided. the instructions for use referenced in the initial report is for the impella rp with smart assist system in the following sections: section: using the automated impella controller with the impella rp with smart assist system catheter section: direct aortic insertion added- b5 mso review d4-corrected model number f6/f8-should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
27242 | 1220648-2024-10561 | 19213640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213640&pc=OZD | 03/23/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/23/2022 | 04/30/2024 |  |  | 08/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2022139163 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Deliver | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/23/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding and a groin hematoma. the physician also reported the first vascular approach failed as the surgeon was unable to pass through the vasculature, the second approach was successful. the cardiologist lowered then stopped the heparin in the purge solution to address the bleeding. in addition, the patient received units of packed red blood cells. the device was explanted and the vascular surgeon repaired the patient's femoral artery. patient outcome was reportedly improved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27243 | 1220648-2024-10562 | 19213695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213695&pc=OZD | 04/18/2024 | 07/15/2024 |  | 04/30/2024 |  |  | 04/18/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1249363 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) alarms were not audible. the aic was replaced. there was no patient harm. |  |  |  |  | The investigation for the console (aic) inaudible alarms has been completed. the data logs did not contain any information relevant to this failure mode. the aic was returned for evaluation and upon functional testing, the failure was reproducible upon boot up in the lab, which confirmed that alarms were not audible. inspection of internal console circuitry revealed a disconnected speaker ribbon cable. therefore, the root cause for the inaudible alarms was a disconnected speaker ribbon cable. added- b5 mso review, d9 device return date d4-corrected model number f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27244 | 1220648-2024-10563 | 19213819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213819&pc=OZD | 04/18/2024 | 05/30/2024 |  | 04/30/2024 |  |  | 04/18/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438126 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support post cardiac surgery. during support, the patient developed a hematoma at the access site. the patient was given blood products, heparin was withheld, and the access site was stitched. the patient was successfully weaned from the impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since adding additional sutures resolved the bleed.
27245 | 1220648-2024-10564 | 19213856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213856&pc=OZD | 03/30/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/30/2022 | 04/30/2024 |  |  | 10/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154562 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/30/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during treatment the patient may have developed thrombocytopenia. the patient had low platelet level and there was no heparin in the purge nor systemic. they were given continuous renal replacement therapy (crrt). additionally, the patient was bleeding from an unknown source and was administered bival systemically. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). this report is being filed as part of a retrospective review of historical records.
27246 | 1220648-2024-10565 | 19213946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213946&pc=OZD | 04/08/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 04/08/2022 | 04/30/2024 |  |  | 09/30/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2022136669 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/08/2022 |  |  |  |  | Complainant reported a patient was implanted with an impella cp for mechanical circulatory support due to heart disfunction. while on support, the patient had a hematoma and oozing at the access site. a new stitch was placed, and manual pressure held followed by femstop to hold pressure. additionally, the patient experienced hemolysis. heparin was not started. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. this report is being filed as part of a retrospective review of historical records.
27247 | 1220648-2024-10566 | 19213978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213978&pc=OZD | 04/05/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 04/05/2022 | 04/30/2024 |  |  | 10/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149082 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, the patient developed a thrombus. clot was removed from bilateral common femoral artery (cfa) and superficial femoral artery (sfa) using fogarty catheter. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
27248 | 1220648-2024-10567 | 19213982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19213982&pc=OZD | 05/09/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 05/09/2022 | 04/30/2024 |  |  | 10/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149692 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/09/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella cp access site after the manta closure was deployed. manual compression was held at the site and a balloon tamponade was attempted without success. a covered stent was deployed on the central femoral artery and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27249 | 1220648-2024-10568 | 19214016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214016&pc=OZD | 04/19/2024 | 06/06/2024 |  | 04/30/2024 |  |  | 04/19/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025456777 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported a 54-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while obtaining access to the right femoral artery for impella cp implant, a dissection of the external iliac artery occurred. a stent was placed to manage the dissection and access for implant was obtained in the left femoral artery. the subsequent delivery was successful. the patient was reported as stable. | [5 empty] | The investigation into the vascular damage - peripheral vessel issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: warnings & cautions: to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
27250 | 1220648-2024-10569 | 19214052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214052&pc=OZD | 04/06/2022 | 03/07/2025 |  | 04/30/2024 |  |  | 04/06/2022 | 04/30/2024 |  |  | 01/31/2024 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187016 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2022 |  |  |  |  | The user facility reported a 76-year-old female on impella cp for mechanical circulatory support. while on support, it was reported the patient experienced a small amount of oozing. the impella cp was weaned and removed in cardiac cath lab. perclose x2 deployed and a femstop applied to right groin. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
27251 | 1220648-2024-10570 | 19214056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214056&pc=OZD | 03/28/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/28/2022 | 04/30/2024 |  |  | 12/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173748 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/28/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during treatment the patient experienced limb ischemia and had difficulty inserting the introducer during insertion of the pump through axillary access. upon insertion of the sheath, it was noted that it was occlusive. the decision was made to create an external left common femoral artery (cfa) right radial bypass. the peel-away sheath was removed, and the repositioning sheath was inserted, which improved distal flows. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
27252 | 1220648-2024-10571 | 19214081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214081&pc=PYX | 04/20/2024 | 06/27/2024 |  | 04/30/2024 |  |  | 04/23/2024 | 04/30/2024 |  |  | 11/30/2025 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024431726 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Hemolysis | Tricuspid Valve Insufficiency/ Regurgitation | Renal Impairment | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella rp flex for mechanical circulatory support. it was initially stated that the pump was unable to reach expected flow rates despite being in an optimal position. it was then documented that the patient's urine became red with increasing lactate dehydrogenase levels. upon confirmation of hemolysis, attempts to reposition the pump were made and two units of blood were given. the patient was subsequently taken to the operating room where a dialysis line was placed. despite the interventions performed, the decision was ultimately made to explant the pump to prevent further hemolysis/kidney failure. while assessing the ability to wean the patient, the surgeon reported that the pump cannula was also causing moderate to severe tricuspid regurgitation. the pump was eventually explanted with the patient on medical therapy. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. rp flex with smart assist system with automated impella controller section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿ section: warnings & cautions: cautions ¿patients with tricuspid or pulmonary valve stenosis or insufficiency, and patients with prosthetic tricuspid or pulmonary valves, may be compromised by the use of the impella rp flex with smartassist system catheter.¿ | The investigation into the reported issues has been completed since the original report was filed. the pump was returned for investigation. no physical abnormalities were reported on the returned product. biomaterial was found wrapped around the impeller. blue fibers were found during visual inspection. no cannula kink was observed. the pump passed hemolysis testing. the data logs showed the pump to be flowing on the lower side of the metrics for the entire duration of the case. there were no motor current spikes outside of the p-level changes and no ingestion events. the fibers/biomaterial was most likely from after explant when wiped with a cloth. there are suction alarms consistent throughout support. the cause of the heart valve damage issue was not determined since sufficient clinical information was not provided. the cause of hemolysis, low pump flow and renal failure was improper positioning of the device. d.1 brand name and d.4 model number were revised from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 medical device problem code was revised from the initial submission in accordance with updated procedures.
27253 | 1220648-2024-10572 | 19214236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214236&pc=OZD | 04/23/2024 | 05/21/2024 |  | 04/30/2024 |  |  | 04/23/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1302710 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | The complainant reported that when trying to change the purge lines for their automated impella controller (aic), the console would not recognize the purge disc. the medical staff attempted three different purge cassettes without success, but on another aic, the cassettes worked fine. the aic was exchanged and the issue was resolved. no patient harm has been reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the logs were reviewed and the issues with the aic not being able to detect the purge disc was confirmed. the purge disc was found to be broken. the cause of the issue was a defective component. a.3 revised in accordance with updated procedures.
27254 | 1220648-2024-10573 | 19214365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214365&pc=OZD | 04/06/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 04/06/2022 | 04/30/2024 |  |  | 12/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174258 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, motor current was over 1000 and the pump alarmed saying it stopped but then resolved. additional impella device was needed. patient remained on impella support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27255 | 1220648-2024-10574 | 19214380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214380&pc=OZD | 03/23/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/23/2022 | 04/30/2024 |  |  | 11/30/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168953 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 03/23/2022 |  |  |  |  | The user facility reported a 63-year-old patient needing mechanical circulatory support. it was reported that while on impella support, there were complications due to ischemia. a distal leg perfusion catheter was placed to manage the issue. no further information was available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27256 | 1220648-2024-10575 | 19214412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214412&pc=OZD | 03/24/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/24/2022 | 04/30/2024 |  |  | 12/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177284 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 03/24/2022 |  |  |  |  | The user facility reported a 64-year-old patient was placed on impella cp support for high-risk percutaneous cardiac insertion (hrpci). it was reported that there was dissection of the right femoral artery after the sheath was inserted. the artery was heavily calcified. initially a long 14x25 cm sheath was chosen with serial dilation and a 14 fr dilator. the team attempted to insert the long sheath, but it would not pass completely. a shorter 14x13 cm sheath was chosen and passed, but angio showed possible dissection in the area. the wire was passed. multiple images showed the dissection was stable, the procedure was continued, and stents were placed. the impella was placed, procedure completed, and the pump was weaned and explanted. after the case, an angio showed more prominent dissection. the artery was stented with good results. the site was closed appropriately with surgical intervention. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records.
27257 | 1220648-2024-10576 | 19214439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214439&pc=OZD | 04/23/2024 | 06/24/2024 |  | 04/30/2024 |  |  | 04/23/2024 | 04/30/2024 |  |  | 06/30/2025 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024375554 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | The user facility reported a 54-year-old male with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the impella cp pump came back across the aortic valve during support. the medical staff attempted to reposition the pump unsuccessfully and decided to explant the impella cp. no patient harm has been reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the positioning issue most likely was patient movement. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures.
27258 | 1220648-2024-10577 | 19214453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214453&pc=OZD | 04/07/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 04/07/2022 | 04/30/2024 |  |  | 11/30/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168174 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient's right lower extremity was dopplered, and no distal pulse was found. additionally, after the impella cp was placed and penumbra aspiration failed to aspirate a clot. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. thrombus. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
27259 | 1220648-2024-10578 | 19214515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214515&pc=OZD | 03/24/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/24/2022 | 04/30/2024 |  |  | 01/31/2024 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181730 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 03/24/2022 |  |  |  |  | The user facility reported a 78-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma at the insertion site and was brought back to the lab and received one covered stent to the axillary artery. patient was reportedly stable following intervention. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27260 | 1220648-2024-10580 | 19214542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214542&pc=OZD | 04/19/2024 | 06/04/2024 |  | 04/30/2024 |  |  | 04/19/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025458048 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported a 53-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella cp support lost pulse in the impella leg. the physician was notified, and the decision was made to pull the impella overnight. in the morning the patient was taken for a thrombectomy. perclose created a narrowing and due to low flow, clot formed. patient hemodynamically stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details. impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d4 and h10 have been revised from the initial submission.
27261 | 1220648-2024-10579 | 19214543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214543&pc=OZD | 03/28/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 03/28/2022 | 04/30/2024 |  |  | 12/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171256 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was shocked three (3) times for ventricular fibrillation after bleeding episode. the patient developed a hematoma which improved after manual pressure held and high mean arterial pressure decreased to within normal range. patient received two (2) units of packed red blood cells post ecmo cannulation due to blood loss from impella cp site, ecmo cannulation sites and left groin hematoma. due to blood loss during transport patient received two (2) more units of prbc. patient remained on support and was sent to intensive care unit. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined
27262 | 1220648-2024-10581 | 19214578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214578&pc=OZD | 04/08/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 04/08/2022 | 04/30/2024 |  |  | 09/30/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022145501 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was reported that the pump stopped and was restarted without issue. the pump was removed. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27263 | 1220648-2024-10583 | 19214586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214586&pc=OZD | 04/19/2024 | 06/11/2024 |  | 04/30/2024 |  |  | 04/19/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025485222 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Insufficient Flow or Under Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported a 65-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed hemolysis during impella cp support coinciding with persistent suction alarms. repositioning was attempted, but the suction alarms continued. the patient was given three units of red blood cells, at which point the suction resolved. it was then stated that no pulse was found in the right upper extremity distal to the insertion site. due to the noted conditions, the decision was made to explant the pump. the patient was reported to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure | The investigation for the reported hemolysis, limb ischemia, and low pump flow has been completed. data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in aorta. logs also showed low flow occurred for 17 hours since the pump started that triggered auto suction response and suction alarms. product was not returned for review. based on clinical description & data analysis, the root cause of hemolysis was most likely due to patient condition as the pump could run at higher p-level after adding the volume. based on clinical description and data analysis, the root cause of low flow was most likely due to patient condition as the pump could run at higher p-level after adding the volume. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the instructions for use referenced in the initial report is for the impella cp with smartassist system in section: use of echocardiography for positioning of the impella catheter, section: potential adverse events (united states), section: hemolysis, and section: suction. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27264 | 1220648-2024-10584 | 19214653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214653&pc=OZD | 04/14/2024 | 06/13/2024 |  | 04/30/2024 |  |  | 04/20/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1791478 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure of Device to Self-Test | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/20/2024 |  |  |  |  | The abiomed clinical representative brought an automated impella controller (aic) to the user facility. upon startup, the aic displayed a "self-check error". the aic was replaced. there was no patient involvement. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console (aic) self check error has been completed. the console was returned for evaluation. during testing, the reported failure mode was reproduced upon turning on the console. the cable between pud and the 4-in-1 was not fully seated at the 4-in-1 side. data logs were reviewed findind there was a communication issue with the pud during boot-up, so the console was rebooted automatically. after reboot, since the communication issue with the pud was persistent, the boot-up process resulted in a ¿system self check failed¿ screen per the log ¿post: syscall10 checkfirmware failed¿. this confirms the reported failure mode. the root cause for the system self check failure was a loose cable connection at the 4-in-1 connector. d4-updated model number
27265 | 1220648-2024-10585 | 19214724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214724&pc=OZD | 04/21/2024 | 06/07/2024 |  | 04/30/2024 |  |  | 04/21/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025450563 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardia infarction/cardiogenic shock was on impella cp support due to the patient being in ventricular fibrillation. while on support, the pump was slightly deep, and pigtail was stuck on the papillary muscle causing frequent suction alarms. the physician attempted to reposition the impella pump under echo; however, the physician was unable to safely free the impella pump from the papillary muscle without pulling back into the aorta. the physician, therefore, decided to leave the impella pump in the position and ran the pump at p-levels p-5/p-6. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow and the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the positioning issue was most likely use related issue due to improper technique as impella slightly deep and pigtail stuck on a papillary and was unable to safely free from papillary without pulling back into the aorta and let to low pump flow. the cause of the low pump flow was most likely improper positioning due to improper technique.
27266 | 1220648-2024-10586 | 19214751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214751&pc=OZD | 03/25/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/25/2022 | 04/30/2024 |  |  | 12/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171450 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/25/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella was placed without using ultrasound for visualization access. while on support, the echocardiogram showed the impella was sitting on the mitral valve apparatus. the patient was taken off all pressors and the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27267 | 1220648-2024-10587 | 19214771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214771&pc=OZD | 04/22/2024 | 06/24/2024 |  | 04/30/2024 |  |  | 04/22/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1360190 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Tipped Over | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/22/2024 |  |  |  |  | The user facility reported the automated impella controller's (aic) lower leg/wheel assembly became disconnected from console stand. the aic tipped while staff member was rolling it from icu with no patient/pump connected at the time. the lower left-hand corner of aic housing chipped, and a bolt was missing from stand assembly. no patient harm or involvement was reported. the console was pulled from service. | [5 empty] | The investigation has been completed. the data logs were reviewed but did not show the cart breaking or the console being damaged. the console was returned and found damaged in the lower left front of the console. this did not affect functionality and the console performed as expected. it was not determined why the wheel of the cart fell off. the cause of the automated impella controller (aic) damage was due to the detachment of the aic cart's leg. a.1, a.3 and d.2 revised as previously entered incorrectly. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27268 | 1220648-2024-10588 | 19214874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214874&pc=OZD | 03/28/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 03/28/2022 | 04/30/2024 |  |  | 12/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180998 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had labs that were hemolyzed. discussed potential hemolysis and reducing p-level, volume, and rechecking echo. patient was successfully weaned from support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) section: use of echocardiography for positioning of the impella catheter section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ section: use of echocardiography for positioning of the impella catheter impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: section: hemolysis hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
27269 | 1220648-2024-10590 | 19214875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214875&pc=OZD | 04/22/2024 | 06/11/2024 |  | 04/30/2024 |  |  | 04/22/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025485232 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/22/2024 |  |  |  |  | The user facility reported a (b)(6) year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the cp had poor positioning and was removed and replaced. the cp explant caused no harm or injury, and support continues on second cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ | The investigation for the reported positioning issue has been completed. data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle/aorta and suction alarms. no other issues were found on the logs. product was not returned for review. based on clinical description and data analysis, the root cause of positioning issue was most likely use related during cpr.
27270 | 1220648-2024-10589 | 19214884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19214884&pc=OZD | 09/16/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 09/16/2022 | 04/30/2024 |  |  | 05/31/2024 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228960 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/16/2022 |  |  |  |  | The complainant reported a patient in non-st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced oozing at the impella access site. multiple dressing changes were performed, heparin was withheld and the patient received multiple blood products. the patient remained on impella support until they were successfully weaned and the impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27271 | 1220648-2024-10591 | 19215029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19215029&pc=OZD | 03/07/2024 | 06/28/2024 |  | 04/30/2024 |  |  | 04/23/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438176 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | The user facility reported a 62-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that there was access site bleeding. the medical staff performed several dressing changes and withheld heparin. additionally, there was a purge pressure low alarm and leak was observed into the purge side arm retainer. the medical staff tried a purge bypass unsuccessfully. the impella 5.5 was replaced, and the patient survived the event. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the reported purge leak and access site bleed has been completed. the pump and purge cassette were returned for investigation. the purge sidearm of pump was not returned. the returned pump was cut open. a sidearm bypass was seen on the returned product, and upon checking, a loose connection was found on the bypassed sidearm. no further investigation could be performed since the sidearm was not returned. as per clinical details, a purge bypass was performed after a noted purge leak from the sidearm. the medical doctor explanted the pump after the purge bypass did not work. the cause of the purge leak issue was most likely a leak from the sidearm, but the exact location of the leak was unknown. the cause of the access site bleeding issue was most likely patient condition related. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures.
27272 | 1220648-2024-10592 | 19215129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19215129&pc=OZD | 09/15/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 09/16/2022 | 04/30/2024 |  |  | 06/30/2024 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237390 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Injury | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/16/2022 |  |  |  |  | The complainant reported a patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella had sporadic suction events, which were resolved as the p levels were lowered to wean the patient off of the impella. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27273 | 1220648-2024-10593 | 19215202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19215202&pc=PYX | 01/25/2019 | 06/12/2024 |  | 04/30/2024 |  |  | 04/23/2024 | 04/30/2024 |  |  | 03/31/2020 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1335603 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 31 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | In april of 2024 a literature review was conducted for the impella rp. the publication entitled "cardiogenic shock in thyroid story: a biventricular impella" was authored by evan caruso, department of cardiology, cooper university hospital. the publication reported for a 31-year-old male had hyperthyroidism and atrial fibrillation and was suffering from thyroid storm. the team placed impella cp and rp devices and the condition was improving. after three days, the patient had hemolysis and thrombocytopenia that improved after the pump was weaned and explanted. the impella rp was explanted first at day two and the impella cp was explanted at day four. at the six month follow-up, the patient's condition had improved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the thrombocytopenia and hemolysis has been completed. product and data logs were not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of thrombocytopenia was unable to be determined as limited clinical details were provided and no product was returned for review. impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27274 | 1220648-2024-10594 | 19215232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19215232&pc=OZD | 03/25/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 03/25/2022 | 04/30/2024 |  |  | 01/31/2024 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183111 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 03/25/2022 |  |  |  |  | The user facility reported a 67-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during support the patient experienced limb ischemia. additionally, it was reported that the patient had access site oozing. the angle of entry was reinforced and blood products were administered to the patient. the patient improved and the device was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. b1 corrected to serious injury for limb ischemia. d4 model number.
27275 | 1220648-2024-10595 | 19215267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19215267&pc=OZD | 03/29/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/29/2022 | 04/30/2024 |  |  | 08/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022134492 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/29/2022 |  |  |  |  | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. while on support the patient experienced bleeding from an unknown source. heparin was put on hold due to bleeding. additionally, pump flow was low and could not be resolved. the physician decided to remove and replace the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
27276 | 1220648-2024-10596 | 19215330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19215330&pc=OZD | 09/20/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 09/20/2022 | 04/30/2024 |  |  | 05/31/2024 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231086 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella support, oozing at the impella access site was noted. as a result, heparin was stopped. patient condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27277 | 1220648-2024-10597 | 19215626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19215626&pc=OZD | 04/23/2024 | 07/09/2024 |  | 04/30/2024 |  |  | 04/24/2024 | 04/30/2024 |  |  | 11/30/2025 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025437531 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 24 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | The user facility reported a 24-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient developed hemolysis during impella 5.5 support. it was noted that the hemolytic findings were accompanied with increased lactate dehydrogenase levels and lower flow rates. due to the noted complications, the decision was made to replace the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation into the hemolysis and the low pump flow issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of hemolysis and the low pump flow was unable to be determined as limited clinical details were provided and no product was returned for review. f6/f8-should have been left blank in the initial report.
27278 | 1220648-2024-10598 | 19215762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19215762&pc=OZD | 04/21/2024 | 06/28/2024 |  | 04/30/2024 |  |  | 04/24/2024 | 04/30/2024 |  |  | 10/31/2025 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024418715 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  | Black Or African American | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was cannulated for veno-arterial extracorporeal membrane oxygenation and brought to the operating room for impella 5.5 placement for support due to going into cardiogenic shock. while on support, the patient had bleeding at the swan site. anticoagulants were withheld, and the patient received blood transfusions. additionally, high purge pressure was seen on the impella pump. tissue plasminogen activator was added to the purge and the staff monitored. no further updates were noted. | Additional information: it was later reported the impella was removed due to a unique slow move towards withdrawing of care. the patient expired. per the abiomed medical safety office, abiomed does not enough information to associate use of the device with the outcome. |  |  |  |  | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation into the high purge pressure and the access site bleeding - major issues has been completed since the original report was filed. the pump was returned for investigation. the issue was not reproduced during upon testing. however, data logs showed a gradual rise in purge pressure during the case run. unsuccessful attempts were made by replacing the purge cassette and board after a reported high purge pressure alarm. purge pressure was gradually reduced likely after administering tpa. the cause of the high purge pressure issue was biomaterial. the cause of the access site bleeding issue was most likely patient condition related, and its likely relation to impella was unknown.
27279 | 1220648-2024-10599 | 19215805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19215805&pc=OZD | 04/07/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 04/08/2022 | 04/30/2024 |  |  | 01/31/2024 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184085 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the catheter slightly advanced in the left ventricle. the doctor pulled the catheter back across the aortic valve in an attempt to reposition the impella cp at bedside with no success. the impella cp was eventually explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined
27280 | 1220648-2024-10600 | 19215878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19215878&pc=OZD | 04/27/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 04/27/2022 | 04/30/2024 |  |  | 08/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114752 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Female |  |  | Asian | Injury | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/27/2022 |  |  |  |  | .the complainant reported that there were delivery issues and product damage for the cannula with the impella device on a 79 year old female patient. the delivery issues were encountered as several unsuccessful attempts were made to insert the pump as the patient was small and the aortic arch was extremely calcified. additionally, unsuccessful attempts were made when seeking alternatives such as using a buddy wire and constructing and inserting an artificial blood vessel. additionally, due to the amount of force applied, the cannula was damaged. as a result, a new pump was inserted instead. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
27281 | 1220648-2024-10601 | 19215906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19215906&pc=OZD | 04/24/2024 | 06/06/2024 |  | 04/30/2024 |  |  | 04/24/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1141485 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported that the automated impella controller (aic) had a battery failure, and consequently, the aic was replaced. there was no report of patient involvement. | [5 empty] | The investigation for the reported aic (console) battery cell failure has been completed. the console was returned for evaluation. upon testing, battery failure issues were able to be reproduced. issues were determined to be caused by internal battery cell imbalance (found in vcell3 of battery). the data logs were reviewed and the issues with the battery failures on this console were confirmed. there were multiple instances of battery failure alarm (transport connection blown/missing) (event set #360), battery failure alarm (low voltage) [v < 12v] (event set #350), and battery level low (50%) (event set #23) found from the reported occurrence date. the root cause of this issue was internal battery cell imbalance. | The automated impella controller (aic) has been received from the customer. however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
27282 | 1220648-2024-10602 | 19215966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19215966&pc=OZD | 04/08/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 04/08/2022 | 04/30/2024 |  |  | 11/30/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156694 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/08/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be implanted with an impella cp device for mechanical circulatory support. during the placement of the pump, the procedure was aborted as the pump was not able to be advanced beyond the iliac due to the patient's vascular anatomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27283 | 1220648-2024-10603 | 19216179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19216179&pc=OZD | 01/19/2024 | 05/20/2024 |  | 04/30/2024 |  |  | 04/04/2024 | 04/30/2024 |  |  | 09/30/2025 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024397884 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | We received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that stroke occurred with a possible relationship to the impella 5.5 used on this patient. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into the stroke/cva has been completed since the original report was submitted. the device was not returned for investigation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. d6a, d6b
27284 | 1220648-2024-10604 | 19216286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19216286&pc=OZD | 04/23/2024 | 06/24/2024 |  | 04/30/2024 |  |  | 04/24/2024 | 04/30/2024 |  |  | 05/31/2025 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374323 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp support was ongoing for a patient admitted post st elevation myocardial infarction and was needed for care escalation from a community hospital to their tertiary. the tertiary conducted a multivessel angioplasty. the pump supported for one day and then was explanted. the team observed ischemia that prompted the team to perform surgical intervention in the form of limb fasciotomy. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the limb ischemia ¿ irreversible and the thrombus issues has been completed since the original report was submitted. the product was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia and the thrombus was not determined.
27285 | 1220648-2024-10605 | 19216322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19216322&pc=OZD | 06/22/2023 | 05/22/2024 |  | 04/30/2024 |  |  | 04/04/2024 | 04/30/2024 |  |  | 05/31/2025 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024349813 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | A notification was received long-term outcome and quality indicator impella registry that the patient developed thrombocytopenia with a definite relationship to the impella 5.5 used on this patient. the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states). aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported thrombocytopenia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombocytopenia could not be determined.
27286 | 1220648-2024-10606 | 19217171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19217171&pc=OZD | 04/23/2024 | 06/12/2024 |  | 04/30/2024 |  |  | 04/24/2024 | 04/30/2024 |  |  |  |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433142 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient noted as high risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed in routine fashion, but it unable to flow above p-level p-2 without suction. the impella pump was re-implanted three times and despite volume and echocardiogram, fluoroscopy guidance for proper placement, the impella pump was still unable to flow above p-2. pump flows increased when pulled into the aorta. therefore, it was determined that it was not a pump motor issue. the patient's left ventricle cavity was very small and very short aortic root which caused the impella cp cannula to kink which was visible on fluoroscopy on the fourth attempt at implant. the decision was made to proceed with the case despite the inability to flow above p-2. post pci procedure, the impella pump was removed without difficulty. | [5 empty] | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk pci section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the low pump flow, delivery issues, and the introducer damage has been completed. pump was not returned for investigation. data logs were returned and reviewed. low flows of 1.5l/min were seen at p9 along with multiple suction events. per the clinical information provided, the patient had a small lv cavity and short aortic root causing the cannula to kink which resulted in multiple attempts of implanting the pump and getting optimal flows. the root cause of the low flow issue was most likely patient condition related to small patient vessels. in accordance with updated procedures, sections d6 have been revised from the initial submission.
27287 | 1220648-2024-10607 | 19218184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19218184&pc=OZD | 04/08/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 04/08/2022 | 04/30/2024 |  |  | 11/30/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158248 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/08/2022 |  |  |  |  | The user facility reported a 81 year old male patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support the physician experienced difficuty with the impella placement due to complex aortic kinking. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27288 | 1220648-2024-10608 | 19218239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19218239&pc=OZD | 05/11/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 05/11/2022 | 04/30/2024 |  |  | 11/30/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023164139 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/11/2022 |  |  |  |  | The complainant reported a 59 year old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was significant oozing. the patient was hemodynamically stable. the impella was removed at bedside. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27289 | 1220648-2024-10609 | 19218313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19218313&pc=OZD | 06/01/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 06/01/2022 | 04/30/2024 |  |  | 03/31/2022 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1469324 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/01/2022 |  |  |  |  | The user facility reported a 67-year-old male patient on impella cp for mechanical circulatory support. while on support, the patient developed coronary bleeding after the stent post inflation. preclose with perclose proglide times 2 was utilized. the device was explanted. the patient remained stable with no alarms during the case. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27290 | 1220648-2024-10610 | 19218504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19218504&pc=OZD | 08/16/2022 | 03/28/2025 |  | 04/30/2024 |  |  | 08/16/2022 | 04/30/2024 |  |  | 01/31/2024 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173676 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/16/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced bleeding and a hematoma at the repositioning sheath connection. the bleeding stopped after rotating the sheath, but a blood transfusion was required. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27291 | 1220648-2024-10611 | 19218512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19218512&pc=OZD | 09/21/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 09/21/2022 | 04/30/2024 |  |  | 06/30/2024 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235813 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 09/21/2022 |  |  |  |  | The user facility reported a 75-year-old male patient needing mechanical circulatory support. it was reported that while on impella support, ecchymosis was noted around the access site. it was later noted that the patient developed a large hematoma at the access site. it was reported that staff cleaned and irrigated the wound, then sutured vessel was closed following impella explant. no complications reported at the conclusion of this case. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27292 | 1220648-2024-10612 | 19218523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19218523&pc=OZD | 09/20/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 09/20/2022 | 04/30/2024 |  |  | 12/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180673 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/20/2022 |  |  |  |  | The complainant reported a 60 year old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding which was resolved with figure eight sutures. it was also reported that heparin was changed to sodium bicarbonate. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27293 | 1220648-2024-10613 | 19218536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19218536&pc=OZD | 02/10/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 02/10/2022 | 04/30/2024 |  |  | 03/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2022061524 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/10/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella pump experienced purge issues. to resolve the audio purge, the pressure was turned off. the physician was advised to check the parameters, reportedly there were no complications due to the purge issues. the patient was reported to be stable and was in the operating room for mitral valve replacement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
27294 | 1220648-2024-10614 | 19218538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19218538&pc=OZD | 03/28/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/28/2022 | 04/30/2024 |  |  | 12/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | IMPELLA CP |  |  | US |  | 2023180660 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/28/2022 |  |  |  |  | The user facility reported a 71 year old male patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support the physician was unable to reposition due to suspected pigtail in the papillary muscle. the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27295 | 1220648-2024-10615 | 19218550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19218550&pc=OZD | 03/29/2022 | 04/30/2024 |  | 04/30/2024 |  |  | 03/29/2022 | 04/30/2024 |  |  | 12/31/2023 |  | 04/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174804 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/29/2022 |  |  |  |  | The user facility reported a 79 year old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella support, there was a hematoma noted at the access site that appeared to be growing. it was also reported that patient¿s right distal pulses were diminishing. a perclose was used to achieve hemostasis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: access site bleeding. ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ limb ischemia. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27296 | 1220648-2024-10467 | 19200340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19200340&pc=OZD | 03/31/2022 | 03/25/2025 |  | 04/29/2024 |  |  | 03/31/2022 | 04/29/2024 |  |  | 12/31/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172932 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/31/2022 |  |  |  |  | The complainant reported a patient in cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced hemolysis requiring the impella to be pulled back 2 centimeters. it was further reported on (b)(6) 2022 that hemolysis continued despite optimal position being verified by echocardiogram, the clinician discussed lowering the pressure level to help reduce. the patient was transferred to another facility to have a workup done for left ventricular assist device or transplant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: ¿potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27297 | 1220648-2024-10468 | 19200411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19200411&pc=OZD | 03/31/2022 | 03/25/2025 |  | 04/29/2024 |  |  | 03/31/2022 | 04/29/2024 |  |  | 01/31/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185250 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/31/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding, and systemic heparin was withheld. manual compression was performed, and hemostasis achieved. the patient was reported to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27298 | 1220648-2024-10469 | 19200541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19200541&pc=OZD | 03/31/2022 | 03/25/2025 |  | 04/29/2024 |  |  | 03/31/2022 | 04/29/2024 |  |  | 07/31/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022115293 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/31/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support after a motor vehicle accident. while on support, the patient developed limb ischemia. the patient was reported to be stable, remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27299 | 1220648-2024-10470 | 19200698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19200698&pc=OZD | 04/01/2024 | 05/16/2024 |  | 04/29/2024 |  |  | 04/01/2024 | 04/29/2024 |  |  |  |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025456108 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 68-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the impella position in ao. confirmed by care team under x-ray and echo. patient taken to cardiac cath lab for removal and replacement. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. clinical details noted that pump was pulled into the aorta and confirmed by x-ray and echo. the root cause of the positioning issues was most likely due to a use issue as pump was pulled into the aorta. in accordance with updated procedures, sections d6 has been revised from the initial submission.
27300 | 1220648-2024-10471 | 19200736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19200736&pc=OZD | 04/01/2022 | 03/25/2025 |  | 04/29/2024 |  |  | 04/01/2022 | 04/29/2024 |  |  | 12/31/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170836 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2022 |  |  |  |  | The complainant reported a patient, race and ethnicity unknown, patient in st elevation myocardial infarction and cardiogenic shock noted as high risk percutaneous coronary insertion was implanted with an impella cp prior to surgery. while on support, echocardiography was performed and revealed a possible thrombus on the end of the impella catheter. the patient¿s ejection fraction improved to 55 percent and was successfully weaned from impella support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27301 | 1220648-2024-10472 | 19200747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19200747&pc=OZD | 03/15/2024 | 05/24/2024 |  | 04/29/2024 |  |  | 04/01/2024 | 04/29/2024 |  |  |  |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025441353 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 95 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/01/2024 |  |  |  |  | The user facility reported a 95-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. we received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that right common femoral artery thrombus occurred with a possible relationship to the impella cp used on this patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombus issue has been completed since the original report was submitted. the pump was not returned for investigation. as the necessary information was not provided, the root cause of the thrombus was not determined. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ in accordance with updated procedures, section h10 has been revised from the initial submission.
27302 | 1220648-2024-10474 | 19200823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19200823&pc=PYX | 02/24/2024 | 09/17/2025 |  | 04/29/2024 |  |  | 04/09/2024 | 04/29/2024 |  |  | 02/28/2025 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024348556 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/09/2024 |  |  |  |  | The user facility reported a 72-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the placement signal had dropped significantly during impella rp flex support. the patient hemodynamics remained stable, and both the motor current and flows remained unchanged despite the noted issue. the staff was advised that the signal may be unreliable, and that the hemodynamics could still be monitored via the swan catheter. intervention was not required as the mc and flow remained unchanged. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. according to the clinical, immediately after beginning impella rp support the placement signal began trending up and down into the negative region. no attempts were made to resolve the issue as the patient hemodynamics remained normal. no product was returned for a visual inspection. data log analysis confirmed a dp sensor issue with an initial increase and decrease in placement signal mirrored by a decrease and increase in flow. motor current, cvp, and speed remained unchanged. the cause of the placement signal issue could not be determined because no product was returned and not enough clinical information was given. during data log analysis a high motor current spike was also observed. this spike and the shift in ps, decrease in flow and speed deviation that followed are consistent with biomaterial ingestion. the most likely cause of the low pump flow was biomaterial. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d4 model number updated. d6b explant date was corrected from the initial manufacturer device report number 1220648-2024-10474. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures from the initial manufacturer device report number 1220648-2024-10474. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-10474.
27303 | 1220648-2024-10475 | 19200875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19200875&pc=OZD | 04/10/2024 | 05/21/2024 |  | 04/29/2024 |  |  | 04/10/2024 | 04/29/2024 |  |  |  |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024435787 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | The user facility reported a 51-year-old male patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient developed a left ventricle thrombus noted with echo along with a clot in the left atrium (la). the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). | The investigation into the thrombus issue has been completed since the original report was submitted. the pump was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27304 | 1220648-2024-10476 | 19201020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201020&pc=OZD | 03/12/2024 | 04/29/2024 |  | 04/29/2024 |  |  | 03/12/2024 | 04/29/2024 |  |  | 08/31/2025 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024393617 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support, and a large hematoma was noted above the surgical incision where axillary graft was placed. during the procedure the pocket was not fully closed. the activated clotting time (act) was supratherapeutic in the 500s, and the skin was stapled together in anticipation of eventual explant and terminal closure of the graft. upon arrival at intensive care unit, manual pressure was applied to reduce hematoma, pressure dressing was placed overnight, and the hematoma was successfully reduced as well as act down to goal range. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. however, an investigation was conducted and determined the root cause of access site bleeding was anti-coagulation management because of the patient's high activated clotting time (act) values, skin was stapled together in anticipation of eventual explant and terminal closure of the graft. the hematoma was successfully reduced as well as act down to goal range.
27305 | 1220648-2024-10477 | 19201026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201026&pc=OZD | 04/04/2022 | 03/25/2025 |  | 04/29/2024 |  |  | 04/04/2022 | 04/29/2024 |  |  | 06/30/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022099873 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2022 |  |  |  |  | The complainant reported a 63-year-old patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella access site. femostop and manual pressure was applied to the access site and all anti coagulation was withheld to achieve hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27306 | 1220648-2024-10478 | 19201049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201049&pc=OZD | 04/23/2024 | 06/11/2024 |  | 04/29/2024 |  |  | 04/23/2024 | 04/29/2024 |  |  | 02/28/2025 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338517 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/23/2024 |  |  |  |  | The user facility reported a 36-year-old female with myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the patient lost distal pulses, so out of concern for ischemia, the medical staff removed the impella cp. during the explant, the patient had increased aspartate aminotransferase, and there was a concern for hemolysis. the pump was found to be under the mitral valve. when the pump was removed, a thrombus was discovered. the patient survived the event. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. h.10 inadvertently mentioned the device was not returned on the previous manufacture device report and the device was discarded. | The investigation for the reported thrombus, hemolysis, positioning, and limb ischemia has been completed. the impella device was not returned for evaluation. data logs were reviewed finding no positioning issue observed. no motor current spikes occurred. no indication for hemolysis can be confirmed. the cause of the positioning issue most likely use related due to improper technique as echo showed impella under mitral valve. the root cause of the limb ischemia and the thrombus could not be determined.
27307 | 1220648-2024-10479 | 19201059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201059&pc=OZD | 04/16/2024 | 06/03/2024 |  | 04/29/2024 |  |  | 04/16/2024 | 04/29/2024 |  |  |  |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025463862 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an impella 5.5 device implanted as elective placement for mechanical circulatory support. approximately 15 minutes after the impella device was turned on, the motor current spiked followed by a brief controller error which immediately resolved itself, and then an "impella stopped motor current high" alarm occurred. three attempts to restart the pump at p-level p-3 were made before disconnecting impella pump and rebooting automated impella controller (aic). the impella pump was restarted on same aic after it powered back on and was successfully ramped to p-level p-6. all metrics were reported normal from tech support. tech support suspected a clot was ingested into pump; however, the pump was not damaged according to metrics. the neuro function of the patient was unknown due to sedation. the care team continued to monitor patient's neuro status. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock - section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ | The investigation into the temporary pump stop issue has been completed since the original report was filed. the device was returned for investigation. the pump was able to be restarted on the same console and the controller error alarm was also resolved. the pump ran successfully for the remainder of the case. the root cause of the temporary pump stop issue was not determined since the issue could not be reproduced.
27308 | 1220648-2024-10480 | 19201240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201240&pc=OZD | 04/23/2024 | 06/11/2024 |  | 04/29/2024 |  |  | 04/23/2024 | 04/29/2024 |  |  |  |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025489020 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | The user facility reported a 43-year-old male with non-st elevated myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. after delivering to the left ventricle, the pump came out of position. the pump needed to be replaced due to a kink developing when the medical staff tried to redeliver the pump to the left ventricle. the impella cp was explanted and a second impella cp was implanted. the patient survived the event, and no patient harm has been reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. data logs confirmed pump was in aortic area (about 15 minutes into the case). pump then was stopped manually and disconnected from the console. product was not returned for review. based on clinical description and data analysis, the root cause of positioning issue was most likely use related.
27309 | 1220648-2024-10481 | 19201355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201355&pc=OZD | 04/07/2022 | 03/25/2025 |  | 04/29/2024 |  |  | 04/07/2022 | 04/29/2024 |  |  | 11/30/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167163 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. during placement of the device, the device could not advance past the iliac bifurcation. the impella was removed, and the common femoral artery was ballooned. the physician attempted to insert the impella again, but they were unable to advance past the abdominal aorta. the case was aborted, and the sheath was removed with successful hemostasis. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27310 | 1220648-2024-10482 | 19201428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201428&pc=OZD | 04/16/2024 | 06/10/2024 |  | 04/29/2024 |  |  | 04/16/2024 | 04/29/2024 |  |  | 08/31/2025 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025443756 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/16/2024 |  |  |  |  | The user facility reported a 47-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed had limb ischemia. no pulses were found in the patient's right leg (impella side). the purge solution changed to dextose 5% in water (d5w) with heparin and plan to wean off levo, let the heart rest and reassess later. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). | The investigation into the access site bleeding ¿ minor and the limb ischemia - reversible issues has been completed since the original report was submitted. the device was not returned for investigation. per the clinical information provided, the root cause of the access site bleeding issue was not determined since product was not returned and sufficient information is not provided in the clinical description. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27311 | 1220648-2024-10483 | 19201476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201476&pc=OZD | 05/10/2022 | 04/29/2024 |  | 04/29/2024 |  |  | 05/10/2022 | 04/29/2024 |  |  | 02/29/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2023186334 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | A |  |  |  | Initial submission |  |  |  |  | 05/10/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy, race and ethnicity unknown, was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced hemolysis. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
27312 | 1220648-2024-10484 | 19201593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201593&pc=OZD | 05/12/2022 | 03/25/2025 |  | 04/29/2024 |  |  | 05/12/2022 | 04/29/2024 |  |  | 10/31/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149842 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2022 |  |  |  |  | The complainant reported a patient undergoing aortic valvuloplasty was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved with stoppage of heparin. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27313 | 1220648-2024-10485 | 19201614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201614&pc=OZD | 04/08/2022 | 03/25/2025 |  | 04/29/2024 |  |  | 04/08/2022 | 04/29/2024 |  |  | 01/31/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187012 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed oozing at the impella cp access site and the original dressing from the cath lab had become saturated. the dressing was changed and a clotting agent was added to the surface of the access site. the site was re-dressed and the angle entry was maintained. the patient remained on impella cp support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27314 | 1220648-2024-10486 | 19201619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201619&pc=OZD | 04/04/2022 | 03/25/2025 |  | 04/29/2024 |  |  | 04/04/2022 | 04/29/2024 |  |  | 01/31/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186253 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2022 |  |  |  |  | The complainant reported a 63-year-old patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced oozing at the impella access site. pressure was applied to the site, the blue hub was advanced to skin, and angle matching performed to achieve hemostasis and the patient was transfused with two units of packed red blood cells. it was also reported following emesis, the patient experienced ectopy and hypotension with associated suction alarm. the patient was given 250 fluid bolus and 200 mcg of neo-synephrine. echocardiology was performed at bedside and showed the impella was near posterior wall and was advanced from 3cm to 3.5 cm. patient remained hypotensive, and an additional 500 fluid bolus was given. no further episodes of ectopy were reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27315 | 1220648-2024-10487 | 19201664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201664&pc=OZD | 08/23/2022 | 09/26/2024 |  | 04/29/2024 |  |  | 08/23/2022 | 04/29/2024 |  |  | 03/31/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023203581 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/23/2022 |  |  |  |  | The complainant reported a patient with heart failure and stroke was implanted with an impella 5.5 for mechanical circulatory support. while on support, heparin was withheld. hemostasis was achieved through surgical closure. after transfer to the unit, the surgeon noted an unstable clot in the descending aorta. activated clotting time (act) was initiated with full heparin purge. the patient expired while on impella support. no allegations were made towards the impella for cause of death. | The procedural outcome noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-10487 was provided in sections b2, b5, d6a, d6b, h1, and h6.
27316 | 1220648-2024-10488 | 19201679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201679&pc=OZD | 04/04/2022 | 03/25/2025 |  | 04/29/2024 |  |  | 04/04/2022 | 04/29/2024 |  |  | 02/29/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182982 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, the patient experienced bleeding from the gastrointestinal (gi) system and the impella insertion site. as a result, heparin was temporarily removed from the purge solution. patient condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27317 | 1220648-2024-10489 | 19201751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201751&pc=OZD | 09/11/2022 | 04/29/2024 |  | 04/29/2024 |  |  | 09/11/2022 | 04/29/2024 |  |  | 08/31/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022124737 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 09/11/2022 |  |  |  |  | The user facility reported a 44-year-old patient was on impella cp support and the patient's extremities (foot) was reportedly very cold due to extracorporeal membrane oxygenation and impella cannulas (lack of perfusion). the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27318 | 1220648-2024-10490 | 19201827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201827&pc=OZD | 04/08/2022 | 03/25/2025 |  | 04/29/2024 |  |  | 04/08/2022 | 04/29/2024 |  |  | 01/31/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186682 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there was narrow pulse pressure and poor distal flow to pedal pulse from impella groin site. antegrade sheath was placed for improved distal flow to the foot. in addition, there were suction alarms on impella support with elevated central venous pressure. they were unable to resolve by checking position and fluid management. decision was made to escalate care tandem heart and the impella cp was successfully removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27319 | 1220648-2024-10491 | 19201857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201857&pc=OZD | 09/21/2022 | 03/25/2025 |  | 04/29/2024 |  |  | 09/21/2022 | 04/29/2024 |  |  | 02/29/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202873 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, slight oozing at the groin was noted. two preclose sutures were put in place with no pre dilation done before 14f sheath. patient condition was reported as stable after the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27320 | 1220648-2024-10492 | 19201999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19201999&pc=OZD | 04/05/2022 | 03/28/2025 |  | 04/29/2024 |  |  | 04/05/2022 | 04/29/2024 |  |  | 09/30/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142723 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced sepsis. the physician lowered the impella to setting p5 and subsequently explanted it. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | When diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: -patient symptoms and medical history obtained by the treating clinician -physical examination findings -results of special investigations : laboratory test results & imaging studies -microbiological culture and sensitivity results -response to previous treatments and therapies -any relevant epidemiological information or exposure history -concomitant devices in use -preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. -care of access site because this evidence has not been provided, the root cause of the infection cannot be determined. this report is being filed as part of a retrospective review of historical records.
27321 | 1220648-2024-10493 | 19202038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19202038&pc=OZD | 09/23/2022 | 03/31/2025 |  | 04/29/2024 |  |  | 09/23/2022 | 04/29/2024 |  |  | 06/30/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023229192 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/23/2022 |  |  |  |  | The user facility reported a 51-year-old female on impella 5.5 for mechanical circulatory support. it was reported that patient experienced limb ischemia. the patient had vascular consult for necrotic foot amputation. the patient declined and palliative consult was requested. no further information was provided. | [5 empty] | The impella 5.5 device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
27322 | 1220648-2024-10494 | 19202068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19202068&pc=OZD | 04/05/2022 | 05/20/2025 |  | 04/29/2024 |  |  | 04/05/2022 | 04/29/2024 |  |  | 10/31/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156966 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient was started on continuous renal replacement therapy and slow continuous ultrafiltration due to his kidney and liver function getting worse. the impella was pulled back almost 3 centimeters due to echocardiogram revealing the impella being too deep in the ventricle, reportedly this improved the flows. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27323 | 1220648-2024-10497 | 19202167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19202167&pc=OZD | 04/07/2022 | 03/31/2025 |  | 04/29/2024 |  |  | 04/07/2022 | 04/29/2024 |  |  | 10/31/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154731 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the pump flow was low and had suction issues. imaging showed the pump was not out towards the apex and the inlet was possibly obstructed. the impella was explanted, rinsed in dextrose 5% in water, and implanted again. after the procedure, impella was left in position to give patient some rest. however, due to the patient¿s small body habitus, the distal leg perfusion was minimal. as a result, impella was successfully weaned and explanted. patient condition was stable after the procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27324 | 1220648-2024-10498 | 19202169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19202169&pc=OZD | 04/16/2024 | 06/04/2024 |  | 04/29/2024 |  |  | 04/16/2024 | 04/29/2024 |  |  |  |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485145 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/16/2024 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced ectopy. the physician attempted to reposition but was unable to get into adequate position to resolve ectopy. the decision was made by the medical doctor to remove and rewire with the same result. ectopy caused hypotension so the impella device was removed with an increase in blood pressure. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported arrhythmia and positioning issue has been completed. the device was not returned for evaluation. based on clinical description and data analysis, the root cause of positioning issue was most likely use related as they were unable to successfully reposition the device into an acceptable position. the root cause of the arrhythmia could not be determined. added- h6 impact code 4628.
27325 | 1220648-2024-10499 | 19202413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19202413&pc=OZD | 04/04/2022 | 03/31/2025 |  | 04/29/2024 |  |  | 04/04/2022 | 04/29/2024 |  |  | 01/31/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183753 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient began to experience ectopy. during attempts to reposition the pump, the pump was pulled back entirely into aorta inadvertently. the pump was stopped, removed and reinserted without further issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27326 | 1220648-2024-10500 | 19202497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19202497&pc=OZD | 04/16/2024 | 07/02/2024 |  | 04/29/2024 |  |  | 04/17/2024 | 04/29/2024 |  |  |  |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2025441900 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Life Threatening | Disability | Required Intervention | Death |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported the user facility had a patient on impella 5.5 support in need of care escalation from the prior impella cp device and extracorporeal membrane oxygenation (ecmo) supports. the impella 5.5 was placed and supported for just under three days. the patient then expired. the death occurred after a cardiac rupture. the rupture/cardiac tamponade was treated by pericardial drainage and blood infusion back to the patient. it was noted the death was from the underlying condition and multiple organ failure. the tamponade cause and relationship to the use of impella device could not be denied. | [5 empty] | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torqueing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation into the ventricle perforation issue has been completed since the original report was submitted. no product was returned for evaluation. data logs were not returned. clinical details noted that cardiac tamponade was noted after impella 5.5 procedure and pericardial effusion was performed. active bleeding was still noted after drainage and blood products were administered. no impairment of impella operations were noted. the root cause of the ventricle perforation could not be determined as no product or logs were returned for evaluation. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ e.1 first name was revised as previously entered incorrectly on the initial manufacturer report number 1220648-2024-10500. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer report number 1220648-2024-10500 was submitted. h.10 additional manufacturer narrative instructions for use - a portion was omitted from the initial manufacturer report number 1220648-2024-10500 and has been added.
27327 | 1220648-2024-10502 | 19202664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19202664&pc=OZD | 04/02/2022 | 03/31/2025 |  | 04/29/2024 |  |  | 04/02/2022 | 04/29/2024 |  |  | 12/31/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180157 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2022 |  |  |  |  | The user facility reported a 50-year-old female was implanted with impella cp for mechanical circulatory support. it was reported that while on impella support, the patient developed a small hematoma at the groin site. the sheath did not remain completely in the artery, and a great amount of bleeding into the thigh was observed as a result. it was reported the medical doctor (md) made cuts to allow for increased sheath length, then the pump was repositioned under echocardiogram (echo). it was noted that bleeding seemed to be under control. two units of packed red blood cells (prbc) were given. it was later reported that the patient continued to bleed, and the md was concerned that the sheath was not in the artery due to the patient¿s size. the impella cp was weaned and removed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27328 | 1220648-2024-10501 | 19202666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19202666&pc=OZD | 04/25/2022 | 04/29/2024 |  | 04/29/2024 |  |  | 04/25/2022 | 04/29/2024 |  |  | 11/30/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165169 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/25/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the position of the impella was deep. the device was repositioned under echocardiogram (echo) and flows improved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27329 | 1220648-2024-10503 | 19202704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19202704&pc=OZD | 04/17/2024 | 06/05/2024 |  | 04/29/2024 |  |  | 04/17/2024 | 04/29/2024 |  |  |  |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438779 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2024 |  |  |  |  | The user facility reported a 55-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that signs of ischemia were observed. pulses were lost. the pump was explanted after three days and replaced. the patient was reported as stable. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
27330 | 1220648-2024-10504 | 19202801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19202801&pc=OZD | 04/24/2022 | 03/31/2025 |  | 04/29/2024 |  |  | 04/24/2022 | 04/29/2024 |  |  | 11/30/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168147 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. at p2 the cp had a series a suction alarms. after continuous suction the cp was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27331 | 1220648-2024-10505 | 19202960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19202960&pc=OZD | 04/18/2024 | 05/22/2024 |  | 04/29/2024 |  |  | 04/18/2024 | 04/29/2024 |  |  |  |  | 04/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1335325 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | Component Missing | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility's biomedical engineer reported that the automated impella controller (aic) had a purge disc that was cracked and was missing a piece. the issue was discovered post impella pump removal. there was no patient harm or injury. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the console was evaluated and the purge disc retainer was found to be broken. the root cause of the issue was a defective component. e.2 and e.3 revised as previously entered incorrectly.
27332 | 1220648-2024-10506 | 19202964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19202964&pc=OZD | 05/11/2022 | 04/29/2024 |  | 04/29/2024 |  |  | 05/11/2022 | 04/29/2024 |  |  | 01/31/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180349 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/11/2022 |  |  |  |  | The complainant reported a patient required a second impella device for support. the initial device had a suction alarm and massive flow reduction; it was suspected to be from a clot in the cannula. the clot was confirmed, the impella was explanted, and support was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) section: ¿¿potential adverse events (united states) . acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. this report is being filed as part of a retrospective review of historical records.
27333 | 1220648-2024-10507 | 19202978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19202978&pc=OZD | 05/01/2022 | 03/31/2025 |  | 04/29/2024 |  |  | 05/01/2022 | 04/29/2024 |  |  | 09/30/2022 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1505023 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, it was found that the doppler was unable to find lower extremity pulses bilaterally. it is unknown how the ischemia was resolved. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia: when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
27334 | 1220648-2024-10508 | 19203075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19203075&pc=OZD | 10/24/2022 | 05/28/2025 |  | 04/29/2024 |  |  | 10/24/2022 | 04/29/2024 |  |  | 03/31/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214345 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2022 |  |  |  |  | The user facility reported a 68-year-old male on impella cp for mechanical circulatory support. the patient experienced hemolysis. it was reported that the patient was maxed on epi, levo and vaso; the patient was only on p6 due to hemolysis. haptoglobin was less than one multiple times. escalation of the patient to impella 5.5 was discussed. no further information was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27335 | 1220648-2024-10509 | 19203129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19203129&pc=OZD | 02/21/2022 | 03/31/2025 |  | 04/29/2024 |  |  | 02/21/2022 | 04/29/2024 |  |  | 11/30/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161402 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2022 |  |  |  |  | It was reported that the placement procedure for a patient needing an impella cp for mechanical circulatory support was aborted due to the patient's anatomy. no further details were provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27336 | 1220648-2024-10510 | 19203165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19203165&pc=OZD | 09/20/2022 | 03/31/2025 |  | 04/29/2024 |  |  | 09/20/2022 | 04/29/2024 |  |  | 05/31/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227681 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient a small amount of oozing coming from groin impella access site. the perclose devices were tightened however one broke and heparin was withheld. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27337 | 1220648-2024-10511 | 19203256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19203256&pc=OZD | 05/10/2022 | 03/31/2025 |  | 04/29/2024 |  |  | 05/10/2022 | 04/29/2024 |  |  | 02/29/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195642 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/10/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a positioning issue in which the cp pulled back into the aorta. the physician decided to remove the cp, rewired the left ventricle, and reinserted the cp. the following day, it was found that there was no pedal pulse in one of the patient's feet. the resolution is unknown, but the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27338 | 1220648-2024-10512 | 19203262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19203262&pc=OZD | 04/18/2024 | 05/30/2024 |  | 04/29/2024 |  |  | 04/18/2024 | 04/29/2024 |  |  |  |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447391 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported a 57-year-old male of unknown race in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient developed a hematoma during impella cp support. heparin was discontinued in response to the condition and a sandbag was placed to assist with managing the hematoma. the patient was given one unit of blood and support with the implanted pump continued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
27339 | 1220648-2024-10513 | 19203391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19203391&pc=OZD | 04/18/2024 | 06/10/2024 |  | 04/29/2024 |  |  | 04/18/2024 | 04/29/2024 |  |  |  |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447111 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | Thrombocytopenia | Thrombosis/Thrombus | [7 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 04/18/2024 |  |  |  |  | The user facility reported a 64-year-old female of unknown ethnicity in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a hematoma and oozing at the access site. the bleeding was resolved with a femstop and the patient was given multiple units of packed red blood cells. additionally, the patient lost pulses in the left lower extremity. the sheath was repositioned to regain pulses. the patient was reported as stable. | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured acute blood loss anemia with a "definite" relationship to the impella device used. 2 units of blood transfused. anemia likely due to acute blood loss from left groin hematoma + right groin oozing. it was reported that the patient is not recovered, and the event is not resolved. | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured acute blood loss anemia with a "definite" relationship to the impella device used: anemia likely due to acute blood loss from left groin hematoma + right groin oozing. hgb levels remained low but stable through discharge. it was reported that the patient is not recovered, and the event is not resolved. additionally, patient endured thrombus and thrombocytopenia with a "probable" relationship to the impella device used: ¿upon impella removal thrombus was noted at access sheath site. thrombectomy performed with stent placement. the patient recovered with no sequelae. 2 units of blood transfused. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Section b5 and h6 were updated as per additional information received. | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. per the clinical information provided, the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the root cause of the thrombocytopenia was unable to be determined as insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and the thrombus was not determined. | B5 and h6 were updated as per additional information received.
27340 | 1220648-2024-10514 | 19203424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19203424&pc=OZD | 03/22/2022 | 04/29/2024 |  | 04/29/2024 |  |  | 03/22/2022 | 04/29/2024 |  |  | 11/30/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168483 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/22/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed limb ischemia. the pulses were absent on the bilateral lower extremity but were warm to the touch. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿. this report is being filed as part of a retrospective review of historical records.
27341 | 1220648-2024-10515 | 19203559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19203559&pc=OZD | 04/15/2022 | 03/31/2025 |  | 04/29/2024 |  |  | 04/15/2022 | 04/29/2024 |  |  | 04/30/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022076650 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced suction alarms. as a result, pump was explanted. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27342 | 1220648-2024-10516 | 19203647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19203647&pc=OZD | 04/19/2024 | 06/06/2024 |  | 04/29/2024 |  |  | 04/19/2024 | 04/29/2024 |  |  | 08/31/2025 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024384548 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient undergoing high-risk percutaneous coronary intervention was planned for an impella cp device placement for mechanical circulatory support. however, the delivery of the impella cp pump was postponed after the staff was unable to advance the 10 french dilator. it was stated the patient had difficult, calcific plague within the vessel which caused too much resistance for insertion. as such, the implant was aborted prior to the pump being inserted. there was no patient harm. | [5 empty] | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the difficulty inserting/removing introducer issue has been completed since the original report was submitted. the device was not returned for investigation. as per the given clinical, due to a difficult calcific plaque within the vessel, the decision was made to postpone the case. the cause of the difficulty inserting introducer was patient condition due to diseased vessels, based on given clinical details.
27343 | 1220648-2024-10517 | 19203713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19203713&pc=OZD | 04/11/2022 | 03/31/2025 |  | 04/29/2024 |  |  | 04/11/2022 | 04/29/2024 |  |  | 10/31/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154095 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient's feet were found to be mottled bilaterally. flow was confirmed via arterial duplex. the impella was weaned, and upon removal, some clot burden was noted on the pigtail and repositioning sheath. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records. | Thrombus: when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions. ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts. ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. limb ischemia. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
27344 | 1220648-2024-10518 | 19203813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19203813&pc=OZD | 04/19/2024 | 05/28/2024 |  | 04/29/2024 |  |  | 04/19/2024 | 04/29/2024 |  |  | 11/30/2025 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024426412 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported a 47-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that while on impella cp support, the patient developed limb ischemia. a femorofemoral bypass was performed to resolve the ischemia. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the pump was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
27345 | 1220648-2024-10519 | 19203845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19203845&pc=OZD | 04/20/2024 | 06/03/2024 |  | 04/29/2024 |  |  | 04/20/2024 | 04/29/2024 |  |  |  |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442913 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2024 |  |  |  |  | The user facility reported a 63-year-old male was implanted with an impella 5.5 as a bridge to transplant. it was reported that he patient exhibited a thrombus on the catheter which was observed on an echocardiogram. heparin was administered to successfully treat the thrombus. the patient remains on support and is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation into the thrombus issue has been completed since the original report was submitted. the pump was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
27346 | 1220648-2024-10520 | 19204370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19204370&pc=OZD | 04/21/2022 | 04/29/2024 |  | 04/29/2024 |  |  | 04/21/2022 | 04/29/2024 |  |  | 01/31/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189241 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/21/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced limb ischemia which was resolved with fem to fem bypass. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
27347 | 1220648-2024-10521 | 19204443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19204443&pc=OZD | 04/22/2024 | 06/10/2024 |  | 04/29/2024 |  |  | 04/22/2024 | 04/29/2024 |  |  | 11/30/2025 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434127 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/22/2024 |  |  |  |  | The user facility reported a 58-year-old male undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient was found to have left side weakness. a computed tomography scan revealed acute right middle cerebral artery infarct. the impella 5.5 was explanted. some clots on the catheter and in the outlet cage were found following removal. the patient had no change in neurology status or hemodynamics. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the reported stroke and thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke and thrombus could not be determined. d6.
27348 | 1220648-2024-10522 | 19205333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19205333&pc=OZD | 05/11/2022 | 04/29/2024 |  | 04/29/2024 |  |  | 05/11/2022 | 04/29/2024 |  |  | 02/29/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196274 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Pump | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/11/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump stopped and did not restart. no additional information was available. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
27349 | 1220648-2024-10523 | 19205477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19205477&pc=OZD | 04/22/2024 | 06/11/2024 |  | 04/29/2024 |  |  | 04/22/2024 | 04/29/2024 |  |  |  |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025458198 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/22/2024 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that they were unable to advance through the patient's anatomy during attempted delivery due to resistance in the innominate artery. after multiple failed attempts, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torqueing the impella, adjustments should be performed under imaging guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most patient condition since the pump could not be implanted due to the patient anatomy.
27350 | 1220648-2024-10524 | 19205601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19205601&pc=OZD | 04/07/2022 | 03/31/2025 |  | 04/29/2024 |  |  | 04/07/2022 | 04/29/2024 |  |  | 01/31/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184091 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a diagnostic identified a clot in the left main ventricle (lm), left anterior descending (lad) and circumflex arteries (cx). the plan was to place the impella to support percutaneous coronary intervention (pci). penumbra system was utilized to extract a huge amount of the clot from the lad and cx. the suction was ongoing due to hypovolemic state. lactated ringer¿s solution (lr), blood products and albumin were provided to the patient. volume loss was resolved, however despite fem-stop and access site being redressed numerous times, bleeding was continuous. the patient stabilized and the bleeding discontinued. it was also reported that the lad was stented on impella support, and the patient experienced multiple episodes of ventricular tachycardia (vt) and shocks in the cardiovascular lab (cvl) and another episode while in the intensive care unit (icu). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27351 | 1220648-2024-10525 | 19205633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19205633&pc=OZD | 05/26/2022 | 03/28/2025 |  | 04/29/2024 |  |  | 05/26/2022 | 04/29/2024 |  |  | 08/31/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022129797 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced a small hematoma at the impella cp insertion site at the groin. heparin was discontinued from the purge solution and systemically after the patient's activated clotting time (act), international normalized ratio (inr) and anti¿factor xa coagulation results were higher than expected. the groin site began bleeding after patient was turned and moved. the site was redressed and manual pressure held for 20 minutes, but the site was still bleeding so vascular surgery was consulted. vascular surgery placed a suture at the groin site and the patient was transfused with fresh frozen plasma to treat the high inr. the patient was successfully weaned and taken to the operating room where a cut down was performed on the left femoral artery. the impella cp was removed with the artery and wound sutured close. patient remained stable after explant and extubation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27352 | 1220648-2024-10526 | 19205694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19205694&pc=OZD | 04/07/2022 | 03/28/2025 |  | 04/29/2024 |  |  | 04/07/2022 | 04/29/2024 |  |  | 12/31/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178839 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/07/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while inserting a seven by forty-five-inch destination sheath over the table wire dilator, it was noted that the wire prolapsed and caused a dissection of the right common iliac (rci). the patient decompensated which required pressors and a blood transfusion to be performed. an unsuccessful attempt of a balloon tamponade to the rci was made. post dissection and unsuccessful tamponade, the impella was removed so that three (3) covered stents and one (1) vascular plug could be placed to achieve hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27353 | 1220648-2024-10527 | 19205744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19205744&pc=OZD | 03/19/2024 | 04/30/2024 |  | 04/29/2024 |  |  | 04/23/2024 | 04/29/2024 |  |  | 10/31/2025 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445631 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 94 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | The user facility reported a 94-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. per medwatch report 2301650000-2024-8005, it was reported that the left ventricle was perforated by impella cp requiring left ventricle repair. the impella cp was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella 5.5 with smart assist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torqueing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | E.4 was revised to reflect that the initial reporter also sent the report to the fda.
27354 | 1220648-2024-10528 | 19205868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19205868&pc=OZD | 03/28/2022 | 03/28/2025 |  | 04/29/2024 |  |  | 03/28/2022 | 04/29/2024 |  |  | 12/31/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181128 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon placement of the impella cp, the patient sustained ventricular tachycardia and required 1 shock at 200 joules which converted to normal sinus rhythm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | To make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation -timing of the stroke in relation to impella support (during or post-implant) -detailed description of the symptoms experienced by the patient -neurological examination findings and any focal deficits observed -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic) -laboratory test results, including coagulation profiles and cardiac markers -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications -response to any previous anticoagulation or antiplatelet therapies -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. this report is being filed as part of a retrospective review of historical records.
27355 | 1220648-2024-10529 | 19205889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19205889&pc=OZD | 05/11/2022 | 04/29/2024 |  | 04/29/2024 |  |  | 05/11/2022 | 04/29/2024 |  |  | 01/31/2024 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190903 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/11/2022 |  |  |  |  | The complainant reported a patient presented for coronary artery bypass graft (cabg) was implanted with an impella cp device for mechanical circulatory support. during support, the patient underwent off-pump coronary artery bypass grafting x4. the patient was very labile coming out of or and experienced ventricular tachycardia prior to leaving the or. frequent suction during hypotensive episodes, however blood and blood products were given. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
27356 | 1220648-2024-10530 | 19206387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19206387&pc=OZD | 03/28/2022 | 03/28/2025 |  | 04/29/2024 |  |  | 03/28/2022 | 04/29/2024 |  |  | 11/30/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023169192 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, the patient experienced left middle cerebral artery (mca) occlusion and had a stroke. as a result, impella cp pump was explanted. patient's condition was stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | To make a determination as to the cause of the stroke, the following clinical information must be considered: patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation. timing of the stroke in relation to impella support (during or post-implant). detailed description of the symptoms experienced by the patient. neurological examination findings and any focal deficits observed. diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic). laboratory test results, including coagulation profiles and cardiac markers. any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications. response to any previous anticoagulation or antiplatelet therapies. evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. this report is being filed as part of a retrospective review of historical records.
27357 | 1220648-2024-10531 | 19207136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19207136&pc=OZD | 04/23/2024 | 06/25/2024 |  | 04/29/2024 |  |  | 04/23/2024 | 04/29/2024 |  |  | 09/30/2025 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403175 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | The user facility reported a 66 year-old male with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the patient developed a hematoma on the right thigh, at the distal site of the distal perfusion catheter. the medical staff identified extravasation and consulted the vascular staff. the medical staff then performed a cut down, alongside a vascular repair and one unit of blood product. the patient remains on support. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for the reported vascular damage has been completed. the product was not returned for investigation. a right thigh hematoma developed at the impella access site after inserting a 6fr distal perfusion catheter. contrast injection through the impella repo sheath guidewire reaccess port revealed bleeding in a distal vessel. vascular surgery repaired the unrelated bleeding sites and closed the surgical site, keeping the impella in place. the cause of the vascular damage issue was not determined and unrelated to the impella since bleeding/vascular repair at non-impella site per clinical review. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures.
27358 | 1220648-2024-10533 | 19207581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19207581&pc=OZD | 04/24/2024 | 06/04/2024 |  | 04/29/2024 |  |  | 04/24/2024 | 04/29/2024 |  |  |  |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438837 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2024 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a hematoma at the insertion site. attempts were made to reposition and angle match, however the decision was made to explant due to growing hematoma. the patient was reported as stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to anticoagulation management since the act the time of the bleed was 300 which is more than the range limit.
27359 | 1220648-2024-10532 | 19208141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19208141&pc=OZD | 05/08/2022 | 03/28/2025 |  | 04/29/2024 |  |  | 05/08/2022 | 04/29/2024 |  |  | 12/31/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171922 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/08/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient's foot was cold. the staff was unable to obtain the right lower extremity doppler pulse. return of spontaneous circulation was achieved prior to going to catheterization lab. patient remained on support and was sent to unit. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
27360 | 1220648-2024-10534 | 19209322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19209322&pc=OZD | 05/12/2022 | 03/10/2025 |  | 04/29/2024 |  |  | 05/12/2022 | 04/29/2024 |  |  | 12/31/2023 |  | 04/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023171561 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/12/2022 |  |  |  |  | The complainant reported the 73 year old male patient was escalated to 5.5 for hemodynamic support. it was reported while on support the patient was transfused with a unit of packed red blood cells after achieving hemostasis from the impella 5.5 access site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27361 | 1220648-2024-10422 | 19190638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19190638&pc=OZD | 05/04/2022 | 04/26/2024 |  | 04/26/2024 |  |  | 05/04/2022 | 04/26/2024 |  |  | 12/31/2023 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180805 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Low Blood Pressure/ Hypotension | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient became hypotensive and bradycardia, during case perforation noticed on angiogram. the nurse updated the levophed medication level from 6 to 8 to address the hypotension. it was reported the patient was tolerating the impella cp very well. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation...¿ ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿
27362 | 1220648-2024-10423 | 19190848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19190848&pc=OZD | 06/05/2023 | 06/10/2024 |  | 04/26/2024 |  |  | 03/27/2024 | 04/26/2024 |  |  | 12/31/2024 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024306888 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 53 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Retroperitoneal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/27/2024 |  |  |  |  | The user facility reported a 53-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. we received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that left lower extremity limb ischemia occurred with a possible relationship to the impella 5.5 used on this patient. | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured compartment syndrome with a "possible" relationship to the impella device used: ¿compartment syndrome after retroperitoneal bleed may have been cause by placement of arterial femoral line. hgb dropped from 13.6 to 7.8. multiple blood products were given + laparotomy was performed¿. the patient recovered with no sequelae. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into the limb ischemia ¿ reversible and the vascular damage - peripheral vessel issues has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the root cause of the retroperitoneal bleed was not determined as no product and insufficient clinical details were provided. impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | Section b5 and h6 were updated as per new additional information received.
27363 | 1220648-2024-10424 | 19190889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19190889&pc=OZD | 04/14/2024 | 11/26/2024 |  | 04/26/2024 |  |  | 04/14/2024 | 04/26/2024 |  |  |  |  | 04/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1189539 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/14/2024 |  |  |  |  | No patient involvement: the us complainant had a console malfunction. the aic (automated impella controller) rotary knob was not operating as specified and so a loaner console will be utilized for patient use. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into automated impella controller (aic) - kbd/rotary knob issues has been completed. the console logs did not reveal any information relevant to the complaint. the failure of the rotary encoder being unresponsive to push input was reproducible upon boot up in the lab. the rotation input of the encoder was functional. visual inspection of the internal circuitry of the revealed no loose, disconnected, or broken cables. power supply from the pbm to the impellatronic board was found to be within specification. replacing only the rotary encoder with a known good one did not resolve failure. replacing only the impellatronic board with a known good one did not resolve failure. replacing only the 4 in 1 with a known good one did not resolve failure. replacing both the 4 in 1 and impellatronic board resolved failure. the root cause for the rotary encoder not working was a defective impellatronic board and defective epc( 4in1) . d9 date device returned to manufacturer added. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-10424. d2 common device name. d4 model number. f6/f8-should have been left blank.
27364 | 1220648-2024-10425 | 19190895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19190895&pc=OZD | 04/04/2022 | 04/26/2024 |  | 04/26/2024 |  |  | 04/04/2022 | 04/26/2024 |  |  | 06/30/2023 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2022095294 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support hemolysis was identified and the patient experienced limb ischemia. the impella was explanted. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
27365 | 1220648-2024-10427 | 19190923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19190923&pc=PYX | 04/17/2024 | 06/06/2024 |  | 04/26/2024 |  |  | 04/17/2024 | 04/26/2024 |  |  | 10/31/2025 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024413711 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2024 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had rp flex support ongoing for a patient admitted with rv (right ventricle) failure and effusion. after 7days of support there was an observation made of thrombosis within the ra (right atrium). | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. the device was not returned for evaluation. as the necessary information was not provided, the root cause of the thrombus was not determined. d.4 revised model number from the initial submission in accordance with updated procedures. h.6 added code 925 as it was inadvertently left off the previously submitted manufacture device report.
27366 | 1220648-2024-10426 | 19190924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19190924&pc=OZD | 04/04/2022 | 04/26/2024 |  | 04/26/2024 |  |  | 04/04/2022 | 04/26/2024 |  |  | 12/31/2023 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174832 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/04/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced access site bleeding. sutures were placed and heparin was stopped due to the oozing. the patient was given two units of packed red blood cells. it was also reported that an echocardiogram revealed a left ventricular outflow tract thrombus or vegetation. the physician decided to wait on a performing a trans-esophageal echocardiogram until the impella was explanted for better evaluation of the thrombus or vegetation. the patient was reportedly stable, remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
27367 | 1220648-2024-10428 | 19190928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19190928&pc=OZD | 04/05/2022 | 05/16/2025 |  | 04/26/2024 |  |  | 04/05/2022 | 04/26/2024 |  |  | 12/31/2023 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162605 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 30 YR |  | Female |  |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Vascular Problem | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2022 |  |  |  |  | The complainant reported a 30-year-old patient, of unknown race and ethnicity, in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient had complications with diffuse bleeding after the impella cp was inserted. it was also reported during impella cp support, the impella had alarms for purge block and high pressure and the purge cassette was leaking. the purge cassette was replaced and did not resolve. the medical team was advised to administer a tissue plasminogen activator, which resolved the high purge issues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torqueing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27368 | 1220648-2024-10429 | 19190964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19190964&pc=OZD | 04/18/2024 | 05/16/2024 |  | 04/26/2024 |  |  | 04/18/2024 | 04/26/2024 |  |  |  |  | 04/26/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1613085 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, a alarm stating "purge fluid not detected" occurred. the purge bag was replaced and the automated impella controller (aic) was swapped out with a new aic. the aic was returned for evaluation. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge issue ¿ other has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the purge disc retainer was found to be broken. the root cause of this issue was a broken purge disc retainer. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27369 | 1220648-2024-10430 | 19190972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19190972&pc=OZD | 04/06/2022 | 04/26/2024 |  | 04/26/2024 |  |  | 04/06/2022 | 04/26/2024 |  |  | 07/31/2023 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022106545 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/06/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the impella cp had low flow issues and the device was replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27370 | 1220648-2024-10431 | 19191034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19191034&pc=OZD | 04/19/2024 | 06/07/2024 |  | 04/26/2024 |  |  | 04/19/2024 | 04/26/2024 |  |  |  |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025482946 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella pump was supporting; however, it was explanted due to low flows from a kink of the catheter. according to physician, due to patient¿s short ascending aorta, a catheter kink occurred. the physician tried to remedy the kink by pulling the catheter back a little. however, the issue did not resolve. the procedure continued with impella pump support during the case. it was noted no further issues occurred. the impella pump was removed after three hours of support. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (cannula kink) issue has been completed since the original report was filed. the device was returned for investigation. cannula kink was observed near outflow cage. the cause of the low pump flow issue was cannula kink due to patient anatomy condition as the patient had short ascending aorta and kink occurred while attempted to pull the catheter back.
27371 | 1220648-2024-10432 | 19191122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19191122&pc=OZD | 04/06/2022 | 04/26/2024 |  | 04/26/2024 |  |  | 04/06/2022 | 04/26/2024 |  |  | 10/31/2023 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154753 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Injury | Yes | No | Insufficient Flow or Under Infusion | Obstruction of Flow | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/06/2022 |  |  |  |  | The user facility reported a 53-year-old patient needing mechanical circulatory support. while on support, the patient experienced no doppler dorsalis pedis pulse and prothrombin time pulse was weak per doppler signal. additionally, there was a sudden onset of constant suction alarms, flows would not go over 0.8; there was concern over potential clot in the cannula, especially with an activated clotting time of 120. later, the impella cp was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27372 | 1220648-2024-10433 | 19191595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19191595&pc=OZD | 04/06/2022 | 03/25/2025 |  | 04/26/2024 |  |  | 04/06/2022 | 04/26/2024 |  |  | 09/30/2023 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022143884 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/06/2022 |  |  |  |  | The user facility reported an attempt was made to implant the impella cp for mechanical circulatory support, the patient noted as high risk percutaneous coronary intervention. during initial placement, the impella could not be advanced past the distal abdominal aorta/ostial iliac stenosis. implantation was aborted due to the patient's anatomy. the physician proceeded to high-risk percutaneous coronary intervention without impella support, and no further issues were reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. h1 type of reportable event was incorrectly reported as serious injury in initial mfr 1220648-2024-10433 and has been corrected to product malfunction.
27373 | 1220648-2024-10434 | 19191642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19191642&pc=OZD | 04/19/2024 | 06/05/2024 |  | 04/26/2024 |  |  | 04/19/2024 | 04/26/2024 |  |  | 09/30/2025 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024400540 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported an 83-year-old male patient of unknown ethnicity noted as high risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. it was reported that impella implant was aborted due to unable to pass pump through patient¿s tortuous iliac anatomy. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue is determined to be patient condition because the pump was unable to pass through the vasculature due to the tortuous iliac anatomy. the instructions for use for the impella cp with smartassist system in the following sections states: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
27374 | 1220648-2024-10435 | 19191672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19191672&pc=OZD | 04/12/2024 | 06/04/2024 |  | 04/26/2024 |  |  | 04/12/2024 | 04/26/2024 |  |  |  |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485262 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2024 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was admitted with angina and stemi. the 14 french was occlusive and flow limiting and so the team placed distal perfusion. the support was for three days when then weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ the instructions for use referenced in the initial is for the impella cp with smart assist system in sections: general patient care considerations, entering cardiac output, and potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).
27375 | 1220648-2024-10436 | 19191748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19191748&pc=OZD | 04/20/2024 | 06/24/2024 |  | 04/26/2024 |  |  | 04/20/2024 | 04/26/2024 |  |  |  |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025459922 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. multiple complications were encountered during impella cp support. shortly following implant, suction was encountered on the pump. positioning was verified via echocardiogram, and the possibility of a thrombus was discussed with the site. additional attempts to reposition the pump were unsuccessful in restoring adequate flows, and the patient was transferred to a different site. during the transfer, the patient was switched to another automated impella controller, and the suction events resolved in the process. the staff was still made aware of concerns of a potential thrombus despite resolution of the suction event. roughly 40 minutes later, the registered nurse noted left sided weakness and a code stroke was called. two hours later, bleeding at the groin site was discovered and a massive blood transfusion was called. the patient coded once more 15 minutes later, and the decision was ultimately made to halt support. the patient passed away following the complications. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation into the reported issue has been completed since the original report was submitted. the device was not returned for investigation. the data logs confirmed the failure mode and clinical narrative. the logs showed after pump started, motor current spiked following low flow that triggered auto suction response and suction alarms. this event remained for 70 minutes, then resolved when pump was run on p-2. product was not returned for review. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial ingestion. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. as limited clinical information was provided, the true root cause of the stroke/cerebrovascular accident could not be determined. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ d.4 revised model number from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 codes 4440, 2954 and 4628 were reported incorrectly on the previously submitted report per updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
27376 | 1220648-2024-10437 | 19191809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19191809&pc=PYX | 03/29/2024 | 09/19/2025 |  | 04/26/2024 |  |  | 03/29/2024 | 04/26/2024 |  |  |  |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2025476510 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient post-op ct surgery was successfully implanted with an impella rp device for support. while on support, the access site mattress sutures were tightened to maintain hemostasis around repo sheath. | Additional information indicated that, during support, the pump experienced placement signal issues. |  |  |  |  | The impella device has been received from the customer, however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ | The investigation into the optical signal issue and the access site bleeding ¿ major issue has been completed since the original report was filed. the pump was returned for investigation. the device parameters were within normal specification. the patient was decompensating and expired despite best efforts. the root cause of the optical signal issue was most likely patient condition related as the patient was decompensating. the root cause of the access site bleeding major issue was most likely patient condition related. in accordance with updated procedures, section d6 has been revised from the initial submission. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs show that the optical sensor went out of range and triggered the placement signal not reliable alarm. the returned pump passed the laser continuity test and electrical testing. no biomaterial was observed. tthe root cause of the optical signal issue was most likely patient condition related as the patient was decompensating. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. b5 updated with additional information. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures from the initial manufacturer device report number 1220648-2024-10437. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-10437.
27377 | 1220648-2024-10438 | 19191854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19191854&pc=OZD | 04/22/2024 | 07/16/2024 |  | 04/26/2024 |  |  | 04/22/2024 | 04/26/2024 |  |  |  |  | 04/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1417971 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/22/2024 |  |  |  |  | No patient involvement: us complainant reported the automated impella controller (aic) had a controller error. | [5 empty] | The impella device was received from the customer on (b)(6) 2024 and therefore,¿an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) issue had an error. the aic was returned for evaluation and upon functional testing, the reported failure of controller error was confirmed. the console failed system self-check failure on boot-up -in the lab. visual inspection of the pbm revealed corroded a corroded copper. data logs were reviewed which revealed numerous pbm communication controller errors. the root cause for the numerous pbm communication controller error alarms was due to contamination of the pbm boards by leakage of electrolyte from super capacitors c69 and c70. d2-corrected common device name. d4-corrected model number. f6/f8 should have been left blank in the initial report.
27378 | 1220648-2024-10439 | 19191869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19191869&pc=OZD | 04/16/2024 | 05/30/2024 |  | 04/26/2024 |  |  | 04/16/2024 | 04/26/2024 |  |  |  |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450548 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/11/2024 |  |  |  |  | The user facility reported a 79-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that during explant of the pump, there were issues with achieving hemostasis. the perclose failed and bleeding was occurring. the team troubleshot by balloon and placement of a covered stent. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of access site bleeding was determined to be use issue because of the break in one of perclose and one was able to be tightened. later the bleeding was resolved, and the hemostasis was achieved by covered stent on the axillary artery and manual pressure on lfa.
27379 | 1220648-2024-10440 | 19191892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19191892&pc=OZD | 04/22/2024 | 05/15/2024 |  | 04/26/2024 |  |  | 04/22/2024 | 04/26/2024 |  |  |  |  | 04/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1656412 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/22/2024 |  |  |  |  | The complainant reported that during a purge cassette exchange, the automated impella controller (aic) was not recognizing the purge disc. the aic was exchanged. no patient harm has been reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. added- h6 impact code 4629 d4-updated model number.
27380 | 1220648-2024-10441 | 19191931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19191931&pc=OZD | 05/07/2022 | 04/26/2024 |  | 04/26/2024 |  |  | 05/07/2022 | 04/26/2024 |  |  | 10/31/2023 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152489 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/07/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient had narrow pulse pressure, the nurse was unable to locate pulses. records report that heparin was held until next partial thromboplastin time (ptt). it was later reported that the patient¿s groin had a small ooze, and the site was noted to be ecchymotic. it was noted that there were plans to change the dressing by end of shift. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27381 | 1220648-2024-10442 | 19191936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19191936&pc=OZD | 04/23/2024 | 06/12/2024 |  | 04/26/2024 |  |  | 04/23/2024 | 04/26/2024 |  |  |  |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443123 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Embolism/Embolus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/23/2024 |  |  |  |  | The complainant reported a 57 year-old male patient with postcardiotomy cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. the patient had been admitted for aortic valve regurgitation (avr) and aortic repair. the 5.5 was delivered via the 8 x10 gelweave graft at the axillary. after the avr the team observed bleeding and air. after the patient was weaned off the 5.5, the 5.5 was replaced by an intra-aortic balloon pump. the patient survived the event. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation for the reported embolism and access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the embolism and access site bleed could not be determined. the instructions for use referenced in the initial report is the impella 5.5 with smart assist in sections: general patient care considerations, entering cardiac output, potential adverse events (united states). added- h6 clinical code 4438 d4-corrected model number in accordance with updated procedures, sections d6 and g3 have been revised from the initial submission.
27382 | 1220648-2024-10443 | 19192048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19192048&pc=OZD | 04/28/2022 | 03/25/2025 |  | 04/26/2024 |  |  | 04/28/2022 | 04/26/2024 |  |  | 01/31/2024 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184929 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/28/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implantation, the physician was unable to pass the impella cp pump across the aortic valve due to aortic stenosis. the physician attempted to intervene through the left anterior descending, but there was no pulse, so the decision was to stop all intervention in the cath lab. the procedure was aborted due to the patient's anatomy. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27383 | 1220648-2024-10444 | 19192293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19192293&pc=OZD | 04/13/2022 | 04/26/2024 |  | 04/26/2024 |  |  | 04/13/2022 | 04/26/2024 |  |  | 10/31/2023 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023159527 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Restricted Flow rate | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/13/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, it was noted that the p-level could not be increased due to suction issue, patient felt uncomfortable while impella was advancing into lv (left ventricle) and oozing at impella site. the suture was redone by the physician and pressure was manually applied for around 30 minutes, still was unable to stop bleeding. later, femoralstop was applied to the wound. oozing was resolved. a transfusion was also provided, which improved impella flow momentarily; however, the effect eventually disappeared and was still unable to increase p-level further and the impella device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ this report is being filed as part of a retrospective review of historical records.
27384 | 1220648-2024-10445 | 19192391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19192391&pc=OZD | 04/29/2022 | 03/25/2025 |  | 04/26/2024 |  |  | 04/29/2022 | 04/26/2024 |  |  | 09/30/2023 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022146560 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump reportedly stopped working. the patient tolerated being off support, so the pump was explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27385 | 1220648-2024-10446 | 19192524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19192524&pc=OZD | 05/01/2022 | 03/25/2025 |  | 04/26/2024 |  |  | 05/01/2022 | 04/26/2024 |  |  | 10/31/2023 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156147 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/01/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient continued to declined and required thrombectomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27386 | 1220648-2024-10448 | 19193303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19193303&pc=OZD | 05/18/2022 | 04/26/2024 |  | 04/26/2024 |  |  | 05/18/2022 | 04/26/2024 |  |  | 01/31/2024 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191737 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/18/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had no pulses after impella cp removal, the patient was administered heparin and transferred to another department for treatment. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
27387 | 1220648-2024-10449 | 19193476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19193476&pc=OZD | 04/18/2024 | 06/05/2024 |  | 04/26/2024 |  |  | 04/18/2024 | 04/26/2024 |  |  |  |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485868 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported a 64-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp exhibited low pump flow. the pump was explanted and repositioned twice and found to be kinked at the aortic arch. implantation of the impella was aborted. the patient is stable post procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported low pump flow has been completed. the impella has been returned for evaluation. visual inspection found a kink on the cannula about 15 mm from outflow cage and cannula bonding site. no other issues were found on the rest of the pump. pump ran without issue under bench test. the clinical details confirm the cannula kinked. the root cause of low flow was kinked cannula.
27388 | 1220648-2024-10450 | 19193727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19193727&pc=OZD | 04/17/2024 | 06/10/2024 |  | 04/26/2024 |  |  | 04/19/2024 | 04/26/2024 |  |  | 11/30/2025 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435544 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported a 56-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support due to becoming hypotensive. while on support, the impella had low pump flows. blood volume was given and the pump was repositioned, but the pump was removed. when the pump was removed, no clots were seen on the pump and no kinks were seen as well. it was reported that the pump was malpositioned and this was why the low pump flows were observed. also, the patient did have multiple clots removed from the left anterior descending artery and the right axillary artery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported low pump flow, thrombus, and positioning issues has been completed. data logs confirmed the low flow when pump started about 10 minutes and remained to the rest of the case that triggered auto suction response and suction alarms. logs also showed about 6 minutes into the case, pump position was instable (in aorta & correct position) for 3 minutes then resolved. pump was returned for analysis. visual inspection found no issues on the pump. the root cause of the low flow was most likely a patient condition as no issue was found on the pump. the root cause of the positioning issue was most likely patient movement. added- d9 device return date d6.
27389 | 1220648-2024-10451 | 19194115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19194115&pc=OZD | 04/19/2024 | 06/07/2024 |  | 04/26/2024 |  |  | 04/19/2024 | 04/26/2024 |  |  | 11/30/2025 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423271 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was to be implanted with an impella cp device for mechanical circulatory support. impella cp delivery was aborted when the delivery failed due to patient anatomy. the sheath alone was placed. the percutaneous coronary intervention was accomplished without impella support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist for use during cardiogenic shock and high risk pci & section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related, impella implant was unsuccessful as it was unable to cross valve due to patient anatomy.
27390 | 1220648-2024-10452 | 19194273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19194273&pc=OZD | 04/13/2024 | 06/03/2024 |  | 04/26/2024 |  |  | 04/13/2024 | 04/26/2024 |  |  |  |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025469239 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2024 |  |  |  |  | The user facility reported a 69-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that upon arrival at the intensive care unit (icu), blood was noted from the impella access site. heparin was stopped and manual pressure held. one unit of packed red blood cells (prbc) was administered. bleeding continued overnight, an occlusive suture was placed in the morning and bleeding resolved. it was further reported that due to hemolysis and vascular concerns the impella was explanted and transitioned to intra-aortic balloon pump (iabp). the patient was reported to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the access site bleeding ¿ major and the hemolysis issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the access site bleeding and the hemolysis issues could not be determined since the complaint device not returned for investigation and insufficient clinical data provided. impella cp with smart assist system section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections g3 and h10 have been revised from the initial submission.
27391 | 1220648-2024-10453 | 19194385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19194385&pc=OZD | 04/12/2022 | 04/26/2024 |  | 04/26/2024 |  |  | 04/12/2022 | 04/26/2024 |  |  | 11/30/2023 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165798 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/12/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support, and while on impella support, the patient¿s urine showed signs of hemolysis after returning from the operating room. the device was repositioned to 3.7cm, with echocardiography used to verify position, levo and vaso drips weaned off. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
27392 | 1220648-2024-10454 | 19194407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19194407&pc=OZD | 04/17/2024 | 06/05/2024 |  | 04/26/2024 |  |  | 04/17/2024 | 04/26/2024 |  |  | 06/30/2025 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379256 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the cp limb suffered limb ischemia and was explanted after just 8 hours of support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. inspection: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use referenced in the initial report is for the impella cp with smartassist system.
27393 | 1220648-2024-10455 | 19194720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19194720&pc=OZD | 04/17/2024 | 05/03/2024 |  | 04/26/2024 |  |  | 04/17/2024 | 04/26/2024 |  |  | 11/30/2025 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024435544 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2024 |  |  |  |  | Additional information was received. it was reported that manufacturer device report # is a duplication of the same patient and event as reported in manufacturer device report 1220648-2024-10450. | The user facility reported a 56-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and after the implant, there was low pump flow. it was reported that the pump was suctioning. the suction continued despite adjusting the correct position and giving fluids. the clinical representative confirmed the correct pump position; however the pump was explanted. |  |  |  |  | B.5 was revised. any new information will be reported with manufacturer device report 1220648-2024-10450 going forward. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
27394 | 1220648-2024-10456 | 19194996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19194996&pc=OZD | 03/29/2024 | 05/03/2024 |  | 04/26/2024 |  |  | 03/29/2024 | 04/26/2024 |  |  |  |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450604 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/29/2024 |  |  |  |  | The complainant reported an 86 year-old male patient with a high risk of percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the impella device was explanted due to access site bleeding after the peel away removal. the physician performed a cutdown and obtained control of the vessel successfully. the patient later expired while under support. per medical safety review, we do not have enough information to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the access site bleeding was most likely due to anticoagulation management since the act at the time bleed was 300 which is above the limit range. correction in d6a, d6b. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
27395 | 1220648-2024-10457 | 19195114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19195114&pc=OZD | 01/24/2024 | 05/30/2024 |  | 04/26/2024 |  |  | 04/19/2024 | 04/26/2024 |  |  | 06/30/2025 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-CANADA |  |  | ABIOMED, INC. |  |  | US |  | 2024378531 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 68 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported from a study, the patient undergoing high risk percutaneous coronary intervention had an impella cp placed for mechanical circulatory support. the support was successful for under 3 hours and the impella pump was explanted per protocol. however, it was further reported under the study, the patient experienced the following adverse events: cardiogenic shock with hypotension and acute kidney injury (aki). hypotension immediately occurred after placement of the impella sheath (contributing to the hypotension in the setting of acute myocardial infarction left anterior ascending artery). the aki occurred within 24 hours of index procedure and was noted to be likely secondary to hypotension occurring during the index procedure, which was exacerbated by the impella pump placement. the relatedness for both hypotension and aki events was noted as "possible" to the impella cp device procedure. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed and renal failure has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding and the renal failure issue was not determined since product was not returned and much information is provided as per the clinical description provided. the instructions for use reported in the initial report is from ifu for impella cp with smartassist for use during cardiogenic shock and high risk cpi. d4-updated model number.
27396 | 1220648-2024-10458 | 19195405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19195405&pc=OZD | 04/19/2024 | 06/05/2024 |  | 04/26/2024 |  |  | 04/19/2024 | 04/26/2024 |  |  | 05/31/2025 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352059 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support, and after one day of support, a thrombus was found in the left ventricle. as the patient's hemodynamics had improved, the decision was made to wean and explant the impella pump, which was successfully completed. the patient remained hemodynamically stable after removal of the impella pump. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. the device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. instructions for use for the related event are as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." d.4 revised model number from the initial submission in accordance with updated procedures. g.4. revised pma/510(k) number as previously entered incorrectly. h.6 added code 925 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
27397 | 1220648-2024-10459 | 19195526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19195526&pc=PYX | 04/20/2024 | 06/10/2024 |  | 04/26/2024 |  |  | 04/20/2024 | 04/26/2024 |  |  | 06/30/2024 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023229794 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2024 |  |  |  |  | The user facility reported a 70-year-old male patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella rp device for mechanical circulatory support. it was reported that multiple attempts were made to pass the impella rp into the left pulmonary artery, but it would not pass because of the patient tortuous vessel. the case was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella rp smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue is determined to be patient condition because rv was extremely bowed and rpsa would not appropriately conform to anatomy.
27398 | 1220648-2024-10461 | 19195646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19195646&pc=OZD | 04/21/2024 | 06/11/2024 |  | 04/26/2024 |  |  | 04/21/2024 | 04/26/2024 |  |  |  |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454401 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/21/2024 |  |  |  |  | The user facility reported a 49-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the physician attempted to reposition the impella 5.5 and pulled the impella across the valve into the aorta. the patient was taken to the operation room for impella replacement. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist in section: use of echocardiography for positioning of the impella catheter. d4-updated model number. d6-implant and explant date.
27399 | 1220648-2024-10462 | 19195648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19195648&pc=OZD | 04/22/2024 | 06/26/2024 |  | 04/26/2024 |  |  | 04/22/2024 | 04/26/2024 |  |  | 11/30/2025 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434305 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported upon implanting the impella 5.5 device, the guidewire became stuck. the impella 5.5 device was successfully placed. however, when attempting to remove the guidewire, it became stuck inside the impella device. the clinical representative confirmed after looking at the guidewire and seeing that the distal end of the wire was not damaged and intact, the mid portion of the guidewire was kinked/ folded over on itself. the decision was made to remove the impella 5.5. after multiple attempts with transesophageal echocardiogram guidance, the impella 5.5 device and the guidewire were removed as a single unit. there was no report of patient harm. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (guidewire damage - great vessel) has been completed since the original report was filed. the pump and guide wire were returned for investigation. a cannula kink and catheter kink were found on the returned product. multiple bends were noted on the guide wire. as per the clinical details, the impella was successfully inserted into the ventricle, but the guide wire became stuck; after multiple transesophageal echocardiogram-guided removal attempts. the impella and wire were removed together as a single unit. the cause of the guide wire damage issue was use related. d.4 revised model number from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. | H6 code was inadvertently omitted from the previously submitted report.
27400 | 1220648-2024-10463 | 19195809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19195809&pc=PYX | 04/22/2024 | 06/06/2024 |  | 04/26/2024 |  |  | 04/22/2024 | 04/26/2024 |  |  | 05/31/2025 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024355148 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/22/2024 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella rp support and they had access site bleeding. after the impella insertion, an extra mattress suture was applied to achieve hemostasis. there was also some oozing reported when the patient was rolling, but it resolved. there was also some pain at the insertion site. furthermore, when the staff tried to change the purge cassette, it was unable to be clicked into the console. a second purge cassette was used successfully. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported issues has been completed since the original report was filed. the purge cassette was returned. the returned complaint purge cassette was visually inspected. no problem found during priming process and flow was in standard range. the cause of access site bleeding- major cannot be determined since the complaint device not returned for investigation and insufficient clinical data provided.
27401 | 1220648-2024-10464 | 19195938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19195938&pc=OZD | 04/23/2024 | 07/09/2024 |  | 04/26/2024 |  |  | 04/23/2024 | 04/26/2024 |  |  |  |  | 04/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1300881 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller was turning on and off by itself while not in use. there was no patient involvement. | [5 empty] | The automated impella controller (aic) has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - controller error (yellow alarm) has been completed since the initial report was submitted. the console was returned, visually inspected, tested, and reproduced the reported failure mode. upon connecting the ac power, the console booted up by itself without pushing the power button. the front panel optical connector was found contaminated as received and unable to remove the debris by cleaning and was unrelated to the reported failure mode. the root cause of the automatic bootup and restart was pbm malfunction. h4 model number d9 corrected e2.
27402 | 1220648-2024-10465 | 19195946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19195946&pc=OZD | 04/22/2024 | 06/11/2024 |  | 04/26/2024 |  |  | 04/22/2024 | 04/26/2024 |  |  | 10/31/2025 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422209 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/22/2024 |  |  |  |  | The user facility reported a 74 year-old male patient with myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the patient experienced ventricular tachycardia which resolved on its own. the patient remains on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the ventricular fibrillation has been completed. the console was not returned for evaluation. without additional information, the root cause of the issue was not determined because of insufficient clinical details and the product was not returned for investigation. the instructions for use for the impella cp with smart assist system under secondary safety results states one of the 10 major adverse event that occur at any time point with repeat revascularization, which includes cpr or ventricular arrhythmia requiring cardioversion.
27403 | 1220648-2024-10460 | 19196778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19196778&pc=OZD | 04/20/2024 | 06/24/2024 |  | 04/26/2024 |  |  | 04/20/2024 | 04/26/2024 |  |  | 07/31/2025 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392443 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Life Threatening |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced multiple episodes of ventricular tachycardia (vt) during impella cp support. as a result of bleeding, complications at the extracorporeal membrane oxygenation (ecmo) insertion site, impella flows began to drop, and the patient became hypotensive. the left ventricle was believed to be significantly underfilled, and the patient went into multiple episodes of vt requiring shocks. it was stated that the impella pump was thought to be the culprit of the vt, and the pump was subsequently removed. the vt resolved following explant. | [5 empty] | The impella device was not received from the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the ventricular tachycardia/ventricular fibrillation was not determined. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.6 code 4641 was reported incorrectly on the previously submitted report.
27404 | 1220648-2024-10466 | 19197730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19197730&pc=OZD | 05/11/2023 | 05/15/2024 |  | 04/26/2024 |  |  | 03/28/2024 | 04/26/2024 |  |  | 12/31/2024 |  | 04/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024305886 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/28/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. a notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured ventricular tachycardia with a "possible" relationship to the impella device used. a cardioversion was required. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, vascular injury, venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. d4-updated model number added- d6a/d6b
27405 | 1220648-2024-10382 | 19180147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19180147&pc=OZD | 04/30/2022 | 03/26/2025 |  | 04/25/2024 |  |  | 04/30/2022 | 04/25/2024 |  |  | 02/29/2024 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196735 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Sepsis | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/30/2022 |  |  |  |  | The user facility reported a patient in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced signs of sepsis and hemolysis. the patient was treated for sepsis, but no information was provided on the treatment performed. echocardiography was performed to verify impella positioning and the medical team was informed the patient could be escalated to an impella 5.5 to assist in relieving hemolysis in the patient. it is unknown how or if the hemolysis was resolved. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records.
27406 | 1220648-2024-10383 | 19180443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19180443&pc=OZD | 05/09/2022 | 04/25/2024 |  | 04/25/2024 |  |  | 05/09/2022 | 04/25/2024 |  |  | 12/31/2023 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170674 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/09/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced a hematoma. the touhy was tightened, blue hub was sutured, and the impella access site was dressed appropriately, and the hematoma was resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
27407 | 1220648-2024-10384 | 19180550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19180550&pc=OZD | 05/09/2022 | 04/25/2024 |  | 04/25/2024 |  |  | 05/09/2022 | 04/25/2024 |  |  | 11/30/2022 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021017428 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/09/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a left ventricle (lv) thrombus was observed at the patient¿s bedside while utilizing echocardiogram (echo) for verifying impella cp placement. the decision was made to remove the impella cp. it is reported that the patient was weaned successfully, the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27408 | 1220648-2024-10385 | 19181214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19181214&pc=OZD | 06/29/2022 | 03/26/2025 |  | 04/25/2024 |  |  | 06/30/2022 | 04/25/2024 |  |  | 08/31/2022 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1500618 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Vascular Dissection | Thrombosis/Thrombus | [10 empty] | Required Intervention | [5 empty] | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2022 |  |  |  |  | The user facility reported a 57-year-old patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device and venoarterial extracorporeal membrane oxygenation for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced a hematoma at the impella access site. it was reported that the hematoma worsened and the patient was taken to the operating room where transesophageal echocardiography was performed and a left ventricle (lv) thrombus and an aortic dissection (extends to abdominal aorta) was visualized. the impella cp was removed in the operating room and a septostomy was completed. the patient continued to be supported on venoarterial extracorporeal membrane oxygenation. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. b3 date of event was incorrectly reported as 30-jun-2022 on initial mfr 1220648-2024-10385 and has been corrected to 29-jun-2022. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿. this report is being filed as part of a retrospective review of historical records.
27409 | 1220648-2024-10386 | 19181414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19181414&pc=OZD | 01/20/2022 | 03/26/2025 |  | 04/25/2024 |  |  | 01/20/2022 | 04/25/2024 |  |  | 10/31/2023 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154768 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/20/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, patient developed a hematoma at the impella cp access site. manual pressure was applied and four units of packed red blood cells were transfused. the patient was weaned and impella was removed with out issue, but the two perclose devices failed when removing the sheath. a tamponade with balloon was placed with a covered stent. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. this report is being filed as part of a retrospective review of historical records.
27410 | 1220648-2024-10387 | 19181482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19181482&pc=OZD | 04/10/2022 | 09/30/2025 |  | 04/25/2024 |  |  | 04/10/2022 | 04/25/2024 |  |  | 12/31/2022 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022023423 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, there was a low placement signal alert and the impella was repositioned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-10387 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-10387in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-10387. h6 medical device problem code 2906 was erroneously entered on the initial manufacturer device report number 1220648-2024-10387.
27411 | 1220648-2024-10388 | 19181841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19181841&pc=OZD | 04/11/2022 | 04/25/2024 |  | 04/25/2024 |  |  | 04/11/2022 | 04/25/2024 |  |  | 11/30/2023 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163743 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/11/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support due to shortness of breath, lethargy, and heart failure cardiogenic shock. while on support, the patient experienced excessive groin oozing. groin oozing was controlled by placing suture to tamponade. pressure was held with femstop to prevent further bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27412 | 1220648-2024-10389 | 19181946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19181946&pc=OZD | 04/10/2022 | 04/25/2024 |  | 04/25/2024 |  |  | 04/10/2022 | 04/25/2024 |  |  | 11/30/2023 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165048 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/10/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp was continuously bleeding. manual pressure was held, and 3 units of blood products were administered to the patient. however, the impella was explanted and replaced with another impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27413 | 1220648-2024-10390 | 19182050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19182050&pc=OZD | 04/02/2022 | 09/26/2025 |  | 04/25/2024 |  |  | 04/02/2022 | 04/25/2024 |  |  | 12/31/2023 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177636 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2022 |  |  |  |  | The user facility reported a 73-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis. as a result, impella was repositioned with echocardiography. the impella cp was weaned and explanted prior to hemolysis clearing. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27414 | 1220648-2024-10391 | 19182183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19182183&pc=OZD | 04/08/2022 | 03/26/2025 |  | 04/25/2024 |  |  | 04/08/2022 | 04/25/2024 |  |  | 11/30/2023 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167033 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was bleeding at the site. the nurses tried angle matching, manual pressure and femstop, however the bleeding was not stopping. therefore, the impella was removed, and surgical closure was performed. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27415 | 1220648-2024-10392 | 19182445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19182445&pc=OZD | 04/03/2022 | 04/25/2024 |  | 04/25/2024 |  |  | 04/03/2022 | 04/25/2024 |  |  | 08/31/2023 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022134181 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/03/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced bleeding at the impella access site. it was reported that the patient had a very strong and frequent cough which caused him to bear down and increase bleeding at the site. the following was performed to stop the bleeding: manual pressure, an external patch device was attempted, heavy forward tension sutures, the blue portion was removed from the skin, angle matching, lidocaine/epinephrine injection; femstop, and sutured dressing, additionally, heparin was withheld. the bleeding stopped and no additional issues were reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27416 | 1220648-2024-10393 | 19182690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19182690&pc=OZD | 04/11/2022 | 04/25/2024 |  | 04/25/2024 |  |  | 04/11/2022 | 04/25/2024 |  |  | 01/31/2024 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183747 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/11/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. post procedure the patient had pedal pulses absent via doppler and angiogram revealed that the impella peel away sheath was occlusive. they did have distal limb perfusion post implant and was stented, temp pacer sutured in place. the impella was weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27417 | 1220648-2024-10394 | 19182701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19182701&pc=OZD | 04/04/2022 | 03/26/2025 |  | 04/25/2024 |  |  | 04/04/2022 | 04/25/2024 |  |  | 11/30/2023 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165805 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2022 |  |  |  |  | The user facility reported a 65-year-old patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that after a successful percutaneous coronary intervention the impella cp was weaned and explanted in the procedural area and the manta closure was unsuccessful with persistent arterial bleed post deployment. 30mg of protamine was administered and manual pressure held for 10 minutes. hemostasis achieved, site soft, and distal pulses dopplerable at the conclusion of the case. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
27418 | 1220648-2024-10395 | 19182738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19182738&pc=OZD | 03/11/2024 | 06/04/2024 |  | 04/25/2024 |  |  | 03/26/2024 | 04/25/2024 |  |  | 09/30/2025 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403520 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/26/2024 |  |  |  |  | Additional information via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured anemia with a " probable" relationship to the impella device used: ¿baseline hgb 14.0. post pci hgb 7.9. nadir hgb 13mar2024 6.4. 1 unit prbcs transfused. ferrlecit 125mg x1 dose given. bleeding source not identified. hematology note- more suggestive of consumption in setting of blood loss; likely iron deficiency & anemia of chronic disease. hgb at d/c 8.4.¿ the patient recovered with no sequelae. | A notification was received via long-term outcome and quality indicator impella registry regarding a patient that developed right limb ischemia during impella cp support. it was noted that the lower extremity was cold, and the patient complained of pain and numbness in the right leg. to manage the condition, an internal bypass catheter was placed in the left femoral arterial sheath down into the right popliteal. doppler confirmed pulses were present following the intervention. |  |  |  |  | B5 and h6 has been updated as per additional information received. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system: section: potential adverse events (united states) "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia and access site bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and access site bleed could not be determined as the device was not returned nor sufficient clinical details.
27419 | 1220648-2024-10396 | 19182964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19182964&pc=OZD | 04/13/2022 | 04/25/2024 |  | 04/25/2024 |  |  | 04/13/2022 | 04/25/2024 |  |  | 12/31/2023 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176409 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/13/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support the patient experienced ventricular tachycardia. the nurse verified that the patient was stable. the pump was weaned and explanted the next day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27420 | 1220648-2024-10397 | 19183027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19183027&pc=OZD | 04/14/2022 | 03/26/2025 |  | 04/25/2024 |  |  | 04/14/2022 | 04/25/2024 |  |  | 01/31/2024 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188056 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2022 |  |  |  |  | The user facility reported a 70-year-old male on impella cp for mechanical circulatory support. the impella cp was weaned and explanted with perclose. it was reported that there was still a fair amount of bleeding; therefore, a third perclose was deployed. the outcome of the reported event was satisfactory to the physician. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
27421 | 1220648-2024-10398 | 19183046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19183046&pc=OZD | 04/10/2022 | 04/25/2024 |  | 04/25/2024 |  |  | 04/11/2022 | 04/25/2024 |  |  | 04/30/2023 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022080453 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/11/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support and that there was a low pump flow as they encountered continuous suctioning and were unable to increase the p support level of the pump. this was treated by giving patient albumin, and the pump was explanted as a result. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27422 | 1220648-2024-10399 | 19183176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19183176&pc=OZD | 04/14/2022 | 03/26/2025 |  | 04/25/2024 |  |  | 04/14/2022 | 04/25/2024 |  |  | 11/30/2023 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166576 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2022 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a hematoma. manual pressure was applied. the hematoma visually increased. the patient received four units of rapid infused blood and the outcome of the patient was reported as recovered. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
27423 | 1220648-2024-10400 | 19183244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19183244&pc=OZD | 04/15/2022 | 03/25/2025 |  | 04/25/2024 |  |  | 04/15/2022 | 04/25/2024 |  |  | 01/31/2024 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189535 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2022 |  |  |  |  | The user facility reported a 63-year old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis. the pump was repositioned, and the outcome of the event was reported as resolved. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
27424 | 1220648-2024-10401 | 19183269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19183269&pc=OZD | 04/18/2024 | 05/30/2024 |  | 04/25/2024 |  |  | 04/18/2024 | 04/25/2024 |  |  |  |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025449006 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed via the femoral and supported for 37 hours when the team made the decision to explant the impella cp and placed a higher support impella 5.5 pump. the pump was pulled back and the patient arrested. the ventricular fibrillation arrest was treated by 12 shocks/defibrillations. the team then placed an impella 5.5 to treat surgically a groin site hematoma and gave two units of blood. the patient is now on impella 5.5 support and has survived the event. | [5 empty] | The investigation for the reported access site bleed and ventricular fibrillation has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and ventricular fibrillation could not be determined. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
27425 | 1220648-2024-10402 | 19183310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19183310&pc=OZD | 04/15/2022 | 05/30/2025 |  | 04/25/2024 |  |  | 04/15/2022 | 04/25/2024 |  |  | 09/30/2023 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022149877 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2022 |  |  |  |  | The user facility reported a 25-year-old male patient noted with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. impella had low purge flow. heparin was added to the purge solution. no further information was available. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27426 | 1220648-2024-10403 | 19183369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19183369&pc=OZD | 04/15/2022 | 03/25/2025 |  | 04/25/2024 |  |  | 04/15/2022 | 04/25/2024 |  |  | 10/31/2023 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022155589 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/15/2022 |  |  |  |  | The user facility reported a 76-year-old male patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, limb ischemia developed in the patient as there was no pulse found in the right leg. the patient's leg became discolored, started to become mottled and cool to the touch. as a result, the pump was explanted. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27427 | 1220648-2024-10404 | 19183400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19183400&pc=OZD | 04/12/2022 | 03/25/2025 |  | 04/25/2024 |  |  | 04/12/2022 | 04/25/2024 |  |  | 01/31/2024 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188299 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was some bleeding at the site but was resolved by manual pressure and reinforcement of forward tension suture. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27428 | 1220648-2024-10405 | 19183532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19183532&pc=OZD | 04/20/2024 | 07/31/2024 |  | 04/25/2024 |  |  | 04/20/2024 | 04/25/2024 |  |  |  |  | 04/25/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1803713 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/20/2024 |  |  |  |  | The user facility reported a 68-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that during initial insertion set up, the automated impella controller (aic) displayed purge disc not detected despite reinsertion of the purge disc. there was no audible or tactile ¿click¿ during first and subsequent insertions and the purge disc felt loose in the housing, so a cassette change was performed which did not resolve the issue. the aic was exchanged and the issue resolved. there was no harm to the patient. | [5 empty] | The investigation for aic - purge disc/flag issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-10405: | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27429 | 1220648-2024-10406 | 19183640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19183640&pc=OZD | 04/24/2022 | 03/25/2025 |  | 04/25/2024 |  |  | 04/24/2022 | 04/25/2024 |  |  | 01/31/2024 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178349 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 29 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/24/2022 |  |  |  |  | The user facility reported a 29-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. impella cp had delivery issues. the impella was inserted into the patient and the position was confirmed. the impella cp pump was unable to pass through vasculature. procedure was aborted. no additional information was provided. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27430 | 1220648-2024-10407 | 19183859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19183859&pc=OZD | 04/21/2024 | 12/03/2025 |  | 04/25/2024 |  |  | 04/22/2024 | 04/25/2024 |  |  |  |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1193365 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 04/22/2024 |  |  |  |  | The user facility reported an 88-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the alarms on the automated impella controller (aic) were visual with no audio. the aic was swapped out and support continued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27431 | 1220648-2024-10408 | 19184751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19184751&pc=OZD | 04/18/2024 | 06/10/2024 |  | 04/25/2024 |  |  | 04/18/2024 | 04/25/2024 |  |  | 11/30/2025 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024433076 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Failure to Deliver | Difficult to Advance | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella cp insertion was aborted due to the patient¿s anatomy and a hematoma occurred at the impella site. the patient is stable. | [5 empty] | The investigation into the access site bleeding ¿ major and the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. per the clinical information provided, the root cause of the access site bleeding issue was not determined since product was not returned and sufficient information is not provided in the clinical description. the root cause of the delivery issue was most likely patient condition related; cp insertion was aborted due to the patient¿s anatomy as per the clinical description. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
27432 | 1220648-2024-10409 | 19184755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19184755&pc=OZD | 04/08/2022 | 03/25/2025 |  | 04/25/2024 |  |  | 04/08/2022 | 04/25/2024 |  |  | 01/31/2024 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186323 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/08/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the physician thought the impella might be too deep in the ventricle but didn¿t want to reposition due to concerns of causing ectopy which was problematic during the case. bedside echocardiography in intensive care unit confirmed the impella was position was too deep. the physician noted there no signs of hemolysis or other patient complications, so they decided to not reposition the impella. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. h1 type of reportable event was incorrectly reported as serious injury in initial mfr (b)(4) and has been corrected to malfunction. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27433 | 1220648-2024-10410 | 19184833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19184833&pc=OZD | 04/20/2024 | 06/11/2024 |  | 04/25/2024 |  |  | 04/20/2024 | 04/25/2024 |  |  |  |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025485154 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp was exchanged due to bleeding and ischemia. there was excessive bleeding around the repositioning sheath after removal of the peel-away sheath. a surgical cut down to repair the right femoral artery was required. the patient is stable. | [5 empty] | The investigation into the access site bleeding ¿ major and the limb ischemia ¿ reversible issues has been completed since the original report was submitted. the device was not returned for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the root cause of the vascular damage issue was not determined since product was not returned and sufficient information is not provided in the clinical description. impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
27434 | 1220648-2024-10411 | 19184885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19184885&pc=OZD | 04/22/2024 | 06/28/2024 |  | 04/25/2024 |  |  | 04/22/2024 | 04/25/2024 |  |  |  |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025449995 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/22/2024 |  |  |  |  | It was further reported that that care was withdrawn from the patient and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate use of the device with the outcome. | The user facility reported a 67-year-old male patient of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while delivering support to a patient with an impella 5.5 the sterile sleeve pulled out of the suture hub. the issue was rectified with tegaderm to reinforce the suture. there was no harm to the patient. |  |  |  |  | The investigation for the reported product damage has been completed. the device was returned for evaluation. the device was inspected and it was confirmed that the sterile sleeve was completely detached from the cath anchor side. the root cause of the product damage sterile sleeve damage was most likely use related as it was reported to have been torn by the cath grip during use. the instructions for use for the impella cp with smartassist system state in section: warnings & cautions: cautions the following: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ added-b5 mso review, d6b, d9 device return date f6/f8-should have been left blank in the initial report. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27435 | 1220648-2024-10412 | 19184900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19184900&pc=OZD | 04/14/2024 | 06/04/2024 |  | 04/25/2024 |  |  | 04/14/2024 | 04/25/2024 |  |  | 11/30/2025 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430037 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2024 |  |  |  |  | The user facility reported an 85-year-old male patient of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a patient's labs came back hemolyzed during impella cp support. it was additionally noted that the patient's urine had become pink / red. in response to the noted condition, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported hemolysis has been completed. the impella device was not returned for evaluation. data logs were reviewed which showed there were no motor current spikes, no suction alarms, nor positioning alarms to indicate hemolysis issue. the root cause of hemolysis test cannot be determined since the complaint device not returned for investigation and insufficient clinical data provided.
27436 | 1220648-2024-10413 | 19185021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19185021&pc=OZD | 04/22/2024 | 06/10/2024 |  | 04/25/2024 |  |  | 04/22/2024 | 04/25/2024 |  |  |  |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  |  |  | 2025466976 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/22/2024 |  |  |  |  | The user facility reported a 43 year-old male patient with cardiomyopathy was implanted with an impella device for mechanical circulatory support. the patient had been waiting for a left ventricular assist device placement. after 18 days of support, the impella device¿s pump stopped. medical staff attempted to troubleshoot by aic replacement, but instead determined the pump itself had to be replaced. a new impella device was implanted. medical staff alleged that the impella device¿s pump's cord was under the bed wheel which may have caused the pump stoppage. the patient survived the event. | [5 empty] | The investigation for the reported pump stop has been completed. the pump was returned for investigation. a kink was observed at the kink protector on the catheter side. no other physical abnormalities were observed. pump was turning on during the test run. the pump passed optical light testing but failed electrical testing on all three motor lines. further investigation revealed that all three motor lines were found stretched and fractured inside the red handle. according to the clinical details, the pump was tried to be restarted multiple times without success. additionally, it was reported that either the bed or aic cart wheel had been on top of the white impella 5.5 power cable. data logs show the pump ran for 440 hours and reported a pump stop at the end of the case run with impella stopped alarm 115/119. the cause of the pump stop issue was an open electrical line due to excessive force. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock, sections: using the automated impella controller with the impella catheter, using the automated impella controller with the impella rp with smartassist system catheter, and impella stopped. added- d9 device return date d4-updated model number added- d6a/d6b | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
27437 | 1220648-2024-10416 | 19187899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19187899&pc=OZD | 04/19/2024 | 06/05/2024 |  | 04/25/2024 |  |  | 04/19/2024 | 04/25/2024 |  |  |  |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486680 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Material Twisted/Bent | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2024 |  |  |  |  | The user facility reported a 53-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had cp support being placed for a patient admitted in for a hrpci. the team placed perfusion sheath due to known pad and impella in the limb. the pump delivery was complicated, but the pump was inserted. the insertion caused the pump to kink and be replaced. a 2nd cp was placed, and support continues to date. | [5 empty] | The investigation for the reported limb ischemia and product damage has been completed. the impella was not returned for evaluation. the clinical details was sufficient to determine the root cause. the cause of the product damage (cannula kink) most likely was use related due to improper technique. the cause of the limb ischemia has not been determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. impella cp with smartasssistsmartassist system section: potential adverse events (united states) ¿section: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
27438 | 1220648-2024-10417 | 19188115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19188115&pc=OZD | 04/21/2024 | 06/11/2024 |  | 04/25/2024 |  |  | 04/21/2024 | 04/25/2024 |  |  | 07/31/2025 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025441139 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Improper Flow or Infusion | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/21/2024 |  |  |  |  | The user facility reported a 43-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella cp was placed successfully but during the pci (percutaneous coronary intervention) procedure, there was concern with pump saturation due to aortic, left ventricle waveforms, motor current, and flows. after pci procedure was completed the impella cp was exchanged. the patient is stable. | [5 empty] | The investigation into the saturated flow has been completed since the original report was filed. the pump and data logs were returned. data logs confirm saturated flow; however, no biomaterial or blood ingress was found on the returned product. the root cause of the saturated flow issue was not determined since the issues was unable to be reproduced. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
27439 | 1220648-2024-10418 | 19188563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19188563&pc=OZD | 06/04/2023 | 10/16/2024 |  | 04/25/2024 |  |  | 03/26/2024 | 04/25/2024 |  |  | 03/31/2025 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024342503 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Life Threatening | Required Intervention | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | The patient was escalated to an impella 5.5 on the 9th day of impella cp support. it was reported that the patient procedure outcome while supported on the impella 5.5 was expired. medical safety review of the event noted use of the impella cp device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported a female patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported successful placement of impella cp, patient then decompensated requiring epi gtt. pt further decompensated to vt and shocked x 2 to st. pt was non pulsatile with only impella flow 3.5 lpm. decision was made to increase to va ecmo concurrent with impella. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2024-10418 was provided in sections b3 date of awareness was corrected, b5 additional information was provided, d6a date of implant was added, d6b date of explant was added, h1, and h6 codes were all corrected. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. product was not returned for review. the root cause of the pc leak was unable to be determined as limited clinical details were provided and no product was returned for review. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27440 | 1220648-2024-10419 | 19188570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19188570&pc=OZD | 06/14/2023 | 10/19/2024 |  | 04/25/2024 |  |  | 03/26/2024 | 04/25/2024 |  |  | 03/31/2025 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024342501 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Required Intervention | Other | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2024 |  |  |  |  | We received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), indicating an allegation of biventricular failure and rle ischemia with the impella cp device used on this patient. | The decision was made to withdraw care and the patient expired. per medical safety officer review, there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not returned by the customer and therefore, investigation of device was not possible. the product was released with a completed and reviewed device history record under the abiomed quality system. the manufacturer is closing the file on this event. | The investigation has been completed. the impella device was discarded. as the necessary clinical information was not provided and the device was not returned, the root cause of the reported issues was not determined. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.2 revised selection since the initial manufacturer device report number 1220648-2024-10419 was submitted due to incorrect previous selection and new information received from the medical safety officer review. b.3 revised date of event as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-10419. b.5 additional information was received, and abiomed's medical safety officer review was completed since the initial manufacturer device report number 1220648-2024-10419 was submitted. b.7 revised as information was inadvertently omitted from initial manufacturer device report number 1220648-2024-10419. d.3 revised manufacturer name as it was inadvertently omitted from initial manufacturer device report number 1220648-2024-10419. d.4 revised model number as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-10419. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2024-10419 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-10419 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2024-10419 was submitted. h.1 revised due to medical safety officer review since the initial manufacturer device report number 1220648-2024-10419 was submitted h.6 code 925 added since the initial manufacturer device report number 1220648-2024-10419 was submitted in accordance with updated procedures. h.6 health effect - impact code was revised due since the initial manufacturer device report number 1220648-2024-10419 was submitted due to medical safety officer review. h.6 codes 4112, 4114 and 4315 were previously entered incorrectly on the initial manufacturer device report number 1220648-2024-10419. new codes were added. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-10419. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded.
27441 | 1220648-2024-10420 | 19188729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19188729&pc=OZD | 02/01/2022 | 03/25/2025 |  | 04/25/2024 |  |  | 02/01/2022 | 04/25/2024 |  |  | 10/31/2023 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022151867 |  | 0048-00452 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2022 |  |  |  |  | The complainant reported a 68 year old female patient was implanted with an impella cp for mechanical circulatory support. while on support the patient experienced gi bleeding. heparin was withheld and the patient received 1 unit of blood. afterwards, the gi bleed was still apparent and in the patient¿s stool. heparin continued to be withheld and additional blood was given to the patient. manual pressure was held at bedside. the patient was successfully weaned off impella support. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27442 | 1220648-2024-10421 | 19188774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19188774&pc=OZD | 02/07/2022 | 03/25/2025 |  | 04/25/2024 |  |  | 02/07/2022 | 04/25/2024 |  |  | 09/30/2023 |  | 04/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022143054 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced ventricular fibrillation that was resolved with shock. additionally, the patient experienced access site bleeding that was resolved with balloon occlusion and perclose suturing. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27443 | 1220648-2024-10326 | 19171756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19171756&pc=OZD | 04/21/2022 | 05/20/2025 |  | 04/24/2024 |  |  | 04/21/2022 | 04/24/2024 |  |  | 09/30/2023 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022147380 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Dislodged or Dislocated | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 04/21/2022 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during transport of the patient to the intensive care unit (icu), aorta alarms were observed. an echocardiogram (echo) was performed, and it was identified that the impella device had pulled out of the ventricle. the device was successfully repositioned utilizing echo. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | D6: explant date unknown this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27444 | 1220648-2024-10327 | 19171882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19171882&pc=OZD | 05/19/2022 | 03/26/2025 |  | 04/24/2024 |  |  | 04/10/2024 | 04/24/2024 |  |  | 12/31/2023 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180821 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. there was some bleeding noted around the hemostatic valve after the impella was inserted. this was resolved by repositioning the catheter. no further bleeding was noted, and the patient remained on impella support. | It was additionally reported, that during impella insertion, the left femoral artery was punctured and access was obtained with ultrasound guidance. it was the determined it was too low near the bifurcation after imaging was reviewed and the impella was unable to advance through the vasculature. the left femoral access was abandoned, and the right femoral artery access was obtained with ultrasound guided puncture. the impella was then successfully inserted. |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. b1 product problem was not selected and should have been in initial mfr 1220648-2024-10327. b3 date of event was corrected from initial mfr 1220648-2024-10327 to 19-may-2022. b5 describe event or problem: additional narrative added for reported delivery issues that was not reported in initial mfr 1220648-2024-10327. h6: medical device problem code corrected to 2524.
27445 | 1220648-2024-10331 | 19172054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172054&pc=OZD | 06/17/2022 | 03/26/2025 |  | 04/24/2024 |  |  | 04/04/2024 | 04/24/2024 |  |  | 02/29/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194971 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, and the patient experienced a stroke, confirmed with a head computed tomography scan. the patient remained on impella support and was continued to be followed through successful wean and explant. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27446 | 1220648-2024-10332 | 19172189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172189&pc=OZD | 04/13/2022 | 03/26/2025 |  | 04/24/2024 |  |  | 04/13/2022 | 04/24/2024 |  |  | 12/31/2023 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162324 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, hemolysis was noted. this improved after adjustment of the impella support level and subsequent explantation. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27447 | 1220648-2024-10333 | 19172194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172194&pc=OZD | 04/08/2024 | 05/28/2024 |  | 04/24/2024 |  |  | 04/10/2024 | 04/24/2024 |  |  |  |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461489 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2024 |  |  |  |  | The user facility reported a 47-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that impella 5.5 pump support was ongoing for a patient in need of care escalation from a prior impella cp pump. a large hematoma at the surgical site was noted. the hematoma was unresolved and was worsening the next day so the patient was taken to the operating room for a hematoma evacuation and wound vac placement. the impella 5.5 pump support continued for 404 hours until weaned and left ventricle assist device was placed. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is from ifu for the impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states), which now includes statement "(including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
27448 | 1220648-2024-10334 | 19172293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172293&pc=OZD | 09/22/2022 | 03/26/2025 |  | 04/24/2024 |  |  | 09/22/2022 | 04/24/2024 |  |  | 04/30/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226057 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2022 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient had slight oozing at the access site. perclose was used to achieve hemostasis. the patient remained on impella support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27449 | 1220648-2024-10335 | 19172309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172309&pc=OZD | 04/02/2022 | 03/26/2025 |  | 04/24/2024 |  |  | 04/02/2022 | 04/24/2024 |  |  | 10/31/2023 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022153482 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2022 |  |  |  |  | The user facility reported a 50-year-old patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding from multiple sites including the impella access site. coagulation therapy was discontinued until bleeding was resolved. it was also reported there was a pump stop. medical team decided to proactively explant the impella cp as a result noting that the patient was hemodynamically stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27450 | 1220648-2024-10336 | 19172455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172455&pc=OZD | 05/11/2022 | 04/24/2024 |  | 04/24/2024 |  |  | 05/11/2022 | 04/24/2024 |  |  | 02/29/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192835 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/11/2022 |  |  |  |  | The complainant reported a 78 year old male patient was on an impella cp support and some oozing (bleeding) was noted after the sheath was removed. perclose was tightened with sutures and reangling was attempted, but a slight ooze still occurred. the impella was successfully weaned from the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
27451 | 1220648-2024-10337 | 19172461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172461&pc=OZD | 02/28/2023 | 04/24/2024 |  | 04/24/2024 |  |  | 04/15/2024 | 04/24/2024 |  |  | 06/30/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023233909 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 71 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Unspecified Infection | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported a 71-year-old female had prep-operation placement of an impella cp device for mechanical circulatory support. hemolysis was observed one day after using the impella cp pump, which was managed by lowering the performance level. additional information was obtained and further noted the patient developed an infection while on support and the patient's hemodynamics deteriorated. the impella pump was replaced, and extracorporeal membrane oxygenation was added. | [5 empty] | The investigation has been completed. no product was returned. the root cause of the hemolysis and infection could not be determined since the product was not returned and sufficient clinical details were not provided. instructions for use for the related events are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can catheter blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿
27452 | 1220648-2024-10339 | 19172506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172506&pc=OZD | 03/25/2022 | 10/30/2025 |  | 04/24/2024 |  |  | 03/25/2022 | 04/24/2024 |  |  | 09/30/2023 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023168676 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 03/25/2022 |  |  |  |  | The user facility reported a 62-year-old patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced hematuria and hemolyzed laboratory blood samples. the physician repositioned the impella cp multiple times under echocardiography (echo) and transthoracic echocardiography (tee). the pump was found to be positioned under papillary muscle and was repositioned. the patient was placed on dialysis and hematuria continued despite multiple attempts to reposition the pump. hemolysis continued to be present through explant of the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
27453 | 1220648-2024-10340 | 19172618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172618&pc=OZD | 04/13/2022 | 04/24/2024 |  | 04/24/2024 |  |  | 04/13/2022 | 04/24/2024 |  |  | 01/31/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186866 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/13/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support there was access site bleeding noted by the hematoma in the left groin and oozing from the repositioning sheath. the groin dressing was saturated, and the hematoma doubled in size. manual pressure was applied and perclose suture strings were tightened around the repositioning sheath, however, bleeding persisted, therefore the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
27454 | 1220648-2024-10342 | 19172692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172692&pc=OZD | 05/02/2022 | 05/23/2025 |  | 04/24/2024 |  |  | 05/02/2022 | 04/24/2024 |  |  | 01/31/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | GM |  | 2023187320 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/02/2022 |  |  |  |  | The complainant reported the 51 year old male patient needing mechanical circulatory support. while on impella support, complainant reported the patient was having frequent suction alarms. the complainant also reported the patient had a hematoma at the impella sheath site; manual pressure was held to compress. impella support weaned from p9 to p6; flows optimal at 2.6. impella site re-dressed and manual pressure held over site. unable to place fem-stop d/t impedance of venous cannula flow. heparin on hold due to supra therapeutic ptt. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27455 | 1220648-2024-10341 | 19172693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172693&pc=OZD | 05/11/2022 | 04/24/2024 |  | 04/24/2024 |  |  | 05/11/2022 | 04/24/2024 |  |  | 02/29/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP | [5 empty] | US |  | 2023198932 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/11/2022 |  |  |  |  | The complainant reported that a 61 year old male patient was implanted with an impella cp for mechanical circulatory support. while on support, access site bleeding occurred. initially, when the oozing had started, cinching down on the perclose was recommend and it caused the oozing to stop. however, when the patient arrived to the unit, the groin dressing was saturated. the tuohy borst valve was then tightened, and the site was redressed to angle match and the oozing stopped. additionally, when arriving to the unit, it was noted that the patient had dark red urine. labs were performed and the labs were partially hemolyzed. pump was repositioned, however, this did not resolve the hemolysis. the patient was experiencing ventricular tachycardia (vt) episodes over night and defibrillation was unable to break the vt. amio/lido bolus and drips were started and the decision was made to start the patient on extracorporeal membrane oxygenation (ecmo). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
27456 | 1220648-2024-10343 | 19172704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172704&pc=OZD | 04/13/2022 | 04/24/2024 |  | 04/24/2024 |  |  | 04/13/2022 | 04/24/2024 |  |  | 01/31/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186666 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/13/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the impella position was checked under fluoro, and the pigtail appeared to be in the pap muscle. when the staff tried to reposition, it popped back. the decision was made to leave the impella in the pap muscle during treatment. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
27457 | 1220648-2024-10344 | 19172745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172745&pc=OZD | 04/16/2024 | 05/28/2024 |  | 04/24/2024 |  |  | 04/16/2024 | 04/24/2024 |  |  |  |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467914 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/16/2024 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient developed a large hematoma to the axillary site. anticoagulant was held. site was redressed with surgicel. the hematoma was reduced after a washout and pressure bags overnight. the patient is stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to patient movement since the large hematoma developed to the axillary site with frank bleeding whilst ambulating. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 4114 (device not returned) and (3233) results pending completion of investigation were inadvertently omitted from the previously submitted report and have been added to this report. the investigation has been completed since the previously submitted report. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
27458 | 1220648-2024-10345 | 19172765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172765&pc=OZD | 04/14/2022 | 04/03/2025 |  | 04/24/2024 |  |  | 04/14/2022 | 04/24/2024 |  |  | 01/31/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189880 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/14/2022 |  |  |  |  | The user facility reported a 62-year-old female on impella cp for mechanical circulatory support. it was reported that a patient presented with severe heart failure so decision to implant impella cp was made, however the procedure was aborted due to vascular structure. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27459 | 1220648-2024-10346 | 19172771 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172771&pc=OZD | 07/10/2022 | 04/03/2025 |  | 04/24/2024 |  |  | 07/10/2022 | 04/24/2024 |  |  | 03/31/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212036 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/27/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had serious hemolysis. the impella flows were reduced from p-9 to p-4 and reddish urine output was noted. the impella cp was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27460 | 1220648-2024-10347 | 19172924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172924&pc=OZD | 04/13/2022 | 04/03/2025 |  | 04/24/2024 |  |  | 04/13/2022 | 04/24/2024 |  |  | 01/31/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188067 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support the inability to advance once through the valve due to catheter buckling and a significant amount of bleeding despite manual pressure being applied for forty-five (45) minutes was noted. therefore the impella cp was removed. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27461 | 1220648-2024-10349 | 19172948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172948&pc=OZD | 04/10/2022 | 04/03/2025 |  | 04/24/2024 |  |  | 04/10/2022 | 04/24/2024 |  |  | 12/31/2023 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177712 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/10/2022 |  |  |  |  | The user facility reported a 63-year-old patient in st elevation myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella access site. the access site was angle matched and redressed to resolve the oozing and the patient was transfused with one unit of packed red blood cells. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27462 | 1220648-2024-10348 | 19172949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19172949&pc=OZD | 04/17/2024 | 05/29/2024 |  | 04/24/2024 |  |  | 04/17/2024 | 04/24/2024 |  |  |  |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484115 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Cardiac Arrest | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2024 |  |  |  |  | The user facility reported a 37 year-old male with acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the implant, cardiopulmonary resuscitation (cpr) was performed. the impella device was successfully placed, cpr was continued, and cannulation for extra-corporeal membrane oxygenation (ecmo) was started in the left groin. the impella device was found to be in the ascending aorta and was successfully floated back into the left ventricle. when flows were increased, constant suction alarms were present with flows of 0.0 at all p-levels. the impella device had ingested a clot. the impella was exchanged for a new impella device, and the patient remains stable. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was submitted. the pump was not returned for investigation. based on clinical description & data analysis, the root cause of the low flow was most likely biomaterial ingestion. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27463 | 1220648-2024-10350 | 19173041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173041&pc=OZD | 04/17/2024 | 06/04/2024 |  | 04/24/2024 |  |  | 04/17/2024 | 04/24/2024 |  |  | 08/31/2025 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443683 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2024 |  |  |  |  | The user facility reported a 59-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced access site bleeding. manual compression was held, two units of blood products were provided, and dressing was change to treat the bleeding. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was most likely use related since 6fr terumo sheath, non-abiomed product was used. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
27464 | 1220648-2024-10351 | 19173181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173181&pc=OZD | 04/16/2022 | 04/03/2025 |  | 04/24/2024 |  |  | 04/16/2022 | 04/24/2024 |  |  | 01/31/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185925 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/16/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, patient developed limb ischemia. during impella placement, there were limited distal pulses noted in the right leg impella side. patient¿s pulsatility unexpectedly dropped in the intensive critical unit (icu). a distal perfusion catheter was placed, and a clot was noted in the distal perfusion catheter. the nurse aspirated the clot out, distal perfusion intact and flushed every hour. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided.
27465 | 1220648-2024-10352 | 19173260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173260&pc=OZD | 04/18/2024 | 06/10/2024 |  | 04/24/2024 |  |  | 04/18/2024 | 04/24/2024 |  |  |  |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025464502 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported a 36 year-old female with acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the medical staff discovered a cannula kink. the doctor scanned the catheter with fluoroscopy, which determined the kink was due to the patient¿s anatomy. the patient has a short aortic root. the impella cp device was then explanted. the patient survived the event. | [5 empty] | The investigation into the product damage (cannula kink) issue has been completed since the original report was submitted. the device was not returned for investigation. per the clinical information provided, the cause of the delivery issue was patient condition related due to patient anatomy as the patient has short aortic arch and led to catheter kink. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27466 | 1220648-2024-10353 | 19173279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173279&pc=OZD | 05/29/2022 | 10/01/2025 |  | 04/24/2024 |  |  | 08/09/2022 | 04/24/2024 |  |  | 12/31/2023 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167842 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2022 |  |  |  |  | The user facility reported a patient was being implanted with an impella cp device for mechanical circulatory support and at the end of the percutaneous coronary intervention, the position of the device shifted toward the posterior wall of the ventricle resulting in hemolysis requiring haptoglobin to be administered. it was also noted that there was renal failure, thrombocytopenia, and hematoma requiring a blood transfusion. the patient was successfully weaned, the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2024-10353 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-10353 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-10353 in accordance with updated procedures. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-10353. h6 health effect - clinical code 1886 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-10353. h6 medical device problem code 2906 was erroneously entered on the initial manufacturer device report number 1220648-2024-10353.
27467 | 1220648-2024-10354 | 19173292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173292&pc=OZD | 04/17/2022 | 04/03/2025 |  | 04/24/2024 |  |  | 04/17/2022 | 04/24/2024 |  |  | 09/30/2022 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1503049 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had access site bleeding. as a result, heparin concentration went down. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27468 | 1220648-2024-10355 | 19173320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173320&pc=OZD | 04/17/2022 | 04/03/2025 |  | 04/24/2024 |  |  | 04/17/2022 | 04/24/2024 |  |  | 01/31/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189121 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, patient developed hemolysis. the patient received 2 liters of blood with no change in central venous pressure (cvp). patient's urinary output was bloody, and labs confirmed hemolysis. the impella was repositioned, and more blood was given. hemolysis resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27469 | 1220648-2024-10356 | 19173394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173394&pc=OZD | 04/17/2022 | 04/03/2025 |  | 04/24/2024 |  |  | 04/17/2022 | 04/24/2024 |  |  | 01/31/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185319 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/17/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, it was noted that there were no pedal doppler signals. medication was required to get doppler pulses back. patient normalized. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates the cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided.
27470 | 1220648-2024-10357 | 19173456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173456&pc=OZD | 04/20/2022 | 04/03/2025 |  | 04/24/2024 |  |  | 04/20/2022 | 04/24/2024 |  |  | 12/31/2023 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170711 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2022 |  |  |  |  | The complainant reported a patient in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, a hematoma was noted after repo sheath was placed and sutured. physician advanced repo sheath, resutured, and redressed the site. manual pressure was held until femstop was placed at site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27471 | 1220648-2024-10358 | 19173465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173465&pc=OZD | 09/22/2022 | 04/03/2025 |  | 04/24/2024 |  |  | 09/22/2022 | 04/24/2024 |  |  | 11/30/2023 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022148933 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2022 |  |  |  |  | The user facility reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. while on impella support, the patient had severe bleeding from all surgical field sites. a sternotomy procedure was performed but no cause was able to be determined. no information related to the resolution of the bleeding was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27472 | 1220648-2024-10359 | 19173502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173502&pc=OZD | 04/12/2022 | 04/03/2025 |  | 04/24/2024 |  |  | 04/12/2022 | 04/24/2024 |  |  | 10/31/2023 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158835 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2022 |  |  |  |  | The user facility reported a 56-year-old patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, a large, unstable hematoma was noted at the impella access site after the patient was placed on extracorporeal membrane oxygenation. the vascular surgeon was consulted for treatment of the hematoma at the insertion site, and the decision was made to remove the impella cp to repair the vasculature. the patient was transfused with multiple blood products. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27473 | 1220648-2024-10360 | 19173531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173531&pc=OZD | 04/13/2022 | 04/02/2025 |  | 04/24/2024 |  |  | 04/13/2022 | 04/24/2024 |  |  | 11/30/2023 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023162683 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/13/2022 |  |  |  |  | The user facility reported a patient on impella cp for mechanical circulatory support. there was oozing around the site and pressure was held and vascular was consulted. the patient received blood due to a slow bleed at the access site. in addition, the patient experienced hemolysis and the impella cp was successfully weaned and explanted. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27474 | 1220648-2024-10361 | 19173586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173586&pc=OZD | 04/18/2024 | 06/05/2024 |  | 04/24/2024 |  |  | 04/18/2024 | 04/24/2024 |  |  |  |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486526 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Material Deformation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported a 36 year-old female with acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. as soon as the impella cp device was started, the pump generated suction alarms and flows decreased. the impella device was replaced with another impella device, and the same issues persisted. the medical staff determined the patient had a short aortic root and the impella device¿s cannula was kinked in the patient¿s anatomy. an intra-aortic balloon pump was placed instead. the patient survived the event. | [5 empty] | The investigation for the low pump flow and access site bleeding has been completed. the impella device was not returned for evaluation. data logs confirmed the failure mode and clinical narrative. logs showed after pump started, motor current (mc) spiked followed low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. based on clinical description and data analysis, the root cause of low flow was most likely a kinked cannula. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27475 | 1220648-2024-10362 | 19173625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173625&pc=OZD | 09/24/2022 | 04/02/2025 |  | 04/24/2024 |  |  | 09/24/2022 | 04/24/2024 |  |  | 06/30/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235824 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/24/2022 |  |  |  |  | The user facility reported a 67-year-old female placed on impella cp for mechanical circulatory support due to a non-st-elevation myocardial infarction (nstemi). the physician attempted to perclose the right femoral artery (rfa) impella site but failed. the physician was able to obtain contralateral left femoral artery (lfa) access. a thrombus was identified in rfa. the patient was then taken to the or for thrombectomy and possible surgical repair of the artery. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | Section d6a / d6b: added implant and explant dates. the cause of the thrombus was not determined since product was not returned and all the necessary information was not provided.
27476 | 1220648-2024-10363 | 19173683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173683&pc=OZD | 04/18/2024 | 06/10/2024 |  | 04/24/2024 |  |  | 04/18/2024 | 04/24/2024 |  |  |  |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461081 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while attempting to replace an impella cp with an impella 5.5, the patient anatomy prohibited advancing the pump beyond a point within the vessel between the graft anastomosis and the aorta. implantation of the impella 5.5 was aborted and the graft was over sewn. no patient harm was reported. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. per the clinical information provided, the root cause of the delivery issue was most likely patient condition related, impella implant was unsuccessful as patient anatomy prohibited advancing the pump beyond a point within the vessel between the graft anastomosis and the aorta. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
27477 | 1220648-2024-10364 | 19173740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173740&pc=OZD | 09/22/2022 | 04/02/2025 |  | 04/24/2024 |  |  | 09/22/2022 | 04/24/2024 |  |  | 04/30/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217566 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2022 |  |  |  |  | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient was successfully weaned off of the impella, but there was oozing at the access site. perclose x2 and angio seal were used to achieve hemostasis. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27478 | 1220648-2024-10366 | 19173743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173743&pc=OZD | 04/12/2022 | 04/02/2025 |  | 04/24/2024 |  |  | 04/12/2022 | 04/24/2024 |  |  | 12/31/2023 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163407 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/12/2022 |  |  |  |  | The user facility reported a 63-year-old patient in acute myocardial infarction and cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after the impella was implanted, the impella was dislodged from the left ventricle and moved into the aorta. the physician was unable to push the impella back into the left ventricle and the patient needed a short round of cardiopulmonary resuscitation. the physician was unable to pull the impella from the groin after it became stuck under a vascular patch in the femoral artery that had been previously placed from an operation several years prior. a vascular surgeon was able to remove the impella and close the access site and the patient was placed on extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
27479 | 1220648-2024-10365 | 19173744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173744&pc=OZD | 10/04/2022 | 03/26/2025 |  | 04/24/2024 |  |  | 10/04/2022 | 04/24/2024 |  |  | 01/31/2023 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022043912 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2022 |  |  |  |  | The complainant reported a patient with an unspecified primary indication was implanted with an impella cp device for mechanical circulatory support. several days after placement of the device, it began to display a high purge pressure. no further information is available, and the procedure outcome is unknown. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a added implant date. d6b: date of explant is unknown the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
27480 | 1220648-2024-10367 | 19173860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173860&pc=OZD | 04/18/2024 | 06/05/2024 |  | 04/24/2024 |  |  | 04/18/2024 | 04/24/2024 |  |  | 07/31/2025 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442995 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2024 |  |  |  |  | The user facility reported a 72-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was removed due to the patient having limited flow to the distal limb. the patient is stable. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the pump was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 added code 3038 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
27481 | 1220648-2024-10368 | 19173904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=19173904&pc=OZD | 06/29/2022 | 04/24/2024 |  | 04/24/2024 |  |  | 06/29/2022 | 04/24/2024 |  |  | 03/31/2024 |  | 04/24/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209330 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/29/2022 |  |  |  |  | The complainant reported a patient was on impella cp support and the patient experienced ischemia due to heavily diseased vessels throughout the entire left leg, including a chronic occlusion in the popliteal artery that reconstituted tibial arteries via small collaterals. staff were unable to find a pulse in the left leg. the patient was able to be weaned off all pressors and the impella was weaned. the pump was removed in the or and the left leg arteriotomy closed surgically. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.